## Complaint_Report_20180924_0024
id | summary | status | owner | component | resolution | time | cf_occur_date | cf_step_count | cf_affected_units | cf_account_list | cf_contact | cf_email | cf_phone | cf_country | cf_injury_description | cf_complainant | cf_product | cf_reportable_event | cf_service_type | cf_complaint_category | cf_investigation_type | cf_root_cause_type | description | cf_resolution | cf_root_cause | cf_investigation_summary
4108 | Richmond VA, broken flexyfoot (i-4320) | closed | jwitherspoon | UI - Crutch Interface | invalid | 2018-09-21 00:23 | 2018-08-28 10:23 | 257452 | EKSO-0150-00000000000021407005 | Unassigned | Dr. Ashraf Gorgey | ashraf.gorgey@va.gov | 8046755000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned | No investigation needed - see Justification | Wear | Richmond VA reported that the flexyfoot on the crutch had started to crack within the orange plastic section. | Replaced the lower section of the crutch with the flexyfoot pre-installed. | Flexyfoot was damaged due to normal repetitive use. | Flexyfoot was damaged due to normal repetitive use.
4107 | S and L torso pads no longer holding a (i-4311) | closed | jwitherspoon | Other | invalid | 2018-09-21 00:15 | 2018-08-16 12:06 | 442540 | EKSO-0319-00000000000021703002 | Unassigned | Christina |  |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned | No investigation needed - see Justification | Wear | Issue: Christina sent an email saying - I was wondering if you could send a small and large abdominal binder. the velcro is not holding anymore. | Replaced torso pads. | Torso pads beginning to wear due to normal use. | Torso pads beginning to wear due to normal use.
4106 | Phone connectivity at Bridging bionics (i-4309) | closed | jwitherspoon | Other | invalid | 2018-09-21 00:05 | 2018-08-15 13:09 | 383738 | EKSO-100378 | Unassigned | Maria Grufstedt |  |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Unassigned | No investigation needed - see Justification | Process Not Clear/Complete | Issue: Hi Lily and Darrell! We have some questions regarding the data we receive from Ekso pulse. It seems like that data we can find on our log in stops at April 18th of 2018. Is there a new log in and password for our new Ekso GT?? I thought since we are using the same client ID numbers it would just continue?? Let me know! We need the data from the new EKSO GT for a Grant due very soon. Thanks/ Maria Grufstedt Physical Therapist and Clinical Coordinator Bridging Bionics Foundation Follow up: Lily Hey Maria, You are correct - nothing should have changed, we transferred the customer number from you 1.1 device to your GT device. We looked into this issue and think that it is due to connectivity/ bad reception. It may help to move the device into an area that has better reception, if possible. Also, if you guys have access to Wifi then we may be able to connect you this way, I have CC'd Chad so that he can help diagnose as well. Let me know if you have any other questions! | Connected the device to WiFi and confirmed data sessions were shown in pulse. | Customer is located in a region without cell reception. | Customer is located in a region without cell reception.
4104 | side closures no longer holding (i-4304) | closed | jwitherspoon | Other | invalid | 2018-09-20 23:12 | 2018-08-10 14:13 | 237341 | EKSO-0194-00000000000021502003 | Parkwood Hospital - Canada | Barry Lynam |  |  | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned | N/A - Identified known wear item | Wear | Barry called in and reported their side closures are not holding the torso pads as well as they used to. | Replaced side closures. | This is a wear issue | Side closures beginning to wear due to normal use.
4102 | Hip width challenging to adjust (i-4301) | new | jwitherspoon | Adjustments - Hip |  | 2018-09-20 19:48 | 2018-08-09 13:16 | 40092 | EKSO-100376 | Scripps | Ira Keefer |  |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issue: One other component, The PTs also mentioned the Left Hip has been challenging to adjust the hip width. It‚Äôs considerably more difficult to adjust than usual, and it‚Äôs fully unlocked, when I tried it. It‚Äôs possible to adjust, but certainly something the field service technician can take a look at when that visit is scheduled. | The left hip abduction was replaced per MPI 102205/102208. | Root cause of the issue unknown. | A corner within the pocket where the locking key is inserted into the housing was marred. This was causing some stickiness to the locking key. The bearing retainer was also not set properly per manufacturing specifications. This was causing the retainer to rub against the torso bar.
4101 | Foot binding velcro peeling (i-4300) | closed | duanee | Pick an option | invalid | 2018-09-20 19:39 | 2018-08-09 12:46 | 40092 | EKSO-100376 | Scripps |  |  |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned | N/A - Identified known wear item | Wear | Mike reported during his training: One additional component to Scripps. When the field service technician goes to Scripps, the Velcro pieces glued to the foot binding are starting to separate. These could use some glue to increase adherence. | Velcro pads were re-secured to foot binding. | (Required for complaints) | Known design issue.
4095 | controller blacking out during use (I-4336) | closed | jwitherspoon | UI - Controller | fixed | 2018-09-15 04:37 | 2018-09-07 |  |  | Unassigned | Natalie Brandt | natalie.brandt@rehabkc.org | 8167517700 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Determined Unknown | Controller became unresponsive during use. | Replaced UI. | Root cause is undetermined. We will not be receiving the item for full root cause analysis. | Previous investigation completed (INV-0008)
4094 | Charger not holding a charge (I-4310) | closed | jwitherspoon | Battery - Charger | fixed | 2018-09-15 04:16 | 2018-08-15 | 6813 | EKSO-100403 | Unassigned | Dave Hancock | dahancock@dignityhealthevrehab.com | 4805408532 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | N/A - Identified known wear item | Customer Use | Dave at Dignity called and said one charger has not been charging their batteries and keeps giving them 'mismatched batt errors.' The one charger was going straight from orange to green. | Replaced battery charger. | Found that pins on the battery charger were damaged.
4093 | L Leg Rotating Inwardly (i-4289) | new | jwitherspoon | Lower Leg |  | 2018-09-14 23:19 |  | 251560 | EKSO-0228-00000000000021507005 | Unassigned | Krista Henshe | kristina.henshue@allina.com | 6128634200 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issue: I mentioned a month or so ago that our left lower leg was internally rotating with a particular patient. I changed the bolts/screws on the L lower leg but it has continued to happen. It often just happens as she is walking in a straight path. The right lower leg has started to internally rotate as well. This only happens with one patient (currently 3 are using the Ekso). This patient does have some spasticity and she is a bigger woman. However, we have to stop and adjust the legs approx. every 100‚Äô. Any recommendations? Follow up: We had a call with Krista and walked her through testing both legs with the T-Wrenches, the L side was easy to rotate even just by hand. | Directed the customer to clean off the particulate and replaced the adjustment screws. | (Required for complaints) | Lower leg was rotating due to a build up of particulate and well used adjustment screws.
4092 | Serial number logging as 0- (i-4287) | closed | jwitherspoon | Pulse Data | fixed | 2018-09-14 23:08 | 2018-07-06 13:10 | 4794 | EKSO-100397 | Unassigned | Chwee Foon Lim | ChweeFoon@eksobionics.com | 6582237972 | China | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | Issue: I am conducting training at Queen Mary hospital in Hong Kong these few days. From the controller, we noted that the serial number is 00000000, meaning there is no number assigned to this device delivered to Queen Mary (Hong Kong university). I believe this will create issue in data collection in the cloud. Could you kindly look into this urgently, please, as training is ongoing with patients and they may want to access cloud later stage to show them how to use the Pulse, Follow up: Chad - The battery which maintains the serial number likely ran out of charge. The best we can do remotely is to reset the customer number to ensure that pulse works. Please call me when you are available and I can walk you through the steps. | Reset the customer number to ensure that Pulse works. Unable to reset the serial number as that process requires a jumper to be installed. | Unable to determine the root cause. It was found that the coin cell battery which maintains the real clock did not cause this issue. | There is no safety issue associated with the reported symptoms/behavior. The root cause could not be determined. However, the customer issue was resolved by resetting the customer number in Ekso Pulse. There were no reported issues with the performance or safety of the Ekso GT. No further action is necessary. (D.Edmonds ‚Äì 14 Sep 2018)
4091 | Hip Abduction adjustment and Free Hip Ab-(i-4284) | new | jwitherspoon | Adjustments - Hip |  | 2018-09-14 22:36 | 2018-06-13 10:35 | 0 | EKSO-0336-00000000000021706004 | Unassigned | Jessica Spencer | jessica.a.spencer@osfhealthcare.org | 3096552000 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | I‚Äôm at OSF for their AT in Peoria IL. When demonstrating Squats and increasing Hip Abduction to greatest base of support, the right-side abduction adjustment became ‚Äòstuck‚Äô and directly affected the free hip abduction lock and ultimately unlocked the free hip abduction. I have seen this before and instructed the team to not go to the extreme end of +4 for the Hip Abduction when performing squats and approach the near end. This full team does not have a great attention to detail and I‚Äôm concerned this could potentially unlock without them noticing. Any other suggestions for preventing the chance of the Free hip abduction unlocking? Follow up on 7/25 from Jessica Spencer: I just wanted to touch base with you. I don't believe a service tech ever made it over here to check on/fix that one lock on the hip of our Ekso. We did receive the controller but to my knowledge no one physically was here. Is there anything I need to do to check into this? Josh: We have been working with them over the phone and all issues were resolved so no tech was needed. I‚Äôll reach out to Jessica right now. | replaced both left and right hip abductors. | (Required for complaints) | Excess play between the bearing sleeve and control bar was found which allowed for the hip abduction to be back driven.
4089 | Possible brake slip at Gaylord (I-4252) | new | jwitherspoon | Brake / Release / 9V Battery |  | 2018-09-14 22:14 | 2018-07-03 16:07 | 1010627 | EKSO-0165-00000000000021410002 | Gaylord Hospital |  | kcosgrove@eksobionics.com | 17077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor |  | Undetermined | This is coming from my personal email because the file was too large for Gaylord‚Äôs. Please see below for a video of the Ekso malfunction we had today. Please advise. We were not able to do the session today. The video is of one of our aides in it as we didn‚Äôt feel comfortable trying the patient with it. Leslie- it is our study patient who we couldn‚Äôt see. Please tell us what we need to do as far as the study goes. Thanks! | replaced joint board and programmed to latest software. | fluttering brake in left knee. | fluttering brake in left knee.
4086 | #363, Repeated Error (I-4241) | closed | jwitherspoon | Encoder / Potentiometer | fixed | 2018-09-12 19:40 | 2018-05-25 | 5296 | EKSO-100363 | Unassigned | Lisa Hughes |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Materials | Issues with encoder. Logan was here all morning working with our Ekso and we just put a patient in this afternoon. After maybe 40-50 steps including one left turn, we got another error and safe mode situation. This time it happened while walking straight vs turning. Logan is still here and we still have the loaner so we are getting it back out of the crate now and setting up to use while he continues to work on Avon‚Äôs Ekso. Minutes after Logan took our loaner to Beechwood this am after one successful patient session (38 min up time), we put our patient in and received another !#25 LH sensor error and safe mode, ending yet another of this same patient‚Äôs sessions for the day (on his first sit-stand). I spoke to Chad already and he is contacting Logan now to turn him around with the loaner so we can salvage the afternoon. Plan is in place. Thanks | The soft pot was removed and replaced with functioning pot per MPI 102205/102208. The device was then walked per 103939. | It was found that the errors were due to a bad batch of soft pots. | Issues with softpot
4083 | Missing data session in Pulse (i-4220) | closed | jwitherspoon | Pick an option | fixed | 2018-09-12 18:31 | 2018-06-12 14:04 | 2627 | EKSO-100384 | Unassigned | Rachel Wilkes | rachel.wilkes@kindred.com | 5123230756 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Process Not Followed | Last week my facility participated in Phase 1 Ekso training and I set up our Ekso Pulse account. Upon logging in this morning, I noticed an error in the analytics regarding sessions on 6/8/18. A session is missing from around 10:30 am our time with patient 0003. The session beginning at 1:08 states it was with patient 0003 with details from the 10:30 am session with patient 0003. There is a missing session that started around 1:00 that was with patient 0008 and there are no details in pulse regarding that session. Please let me know if this is some error on our part with set-up of the device or if there is some other error. | Changed time zone in Pulse account and verified that data was being shown correctly. | Pulse account was set up incorrectly. | The use of Ekso Pulse is not core to the safety or effectiveness of the Ekso GT. This issue has been corrected (see Resolution Description). No further action is necessary due to the very low risk level (Inconvenience) of this issue. (D.Edmonds ‚Äì 14 Sep 2018)
4080 | R toe guard broken (I-4216) | closed | jwitherspoon | Other | invalid | 2018-09-11 21:05 | 2018-06-12 12:01 | 24783 | EKSO-0312-00000000000021701001 | Loma Linda |  | kcosgrove@eksobionics.com | 17077474777 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Wear | Mike during ITW reported that: Right foot plate Toe guard just slid out. Perhaps the engineer coming down for the service update can glue that back into place. Thanks | re-glued toe guard. | Toe guard came loose and slid off. | Toe guard came loose and slid off.
4079 | #207, Snapped screw -Foot Plate to Ankle (I-4214) | closed | jwitherspoon | Foot Assy | fixed | 2018-09-11 20:57 | 2018-06-11 11:48 | 378810 | EKSO-0207-00000000000021503008 | Ekso Bionics | Kealoha, Logan |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Major | No investigation needed - see Justification | Design | !#207, Snapped screw -Foot Plate to Ankle | All four screws on right side were replaced. Screws on the other leg were verified to be fine. | Known failure point | INV-0001
4078 | Bent Walker (I-4213) | closed | jwitherspoon | Cane / Walker / Crutch | fixed | 2018-09-11 20:48 | 2018-06-08 13:35 | 2627 | EKSO-100384 | Unassigned | Joanna |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor |  | Mishandling | Issue: The left rear leg of the walker is slightly crooked. One of the PTs is very perseverative on this issue. There is also a little extra play in the left side of the walker. Picture attached. Any recommendations are appreciated. <image1.jpeg> Thanks, Joanna Sent from my iPhone | Replaced walker. | Walker frame became bent during shipment. | Walker frame became bent during shipment. The complaint database will be monitored to determine if the frequency of this issue increases and needs to be raised in priority. No further action is necessary at this time. (D.Edmonds ‚Äì 23 Sep 2018)
4077 | Battery not sitting correctly on charger (I-4212) | closed | duanee | Battery - Charger | fixed | 2018-09-11 20:36 | 2018-06-08 10:19 | 12165 | EKSO-0262-00000000000021601001 | Magee | Mary Schmidt PT - Outpatient Rehab Servic | mschmidt@mageerehab.org | (215) 587-3151 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | N/A - Identified known wear item | Wear | Battery not setting correctly on charger. Issue: When at Magee for our Open House on Monday, Liz mentioned that their battery chargers for the GT don‚Äôt often seat correctly on the batteries. When they think they have them set correctly and the orange light comes on, many times they come back to swap batteries and the light is green but one of the batteries may not be fully charged (many times less than 80%). It‚Äôs an intermittent problem and they know to double check the seating before they walk away but sometimes things happen. | Replaced batteries. | Issue was found to be due to normal battery wear. | Battery not sitting correctly on charger.
4075 | #379, Brake slip (I-4200) | closed | jwitherspoon | Brake / Release / 9V Battery | fixed | 2018-09-11 19:47 | 2018-06-05 11:28 | 9051 | EKSO-100379 | Unassigned | Marianne Tokar |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Process Not Followed | #379, Brake slip Issue Reported on 5/24: Good afternoon, We were working with our EKSO this PM and received x2 error messages 137-LK Motor Error with a note to contact EKSO. We were able to restart following the first error and continue. Our session ended with the 2nd error. Please advise. | The brake hub was removed and was reset to the proper manufacturing specification per MPI 102204/102207. | The brake hub was not set to proper manufacturing specifications. | The analysis of brake slip errors in the field resulted in the discovery that the software routine that checks for brake slip (a) checks for slip in both the extension and flexion directions, (b) has been triggered frequently in extension despite the fact that this is not the direction of potential harm, and (c) the number of both background slip detections and those that are presented to the customer across the field has caused more inconvenience to the customer than is warranted. This is supported by the lack of catastrophic failures in the field or safety issues (falls, etc.) caused by brake failure. As a result, the software behavior has been modified to limit the brake slip detection behavior based on data from the field to decrease the number of false detections. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2018)
4074 | Unit #142 Brake error (I-4198) | closed | jwitherspoon | Brake / Release / 9V Battery | fixed | 2018-09-11 19:36 | 2018-06-04 08:38 | 1044667 | EKSO-0142-00000000000021406003 | Unassigned |  | kcosgrove@eksobionics.com | 17077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Specification Not Correct | (Describe issue, including geographic location, if appropriate) | Parts were ordered. Replaced right side signal chain. cleaned and blew out motors, re-grounded all groung lugs. | (see Investigation Summary) | The analysis of brake slip errors in the field resulted in the discovery that the software routine that checks for brake slip (a) checks for slip in both the extension and flexion directions, (b) has been triggered frequently in extension despite the fact that this is not the direction of potential harm, and (c) the number of both background slip detections and those that are presented to the customer across the field has caused more inconvenience to the customer than is warranted. This is supported by the lack of catastrophic failures in the field or safety issues (falls, etc.) caused by brake failure. As a result, the software behavior has been modified (2.1.4) to limit the brake slip detection behavior based on data from the field to decrease the number of false detections. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2018)
4073 | Brake slip [I-4057] | closed | jwitherspoon | Brake / Release / 9V Battery | fixed | 2018-09-10 13:10 | 2018-05-25 | 9051 | EKSO-100379 | Unassigned | Craig Sisak |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Wear | Brake slipping resulting in 133 errors. | Reburnished brake and confirmed that the brake holding torque met specifications after all applicable testing. | The brake was found not to be holding to target torque due to wear. | The brake was found to not hold the target torque upon inspection. The brake was reburnished and retested to confirm that the holding torque met specifications. No further action is necessary. (D. Edmonds ‚Äì 20180910)
4072 | Noise Issue w/ device | new | jwitherspoon | Pick an option |  | 2018-06-21 21:58 | 2018-03-21 14:24 |  | 0196-00000000000021502005 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Noise Issue w/ device Incident 4066 | (Explain actions taken to resolve this issue) | (Required for complaints)
4070 | issue with batteries and/or charger (i-4069) | new | jwitherspoon | Battery - Charger |  | 2018-06-21 21:31 | 2018-03-21 | 0 | 0326-00000000000021704004 | Unassigned | Michal Janus | serwis@technomex.pl | +48 32 23 16 660 w. 62 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | !#326, issue with batteries and/or charger My client has a problem with ekso SN: 021704004: 1. The charger charges up to 40-50% and stops charging (green LED is on). I changed the charger. Old SN charger: C2170519. The new charger has no serial number. 2.The customer said the battery will be discharged very quickly. This may have a reason for the new motherboard, or a battery charger? https://photos.app.goo.gl/tCCR4NU1FQ7IKQmt1 3. They had a single case of locked knees. They do not have a photo of the error. From what they remember the message was: reboot system blocked knees. | (Explain actions taken to resolve this issue) | (Required for complaints)
4069 | issue with knee noise | new | jwitherspoon | Pick an option |  | 2018-06-21 21:26 | 2018-03-07 |  | 0147-00000000000021407002 | Unassigned | Robert Evans |  |  | South Africa | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #147, issue with knee noise Incident 4070 | (Explain actions taken to resolve this issue) | (Required for complaints)
4068 | issue with left leg irregular step | new | jwitherspoon | Adjustments - Leg |  | 2018-06-21 21:21 | 2018-03-19 |  | 0243-00000000000021510006 | Unassigned | Joan Romero | jromero@fundacionstepbystep.org | 34932231663 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #243, issue with left leg irregular step Incident 4071 | (Explain actions taken to resolve this issue) | (Required for complaints)
4067 | hip joint(l), high flex got stuck | new | jwitherspoon | Adjustments - Hip |  | 2018-06-21 21:16 | 2018-03-26 00:47 |  | 0315-00000000000021701004 | Unassigned | Matteo Carelli | m.carelli@emac.it | +39 347 5434417 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #315, hip joint(l), high flex got stuck Incident 4073 | (Explain actions taken to resolve this issue) | (Required for complaints)
4066 | left knee, right ankle | new | jwitherspoon | Lower Leg |  | 2018-06-21 21:10 | 2018-03-26 00:57 |  | 0197-00000000000021502006 | Unassigned |  |  |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #197, left knee, right ankle Incident 4074 | (Explain actions taken to resolve this issue) | (Required for complaints)
4065 | Hip width lock (I-4075) | new | jwitherspoon | Pick an option |  | 2018-06-21 21:06 | 2018-03-26 10:23 |  | 0109-00000000000021401002 | Ekso Bionics EU |  |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | I'm sorry to bother you with that but just this morning we had a problem with the loaner unit #109. As you can see in the picture attached, the bearing is broken (right hip). | (Explain actions taken to resolve this issue) | (Required for complaints)
4064 | Several Issues at Casa Colina | new | jwitherspoon | Pick an option |  | 2018-06-21 21:02 | 2018-03-27 16:24 |  |  | Unassigned |  |  |  |  | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Several Issues at Casa Colina Incident 4076 | (Explain actions taken to resolve this issue) | (Required for complaints)
4062 | Rotating Lower Left Leg | new | jwitherspoon | Adjustments - Leg |  | 2018-06-21 20:51 | 2018-03-12 | 147784 | 0340-00000000000021707001 | Unassigned | T.C. Doyle |  | 308-641-7456 |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Unit 340 - Rotating Lower Left Leg Incident 4084 | (Explain actions taken to resolve this issue) | (Required for complaints)
4061 | Chest Strap Stitching Issue | new | jwitherspoon | Soft Goods - Torso / Sacral |  | 2018-06-21 00:27 |  | 6427 | 0350-00000000000021708005 | Unassigned | Michael Glover | mglover@eksobionics.com | 6142962881 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #350 - Chest Strap Stitching Issue Incident 4085 | (Explain actions taken to resolve this issue) | (Required for complaints)
4060 | AMP board broken | new | jwitherspoon | PCBA |  | 2018-06-21 00:15 | 2018-04-03 01:24 | 92024 | 0242-00000000000021510005 | Ekso Bionics EU | Michael Blaschke |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #242, AMP board broken Incident 4086 | (Explain actions taken to resolve this issue) | (Required for complaints)
4059 | Battery board pin bent (i-4087) | new | jwitherspoon | Battery |  | 2018-06-21 00:09 | 2018-04-03 16:07 | 106423 | 0306-00000000000021611001 | United Hospital |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Process Not Followed | Jenn a PT at United reported the below issue: Writing about an EKSO battery issue we are having here at United with our EKSO GT. Noticed today that we went from 55% to 30% with single sit to stand with patient. So we decided to change out the batteries and when we replaced them and turned the device back on, it said R battery missing. We removed the battery and found a pin that was curved/loose. See picture below. How do we handle this? I am off the next two days so communication will likely need to be with Josh Tanner, our manager. Please advise. Thanks! Josh at United: An FYI-this is the second time this has happened on this right side. Thanks Josh Follow up (Internal at Ekso) Hey, Our warranty for medical is pretty wide open, it‚Äôs not a good idea to fix under warranty once and then charge for a 2nd time, plus having 27 pins without a plug isn‚Äôt the greatest of designs. We have 3 facilities in Minneapolis that need some type of service, Minn VA, CK and now United. We could hit all 3 centers and it wouldn‚Äôt cost us anything additional in terms of flights, hotels, etc. Rather charging, we should educate them on how to properly remove and place batteries on the Ekso to prevent this from happening again. Lily, I think clinical has a video of this in their pre- ITW training pack. Can you reach out to Mike and see if he can send this to you? Then we can send over as a refresher, we had a rocky start with them, I want to make sure they understand this is typically caused by banging the battery into the Ekso even if they can bend the pin back into place. Let‚Äôs try and get them back up working with the below people and see if we need to even visit their facility, we should definitely have a call though to help educate. Thanks everyone. Lily to United:Hey Josh,I hope all is well! When you have a chance, are you able to attempt to bend this one pin back into place, if you haven‚Äôt already attempted to do so. This issue actually has come up at various other centers and usually occurs when the battery is place on the device upside down (I know it can happen fairly easily especially with how busy you all are and how quickly you need to change in between patients.) If you aren‚Äôt able to bend the pin back ‚Äì this will likely need a field service visit to fix.If it is easier feel free to call our medical line or call me directly at 510-847-5602. Hopefully we can get this resolved for you soon. Thank you and talk to you soon! - Waiting customer feedback. | (Explain actions taken to resolve this issue) | (Required for complaints)
4058 | hip width not changeable | new | jwitherspoon | Adjustments - Hip |  | 2018-06-20 23:56 | 2018-04-04 04:39 |  | 0152-00000000000021407007 | BGU Tubingen - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | 4920380000000 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #152, hip width not changeable Incident 4088 | (Explain actions taken to resolve this issue) | (Required for complaints)
4057 | sensor issues (I-4089) | new | jwitherspoon | Encoder / Potentiometer |  | 2018-06-20 23:51 | 2018-04-04 06:35 | 346520 | 0134-00000000000021405001 | Chuck Slyngstad | Chuck Slyngstad |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor |  | Undetermined | unit #134 sensor issues | (Explain actions taken to resolve this issue) | (Required for complaints)
4056 | issue with foot sensors | new | jwitherspoon | Foot Sensor |  | 2018-06-20 23:45 | 2018-03-29 | 1829892 | 0233-00000000000021508004 | Unassigned |  |  |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #251, issue with foot sensors Incident 4090 | (Explain actions taken to resolve this issue) | (Required for complaints)
4055 | right hip cracked | new | jwitherspoon | Pick an option |  | 2018-06-20 23:42 | 2018-04-05 05:59 | 1829892 | 0233-00000000000021508004 | Unassigned |  |  |  | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #233, right hip cracked Incident 4091 | (Explain actions taken to resolve this issue) | (Required for complaints)
4054 | RK sensor errors | new | jwitherspoon | Pick an option |  | 2018-06-20 23:07 | 2018-03-13 | 42450 | 0314-00000000000021701003 | Bensonhurst |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #314, RK sensor errors Ekso 1.2 Operational Qualification Testing Incident 4092 | (Explain actions taken to resolve this issue) | (Required for complaints)
4053 | Controller Issues (I-4093) | closed | jwitherspoon | UI - Controller | fixed | 2018-06-20 23:02 | 2018-04-05 14:36 | 194133 | 0228-00000000000021507005 | Unassigned |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Materials | Krista, the main PT at Courage K has been reporting issues with the controller over the past few months (January~April) Initially the controller went black and we sent another controller to solve the issue. It worked fine for several weeks. The issues w/ the controller then persisted again w/ a slight shock and black screen, we sent a loaner for a different issue and the issue persisted. With the loaner she reported: Of course after I sent this it happened twice with my last patient. It keeps happening with the same patient. She is being seen 3x/week in the Ekso. Next week she is here Monday at 11, Wednesday at 10, Friday at 10 and the following week M,W,F all at 10. | The flat flex cable was replaced. | In addition to the known sensitivity to ESD that can result in the scrambling or clearing of the LCD display (see INV-0008), the flat flex cable was found to be bent upon inspection. This may have exacerbated the problem. | Previous investigation performed (INV-0008). See Root Cause Analysis.
4052 | LK 157 error | new | duanee | Software |  | 2018-06-20 22:59 | 2018-04-05 14:50 |  |  | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | LK 157 error Incident 4094 | (Explain actions taken to resolve this issue) | (Required for complaints)
4051 | No Problem Listed | new | jwitherspoon | Pick an option |  | 2018-06-20 22:54 | 2018-04-05 14:51 |  | 100359 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Ekso 1.2 Operational Qualification Testing incident 4095 | (Explain actions taken to resolve this issue) | (Required for complaints)
4050 | walking issues | new | jwitherspoon | Pick an option |  | 2018-06-20 22:50 | 2018-04-06 03:43 | 445683 | 0153-00000000000021407008 | Ekso Bionics EU |  |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #153, walking issues Ekso 1.2 Operational Qualification Testing Incident 4098 | (Explain actions taken to resolve this issue) | (Required for complaints)
4049 | cracked knee | new | jwitherspoon | Other |  | 2018-06-20 22:46 | 2018-04-09 06:36 | 346520 | 0134-00000000000021405001 | Unassigned | Oliese Stewart | ostewart@eksobionics.com | 5102055919 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | unit #134 cracked knee Incident 4099 | (Explain actions taken to resolve this issue) | (Required for complaints)
4048 | error #121-LR and #121 RK | new | jwitherspoon | Motor |  | 2018-06-20 22:41 | 2018-04-10 06:36 |  | 0138-00000000000021405005 | Unassigned |  |  |  | Switzerland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #138, error #121-LR and #121 RK Ekso 1.2 Operational Qualification Testing Incident 4103 | (Explain actions taken to resolve this issue) | (Required for complaints)
4047 | Right hip ext thrust (I-4104) | new | jwitherspoon | Adjustments - Hip |  | 2018-06-20 22:36 | 2018-04-10 14:58 | 155523 | 0244-00000000000021510007 | Kessler |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Right hip ext thrust Ekso 1.2 Operational Qualification Testing | (Explain actions taken to resolve this issue) | (Required for complaints)
4046 | Controller Issues | new | jwitherspoon | UI - Controller |  | 2018-06-20 22:32 | 2018-04-10 15:40 | 362302 | 0109-00000000000021401002 | Unassigned |  |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Controller Issues Ekso 1.2 Operational Qualification Testing Incident 4105 | (Explain actions taken to resolve this issue) | (Required for complaints)
4044 | issue with mainboard, error 121rk | new | jwitherspoon | Motor |  | 2018-06-20 22:23 | 2018-04-11 06:13 |  | 0326-00000000000021704004 | Unassigned | Michal Janus | serwis@technomex.pl | +48 32 23 16 660 w. 62 | Poland | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | #326, issue with mainboard, error 121rk Ekso 1.2 Operational Qualification Testing Incident 4107 | (Explain actions taken to resolve this issue) | (Required for complaints)
4043 | issue with error 21 HL | new | jwitherspoon | Encoder / Potentiometer |  | 2018-06-20 21:23 | 2018-04-11 07:24 |  | 100374 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 40 10 350 w. 38 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #374, issue with error 21 HL Ekso 1.2 Operational Qualification Testing Incident 4109 | (Explain actions taken to resolve this issue) | (Required for complaints)
4042 | Issues at GS | new | jwitherspoon | Pick an option |  | 2018-06-20 21:17 | 2018-04-12 11:49 |  | 0274-00000000000021602003 | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at GS Ekso 1.2 Operational Qualification Testing Incident 4111 | (Explain actions taken to resolve this issue) | (Required for complaints)
4041 | 121 and 137 error (I-4112) | new | jwitherspoon | Motor |  | 2018-06-20 20:05 | 2018-04-12 11:53 |  | 0134-00000000000021405001 | Chuck Slyngstad |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned | Investigation Needed | Undetermined | 121 and 137 error Ekso 1.2 Operational Qualification Testing | (Explain actions taken to resolve this issue) | (Required for complaints) | The analysis of brake slip errors in the field resulted in the discovery that the software routine that checks for brake slip (a) checks for slip in both the extension and flexion directions, (b) has been triggered frequently in extension despite the fact that this is not the direction of potential harm, and (c) the number of both background slip detections and those that are presented to the customer across the field has caused more inconvenience to the customer than is warranted. This is supported by the lack of catastrophic failures in the field or safety issues (falls, etc.) caused by brake failure. As a result, the software behavior has been modified (2.1.4) to limit the brake slip detection behavior based on data from the field to decrease the number of false detections. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2018)
4040 | Issues at KKI | new | jwitherspoon | Pick an option |  | 2018-06-20 19:49 |  |  |  | Kennedy Krieger Institute |  |  | 7077474777 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at KKI Incident 4113 | (Explain actions taken to resolve this issue) | (Required for complaints)
4039 | cracked knee and hip housing | new | jwitherspoon | Other |  | 2018-06-20 19:46 | 2018-04-12 14:40 | 346520 | 0134-00000000000021405001 | Chuck Slyngstad | Oliese Stewart | ostewart@eksobionics.com | 5102055919 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | unit #134 cracked knee and hip housing Ekso 1.2 Operational Qualification Testing Incident 4114 | (Explain actions taken to resolve this issue) | (Required for complaints)
4037 | issue with right knee | new | jwitherspoon | Pick an option |  | 2018-06-20 19:35 | 2018-04-17 05:35 |  | 0251-00000000000021511006 | Unassigned |  |  |  | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #251, issue with right knee Ekso 1.2 Operational Qualification Testing Incident 4116 | (Explain actions taken to resolve this issue) | (Required for complaints)
4036 | High pitch on device | new | jwitherspoon | Pick an option |  | 2018-06-20 19:27 | 2018-04-17 15:32 |  | 0279-00000000000021602008 | Sunnyview Rehab |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | High pitch on device Ekso 1.2 Operational Qualification Testing Incident 4119 | (Explain actions taken to resolve this issue) | (Required for complaints)
4035 | Patient Issue at United | new | jwitherspoon | Pick an option |  | 2018-06-20 19:21 | 2018-04-17 15:36 |  | 0306-00000000000021611001 | United Hospital |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Patient Issue at United Ekso 1.2 Operational Qualification Testing Incident 4120 | (Explain actions taken to resolve this issue) | (Required for complaints)
4034 | Issues w/ left knee | new | jwitherspoon | Pick an option |  | 2018-06-19 20:23 | 2018-04-17 15:40 |  | 0222-00000000000021505007 | Kennedy Krieger Institute |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues w/ left knee Ekso 1.2 Operational Qualification Testing Incident 4121 | (Explain actions taken to resolve this issue) | (Required for complaints)
4033 | Motor Overheating | new | jwitherspoon | Motor |  | 2018-06-19 20:19 | 2018-04-17 15:45 |  | 0303-00000000000021610002 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Motor Overheating Ekso 1.2 Operational Qualification Testing Incident 4123 | (Explain actions taken to resolve this issue) | (Required for complaints)
4032 | Issues at Marianjoy | new | jwitherspoon | Pick an option |  | 2018-06-19 20:15 | 2018-04-17 15:54 |  | 0282-00000000000021603001 | Marianjoy Rehab Hospital |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Marianjoy Ekso 1.2 Operational Qualification Testing Incident 4124 | (Explain actions taken to resolve this issue) | (Required for complaints)
4031 | Issues at OKC VA (i-4125) | new | jwitherspoon | Pick an option |  | 2018-06-19 20:11 | 2018-04-17 16:06 |  | 0180-00000000000021412003 | Unassigned |  |  | (405) 456-1000 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at OKC VA Ekso 1.2 Operational Qualification Testing Incident (i-4125) | Parts were ordered. Replaced both left and right ankles | issues with ballscrew. | issues with motor.
4030 | Issues at Sunnyview (I-4126) | new | jwitherspoon | Pick an option |  | 2018-06-19 20:07 | 2018-04-17 16:11 |  | 0279-00000000000021602008 | Sunnyview Rehab |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Sunnyview Ekso 1.2 Operational Qualification Testing | (Explain actions taken to resolve this issue) | (Required for complaints)
4029 | hip abduction changing during walk | new | jwitherspoon | Adjustments - Hip |  | 2018-06-19 20:03 | 2018-04-19 02:27 |  | 0260-00000000000021512005 | Unassigned | Michael Blaschke |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #260, hip abduction changing during walk Ekso 1.2 Operational Qualification Testing Incident 4127 | (Explain actions taken to resolve this issue) | (Required for complaints)
4027 | Broken glue bond | new | jwitherspoon | Pick an option |  | 2018-06-19 19:51 | 2018-04-19 16:15 |  | 0180-00000000000021412003 | Unassigned |  |  |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Broken glue bond Ekso 1.2 Operational Qualification Testing Incident 4130 | (Explain actions taken to resolve this issue) | (Required for complaints)
4026 | Issues at Burke | new | jwitherspoon | Pick an option |  | 2018-06-19 19:47 | 2018-04-19 16:25 |  | 0142-00000000000021406003 | Burke |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Burke Ekso 1.2 Operational Qualification Testing Incident 4131 | (Explain actions taken to resolve this issue) | (Required for complaints)
4025 | screws on footplate broken | new | jwitherspoon | Fasteners |  | 2018-06-19 19:43 | 2018-04-20 05:56 |  | 0221-00000000000021505006 | Ekso Bionics EU | Michael Blaschke |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #221, screws on footplate broken Ekso 1.2 Operational Qualification Testing Ekso GMBH Loaner/Demo Units Incident 4132 | (Explain actions taken to resolve this issue) | (Required for complaints)
4024 | serious issue with right knee | new | jwitherspoon | Pick an option |  | 2018-06-19 19:35 | 2018-04-23 06:07 | 1845641 | 0233-00000000000021508004 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 (0)479 62 99 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #233, serious issue with right knee Ekso 1.2 Operational Qualification Testing Incident 4133 | (Explain actions taken to resolve this issue) | (Required for complaints)
4023 | Issue with foot bindings | new | jwitherspoon | Foot Assy |  | 2018-06-19 19:28 | 2018-04-19 |  | 100370 | Unassigned | Martina Achmann | martina.achmann@auva.at | 4304092 | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #370, issue with foot bindings Ekso 1.2 Operational Qualification Testing Incident 4134 | (Explain actions taken to resolve this issue) | (Required for complaints)
4021 | battery charger cable defect | new | jwitherspoon | Battery - Charger |  | 2018-06-19 19:07 | 2018-04-24 03:31 |  | 100371 | Unassigned | Johnny Loughrey |  |  | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #371, battery charger cable defect Ekso 1.2 Operational Qualification Testing Incident 4135 | (Explain actions taken to resolve this issue) | (Required for complaints)
4020 | issue with left knee | new | jwitherspoon | Pick an option |  | 2018-06-19 19:03 | 2018-04-24 06:44 |  | 0183-00000000000021412006 | Unassigned |  |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #183, issue with left knee Ekso 1.2 Operational Qualification Testing Incident 4136 | (Explain actions taken to resolve this issue) | (Required for complaints)
4019 | issue with t-wrenches (I-4137) | new | jwitherspoon | Tools |  | 2018-06-19 18:57 | 2018-04-24 07:17 |  | 100370 | Unassigned | Bernadine |  |  | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #370, issue with t-wrenches Ekso 1.2 Operational Qualification Testing | (Explain actions taken to resolve this issue) | (Required for complaints)
4018 | issue with left hip cracked | new | jwitherspoon | Other |  | 2018-06-19 18:51 | 2018-04-24 08:38 |  | 0147-00000000000021407002 | Walking with Brandon Beack ‚Äì South Africa | Rob Evans |  |  | South Africa | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #147, issue with left hip cracked Ekso 1.2 Operational Qualification Testing Incident 4138 | (Explain actions taken to resolve this issue) | (Required for complaints)
4017 | Issues with loaner device | new | jwitherspoon | Pick an option |  | 2018-06-19 18:45 | 2018-04-24 13:41 |  | 0134-00000000000021405001 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Issues with loaner device Ekso 1.2 Operational Qualification Testing Incident 4139 | (Explain actions taken to resolve this issue) | (Required for complaints)
4016 | Sensory Error | new | jwitherspoon | Pick an option |  | 2018-06-19 18:41 | 2018-04-24 14:56 |  | 0116-00000000000021402005 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Sensory Error Ekso 1.2 Operational Qualification Testing Incident 4140 | (Explain actions taken to resolve this issue) | (Required for complaints)
4015 | Foot sensor issues at providence | new | jwitherspoon | Foot Sensor |  | 2018-06-19 18:37 | 2018-04-24 14:58 |  | EKSO-0286-00000000000021603005 | Providence |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Foot sensor issues at providence Ekso 1.2 Operational Qualification Testing Incident 4141 | (Explain actions taken to resolve this issue) | (Required for complaints)
4014 | L knee lock out, no movement in Free leg | new | jwitherspoon | Lower Leg |  | 2018-06-19 18:30 | 2018-04-24 15:02 |  |  | Unassigned |  |  |  |  | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | L knee lock out, no movement in Free leg incident 4142 | (Explain actions taken to resolve this issue) | (Required for complaints)
4013 | Possible clinical issues at Sunnyview | new | jwitherspoon | Pick an option |  | 2018-06-19 18:27 | 2018-04-25 13:47 |  | 0279-00000000000021602008 | Sunnyview Rehab |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Ekso 1.2 Operational Qualification Testing Incident 4143 | (Explain actions taken to resolve this issue) | (Required for complaints)
4012 | Controller Issue | new | jwitherspoon | UI - Controller |  | 2018-06-19 18:19 | 2018-04-25 13:51 |  |  | Burke |  |  | 7077474777 | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Controller Issue Incident 4144 | (Explain actions taken to resolve this issue) | (Required for complaints)
4011 | hip adjustment problem | new | jwitherspoon | Adjustments - Hip |  | 2018-06-19 17:59 | 2018-04-26 08:23 |  | 0252-00000000000021511007 | Unassigned | Nikolas Malizos |  |  | Greece | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #252, hip adjustment problem Ekso 1.2 Operational Qualification Testi Incident 4145 | (Explain actions taken to resolve this issue) | (Required for complaints)
4010 | No Problem Listed | new | jwitherspoon | Pick an option |  | 2018-06-19 17:55 | 2018-04-26 14:17 |  | 0279-00000000000021602008 | Sunnyview Rehab |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Ekso 1.2 Operational Qualification Testing Incident 4147 | (Explain actions taken to resolve this issue) | (Required for complaints)
4009 | issue with left leg and batteries | new | jwitherspoon | Battery |  | 2018-06-19 17:50 | 2018-04-26 |  | 0251-00000000000021511006 | Unassigned | Rory Brennan | rory.brennan@accuscience.ie |  | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #251, issue with left leg and batteries Ekso 1.2 Operational Qualification Testing Incident 4148 | (Explain actions taken to resolve this issue) | (Required for complaints)
4008 | epoxy on ballscrew broken | new | jwitherspoon | Ball Screw |  | 2018-06-19 17:43 | 2018-04-30 06:55 |  | 0129-00000000000021404004 | Unassigned |  |  |  | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #129, epoxy on ballscrew broken Ekso 1.2 Operational Qualification Testing Incident 4149 | (Explain actions taken to resolve this issue) | (Required for complaints)
4007 | error 43 | new | jwitherspoon | Foot Sensor |  | 2018-06-19 17:37 | 2018-05-02 02:16 |  | 0152-00000000000021407007 | Unassigned |  |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #152, error 43-Left toe pressure Ekso 1.2 Operational Qualification Testing Incident 4150 | (Explain actions taken to resolve this issue) | (Required for complaints)
4006 | Defect t-wrenches | new | jwitherspoon | Other |  | 2018-06-19 17:26 | 2018-05-03 09:08 |  | 0315-00000000000021701004 | Unassigned |  |  |  | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #315, defect t-wrenches Ekso 1.2 Operational Qualification Testing Incident 4152 | (Explain actions taken to resolve this issue) | (Required for complaints)
4005 | Error 27 at Kessler (i-4153) | new | jwitherspoon | Electrical Interconnect |  | 2018-06-19 17:20 | 2018-05-03 12:23 | 336953 | 041-00000000000021510004 | Kessler |  |  | (973) 243-6873 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Materials | Error 27 at Kessler Issue: I wish I was writing to you under better circumstances, but we had a suit error today at the beginning of a session and I was hoping that you could offer some insight/guidance on what we may need to do. For reference, I have attached a screenshot of the session information from Pulse (serial: 021510004). The participant was standing in the suit for about 13 minutes (no steps taken yet), but when they took the first right step, the suit threw an Error 24 (R Knee) and said to contact EksoBionics. We sat the participant manually, restarted the suit, and tried again, but got the same error. To note, this is our non-FES device, and not the suit where we had the right knee rebuilt last year. Lily- I hope you had a good weekend. Sorry again to hear about the issues you guys are having. We decided it would be best to send you guys a loaner and bring your guys device back to HQ. We are preparing a loaner device today and can have it to you by Wednesday morning (hopefully this isn‚Äôt too much of an inconvenience.) Let us know if this works for you and your team or if you have any other questions or concerns. *Loaner device sent and customer device returned to HQ* | The right side signal chain replaced. | The cause was with the wiring, | Follow up: Katie-There‚Äôs some back screen pot/encoder mismatches on the LH so we may want to recal that. The right knee seems to be extra angry though. It‚Äôs indicating a bad A/B transition off the FPGA, communication issue, extra index‚Ä¶ If it‚Äôs recurring, I‚Äôd guess it‚Äôs actually wiring or board. Katie-Eric agrees with my board or wiring assessment. When you service it, be sure to take a look at that LH just in case something is starting to fail there.
4004 | Issues at Bridging Bionics | new | jwitherspoon | Pick an option |  | 2018-06-19 17:13 | 2018-05-03 15:41 |  | 0059-00000000000011306001 | Unassigned |  |  | 7077474777 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Bridging Bionics Ekso 1.2 Operational Qualification Testing Incident 4154 | (Explain actions taken to resolve this issue) | (Required for complaints)
4003 | issue with foot sensors | new | jwitherspoon | Foot Sensor |  | 2018-06-19 00:33 | 2018-05-07 02:29 |  | 0140-00000000000021406001 | Unassigned | Matteo Carelli | m.carelli@emac.it |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #140, issue with foot sensors Ekso 1.2 Operational Qualification Testing Incident 4157 | (Explain actions taken to resolve this issue) | (Required for complaints)
4002 | Issues at Renaissance | new | jwitherspoon | Pick an option |  | 2018-06-19 00:28 | 2018-05-08 14:49 |  | 0305-00000000000021610004 | Doctors Hospital at Renaissance |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Renaissance Ekso 1.2 Operational Qualification Testing Incident 4159 | (Explain actions taken to resolve this issue) | (Required for complaints)
4001 | Leg Errors at Courage K | new | jwitherspoon | Pick an option |  | 2018-06-19 00:23 | 2018-05-08 15:41 |  | 0228-00000000000021507005 | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Leg Errors at Courage K Ekso 1.2 Operational Qualification Testing Incident 4161 | (Explain actions taken to resolve this issue) | (Required for complaints)
4000 | battery release button malfunction | new | jwitherspoon | Battery - Connector |  | 2018-06-19 00:17 | 2018-05-08 15:56 |  |  | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | battery release button malfunction | (Explain actions taken to resolve this issue) | (Required for complaints)
3999 | Motor error at MJ | new | jwitherspoon | Motor |  | 2018-06-19 00:14 | 2018-05-08 15:59 |  | 0202-00000000000021503003 | Marianjoy Rehab Hospital |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Motor error at MJ Ekso 1.2 Operational Qualification Testing Incident 4163 | (Explain actions taken to resolve this issue) | (Required for complaints)
3998 | Hip Width Stuck | new | jwitherspoon | Adjustments - Hip |  | 2018-06-19 00:06 | 2018-05-08 16:02 |  |  | Kennedy Krieger Institute |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Hip Width Stuck | (Explain actions taken to resolve this issue) | (Required for complaints)
3997 | Broken Hip Adjustment | new | jwitherspoon | Adjustments - Hip |  | 2018-06-19 00:03 | 2018-05-08 16:28 |  |  | Rehab Institute of Kansas City |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Broken Hip Adjustment | (Explain actions taken to resolve this issue) | (Required for complaints)
3996 | amp board faulty, noise left knee | new | jwitherspoon | Pick an option |  | 2018-06-18 23:59 | 2018-04-18 |  | 0302-00000000000021610001 | Unassigned |  |  |  | EMEA Region-Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #302, amp board faulty, noise left knee Ekso 1.2 Operational Qualification Testi Incident 4166 | (Explain actions taken to resolve this issue) | (Required for complaints)
3994 | Foot sensor issues at Sunnyview (I-4168) | closed | jwitherspoon | Foot Sensor | fixed | 2018-06-18 23:47 | 2018-05-09 |  | 0279-00000000000021602008 | Unassigned |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Materials | Foot sensor issues at Sunnyview | The foot sensors were replaced on the unit. New sensors were mailed to the customer. | (see INV-0006) | Previous investigation completed: INV-0006
3993 | Threading coming undone on strap | new | jwitherspoon | Soft Goods - Shoulder Straps |  | 2018-06-18 23:42 | 2018-05-10 09:43 |  |  | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Threading coming undone on strap Incident 4170 | (Explain actions taken to resolve this issue) | (Required for complaints)
3992 | Error 27 lk | new | jwitherspoon | Pick an option |  | 2018-06-18 23:39 | 2018-05-11 06:15 |  | 0326-00000000000021704004 | Unassigned |  |  |  | EMEA Region-Poland | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | #326, error 27 lk Ekso 1.2 Operational Qualification Testing Incident 4171 | (Explain actions taken to resolve this issue) | (Required for complaints)
3991 | UI not working | new | jwitherspoon | UI - Controller |  | 2018-06-18 23:32 | 2018-05-11 06:26 |  | 0352-00000000000021708007 | Unassigned |  |  |  | EMEA Region-Poland | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | #352, UI not working Ekso 1.2 Operational Qualification Testing Incident 4172 | (Explain actions taken to resolve this issue) | (Required for complaints)
3990 | Missing Medium Tibial Shims | new | jwitherspoon | Fasteners |  | 2018-06-18 23:24 | 2018-05-16 10:01 |  | 100379 | Unassigned |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Missing Medium Tibial Shims Ekso 1.2 Operational Qualification Testing Incident 4173 | (Explain actions taken to resolve this issue) | (Required for complaints)
3989 | Concern at Madona | new | jwitherspoon | Pick an option |  | 2018-06-18 23:15 | 2018-05-16 10:20 |  | 0280-00000000000021602009 | Madonna |  |  |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Concern at Madonna Ekso 1.2 Operational Qualification Testing Incident 4174 | (Explain actions taken to resolve this issue) | (Required for complaints)
3988 | Pin in ankle exposed | new | jwitherspoon | Fasteners |  | 2018-06-18 21:04 | 2018-05-16 11:02 |  |  | Unassigned |  |  |  |  | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Pin in ankle exposed Incident 4175 | (Explain actions taken to resolve this issue) | (Required for complaints)
3987 | Software Issue at Advanced PT | new | duanee | Software |  | 2018-06-18 20:59 | 2018-05-16 15:30 |  |  | Advanced PT |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Software Issue at Advanced PT Incident 4177 | (Explain actions taken to resolve this issue) | (Required for complaints)
3986 | Foot sensor issues (I-4178) | closed | jwitherspoon | Foot Sensor | fixed | 2018-06-18 19:59 | 2018-05-17 13:03 |  | 0229-00000000000021507006 | TIRR |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Materials | Foot sensor issues | Replaced foot sensors | (see INV-0006) | Previous investigation performed (INV-0006)
3985 | Error 41, toe pressure | new | jwitherspoon | Foot Sensor |  | 2018-06-18 19:50 | 2018-05-18 07:53 |  | 100371 | Unassigned | Michael Blaschke |  |  | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #371, error 41, toe pressure Ekso 1.2 Operational Qualification Testing Incident 4179 | (Explain actions taken to resolve this issue) | (Required for complaints)
3984 | broken hip adjustment unit #263 (I-4180) | closed | jwitherspoon | Adjustments - Hip | fixed | 2018-06-18 19:39 | 2018-05-21 | 556000 | 0263-00000000000021601002 | Rehab Institute of Kansas City | Oliese Stewart |  | 5102055919 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Customer Use | broken hip adjustment unit. hip tube broken pins for brass adjustment inside tube. | replaced hip tube and hip abductors. |  | The issue was isolated to the hip tube and hip abductors. These components were replaced, and the service technician confirmed the device was operating properly prior to return to the customer. No further action is necessary. (D.Edmonds ‚Äì 24 Sep 2018)
3983 | #171, issue with hip width screw | new | jwitherspoon | Adjustments - Hip |  | 2018-06-18 19:27 | 2018-05-23 05:30 |  | 0171-00000000000021410008 | Unassigned | Michael / Rafael Calavia |  | +49 173 155 9837 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #171, issue with hip width screw Ekso 1.2 Operational Qualification Testing Incident 4181 | (Explain actions taken to resolve this issue) | (Required for complaints)
3982 | Crutch Interface Issue | new | jwitherspoon | UI - Crutch Interface |  | 2018-06-18 19:16 | 2018-05-23 13:25 |  | 0312-00000000000021701001 | Loma Linda |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Crutch Interface Issue Ekso 1.2 Operational Qualification Testing Incident 4182 | (Explain actions taken to resolve this issue) | (Required for complaints)
3981 | Hip Width Issue | new | jwitherspoon | Adjustments - Hip |  | 2018-06-18 19:08 | 2018-05-23 13:27 |  | 0312-00000000000021701001 | Loma Linda |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Hip Width Issue Ekso 1.2 Operational Qualification Testing Incident 4183 | (Explain actions taken to resolve this issue) | (Required for complaints)
3980 | Issues at Avon - joint cals | new | jwitherspoon | Pick an option |  | 2018-06-18 18:35 | 2018-05-23 13:34 |  |  | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Unassigned |  | Undetermined | Issues at Avon - joint cals Ekso 1.2 Operational Qualification Testing Incident 4184 | (Explain actions taken to resolve this issue) | (Required for complaints)
3979 | #344, left hip motor hot with smoke | new | jwitherspoon | Motor |  | 2018-06-18 18:10 | 2018-05-22 | 34419 | 0344-00000000000021707005 | Unassigned | Rafael Calavia | rcalavia@rebiotex.com | +34 620 015 761 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #344, left hip motor hot with smoke incident #4186 Ekso 1.2 Operational Qualification Testing | (Explain actions taken to resolve this issue) | (Required for complaints)
3978 | damaged parts (i-4187) | new | jwitherspoon | Other |  | 2018-06-18 17:59 | 2018-05-24 10:55 | 766890 | EKSO-0181-00000000000021412004 | Unassigned | Oliese Stewart | ostewart@eksobionics.com |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | cracked knee housing. broken knee brake switch cable. and Ankle axle flange broken. Damaged Donn Doff links. | (Explain actions taken to resolve this issue) | (Required for complaints)
3977 | #185, soft g., ddl | new | jwitherspoon | Soft Goods |  | 2018-06-18 17:52 | 2018-05-28 09:15 |  | 0185-00000000000021501002 | Unassigned | Michal Janus | serwis.@technomex.pl |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #185, soft g., ddl Ekso 1.2 Operational Qualification Testing Incident 4191 | (Explain actions taken to resolve this issue) | (Required for complaints)
3976 | #113, error #133 and 121 LH | new | jwitherspoon | Motor |  | 2018-06-18 17:44 | 2018-05-29 05:59 |  | 0113-00000000000021402002 | Unassigned | Michael Blaschke |  |  | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #113, error #133 and 121 LH Ekso 1.2 Operational Qualification Testing Incident 4193 | (Explain actions taken to resolve this issue) | (Required for complaints)
3975 | #375, err 135lh, smoke from motor | new | jwitherspoon | Motor |  | 2018-06-18 17:33 | 2018-06-01 06:19 |  | 100375 | Unassigned |  |  |  | Switzerland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | #375, err 135lh, smoke from motor Ekso 1.2 Operational Qualification Testing Incident 4195 | (Explain actions taken to resolve this issue) | (Required for complaints)
3974 | Ankle Sprain (I-4242) | closed | jwitherspoon | Pick an option | fixed | 2018-06-13 17:39 | 2018-03-18 | 653112 | EKSO-0016-00000000000011203002 | WakeMed | Nicolle Lang |  | (516) 659-6931 | United States | ankle sprain | Same as contact listed | ekso 1.1 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | Pt is a T4 SCI. She has a regular standing program and walks with KAFOs. She had used the EKSO last June 2017 and did well but had a R ankle sprain as her R ankle tried to supinate in the EKSO. She returned to try the EKSO again in Dec 2017 and this time has been wearing a R ankle brace to prevent this and has had no issues. She has been using the EKSO 1 x week most of this time. She has been doing well with no issues. On 5/15/18, we had difficulty with sit to stand with not being able to complete standing. On the third attempt we were able to lean forward more and push up to stand with the device slow to acknowledge standing. Pt was able to ambulate 60 ft. Pause due to feeling of hips being off. PT adjusted L leg length 4 points as is was slightly longer than the R. (in sitting). Again had difficulty with sit to stand, but able to push up with extra assist by PT and pt on 2nd attempt. Once up pt was able to ambulate another 100ft. Lowered to sit and EKSO removed. Noted R knee swelling. Swelling remained overnight and pt sent to Dr. X-ray showed R tibia Plateau fx. | Patient was advised to discontinue use of exoskeletons. No further action required unless adverse trends are identified. (see Investigation Summary) 100 point check executed. | Cannot be positively determined. The patient‚Äôs fractured tibia was likely due to minor trauma to fragile osteoporotic bones. The medical device did not malfunction, was used on label and in a supervised clinical environment. (see Investigation Summary) | (see attached Complaint Investigation WakeMed 20180612.pdf)
3973 | Re-set at advanced PT | new | lkealoha | Software |  | 2018-06-07 20:05 | 2018-05-10 |  | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Unassigned |  | Undetermined | Issue: There was a data reset on a 2.1.2 device, phone ID 0017ea1f30d1. This reset affected the serial number, customer number, lockout fields, and patient ID fields. Attached is the header where the device first reported that the fields had changed. Is there anything in the 2.1.2 software that would do this? There isn‚Äôt anything in the error log which stated that someone accessed one of the back menus, so I‚Äôm not sure what is going on. * 2 loaners sent for this issue. | (Explain actions taken to resolve this issue) | (Required for complaints)
3972 | Unable to move left knee | new | jwitherspoon | Pick an option |  | 2018-06-05 01:37 |  |  | 100361 | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor |  | Undetermined | Unable to get movement in Free Legs when device is on. Left knee will not release. Arlington (Kindred Site) - found on 4/24/2018 in Americas Region (Device Serial Number 100361) Possible external cable damage Background ‚Ä¢ Found during initial training. Device manual tells not to pull on these cables. Training also emphasizes the fact that the cables are not to be used to handle the device. Key Goals ‚Ä¢ Prevent customer from pulling on the external cables to adjust the device. | (Explain actions taken to resolve this issue) | (Required for complaints)
3971 | Foot Sensor Error (Error 41) (4157) | new | jwitherspoon | Foot Sensor |  | 2018-06-05 01:27 | 2018-04-11 |  | 100367 | Unassigned | Sebastian? |  |  | EMEA Region |  | Same as contact listed | Pick an option | FALSE | Unassigned | Minor |  | Undetermined | Background ‚Ä¢ Found during initial training Root Cause Analysis ‚Ä¢ Process not followed. Incorrect cable wiring of the foot sensor ‚Ä¢ Found that the adhesive strain relief was not done per the process ‚Ä¢ This error was embedded by a technician in training. ‚Ä¢ We will make sure the MPI is clean and that we are following the correct process. ‚Ä¢ Key Goals ‚Ä¢ ÔÉºWe will make sure the MPI is clean and that we are following the correct process going forward. Completed ‚Ä¢ Ensure that the process is being followed and all technicians are properly trained. Concerns ‚Ä¢ We can‚Äôt inspect for these issues without disassembling the current finished goods in a manner that would cause damage to the product. Strong Points ‚Ä¢ A foot sensor replacement can be performed by the customer at low cost. The part is mailed along with a kit for replacement. Accuscience ‚Äì found on 4/11/2018 in EMEA Region (Device Serial Number 100367) Background ‚Ä¢ Found during initial training Root Cause Analysis ‚Ä¢ Process not followed. Incorrect cable wiring of the foot sensor ‚Ä¢ Found that the adhesive strain relief was not done per the process ‚Ä¢ This error was embedded by a technician in training. Key Goals ‚Ä¢ ÔÉºAssure the MPI is correct and that the process is being followed and all technicians are properly trained. Completed Concerns ‚Ä¢ We can‚Äôt inspect for these issues without disassembling the current finished goods in a manner that would cause damage to the product. Strong Points ‚Ä¢ A foot sensor replacement can be performed by the customer at low cost. The part is mailed along with a kit for replacement. | (Explain actions taken to resolve this issue) | (Required for complaints)
3970 | Multiple error 27 and 51 during use. (4184?) | new | duanee | Encoder / Potentiometer |  | 2018-06-04 19:43 | 2018-05-23 |  | 100363 | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor |  | Undetermined | Avon (Kindred Site) ‚Äì Occurred on 5/23/2018 in Americas Region (Device Serial Number- 100363) Issue: ‚Ä¢ Multiple error 27 and 51 during use. Background ‚Ä¢ Issue found during initial training. This device is still having sensors issues even after the replacement. Key Goals ‚Ä¢ ÔÉºAll finished goods must be checked for the encoder cable damage. Completed by Al 5/30/2018 ‚Ä¢ Work with Spectra to determine a method for testing for potential signal issues with the softpot. ÔÉò Meeting scheduled with supplier for 6/4/2018. ‚Ä¢ ÔÉºA fixture has been created to test all potentiometers after install. ÔÉò We are currently developing the code and setting parameters for failure. ÔÉò The fixture will be implemented in the next few weeks. ÔÉò Working with Spectra to determine an accurate way of identifying and testing these issues after installation. A fixture has been created to test all potentiometers after install. ÔÉò We are currently developing the code and setting parameters for failure. ÔÉò The fixture will be implemented in the next few weeks. ÔÉò Working with Spectra to determine an accurate way of identifying and testing these issues after installation. ‚Ä¢ ‚Ä¢ ÔÉºA fixture has been created to test all potentiometers after install. ÔÉò We are currently developing the code and setting parameters for failure. ÔÉò The fixture will be implemented in the next few weeks. ÔÉò Working with Spectra to determine an accurate way of identifying and testing these issues after installation. ‚Ä¢ ÔÉºA fixture has been created to test all potentiometers after install. ÔÉò We are currently developing the code and setting parameters for failure. ÔÉò The fixture will be implemented in the next few weeks. ÔÉò Working with Spectra to determine an accurate way of identifying and testing these issues after installation. A fixture has been created to test all potentiometers after install. We are currently developing the code and setting parameters for failure.The fixture will be implemented in the next few weeks. Working with Spectra to determine an accurate way of identifying and testing these issues after installation. The potentiometer issues have been presenting themselves in a new manner and we have not built in anything to look for this failure mode. This is actively being worked on. Main Concerns ‚Ä¢ There have been 3 devices that have had entire signal chain replacements for these issues, costing money in parts and labor. ‚Ä¢ The tang for the softpot not fully seating was an issue with the encoder cable (102504). The encoder cable issue was brought up by Chad for the Avon site due to a 27 error. This was the issue with the crimp on the cable not fully being seated onto the jacket of the cable. It is undetermined if this was the cause of the issue because they started having a 51 error shortly after replacement. There have been no field failures directly related to the tang not fully seating. ‚Ä¢ The encoder crimp issue is a one-off that we are unsure if an error was even caused by as it passed all applicable testing in house and the device was still having issues after replacement. We will not be able to accurately determine a root cause of this issue because we do not have the device and the parts that were replaced may have been damaged during removal. | (Explain actions taken to resolve this issue) | Root Cause Analysis ‚Ä¢ Materials issue. Found a bad crimp in encoder cable. The supplier has been given a SCAR form to identify the root cause. (Spectra Symbol Corp. ‚Äì SCAR #12635) ‚Ä¢ Without receiving the failed device or components in question, we will not be able to properly determine the root cause of this issue. The device was still having errors (51) after replacement of the encoder cable and continued to have these errors until the softpot was replaced. The potentially defective softpot will be returned to HQ and inspected on 6/4/18.
3969 | 137 Error (brake slip) asking to reboot the device. | new | duanee | Brake / Release / 9V Battery |  | 2018-06-04 19:11 | 2018-05-24 |  | 100379 | Unassigned |  |  |  | United States | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor | Investigation Needed | Undetermined | Beechwood (Kindred Site) - Occurred on 5/24/2018 in Americas Region (Device Serial Number 100379) Issue ‚Ä¢ 137 Error (brake slip) asking to reboot the device. This error and impact will not harm a patient. ‚Ä¢ Brake hub was installed at a gap larger than specified in the manufacturing process Background ‚Ä¢ Issue found 1 week after initial training | (Explain actions taken to resolve this issue) |  | The analysis of brake slip errors in the field resulted in the discovery that the software routine that checks for brake slip (a) checks for slip in both the extension and flexion directions, (b) has been triggered frequently in extension despite the fact that this is not the direction of potential harm, and (c) the number of both background slip detections and those that are presented to the customer across the field has caused more inconvenience to the customer than is warranted. This is supported by the lack of catastrophic failures in the field or safety issues (falls, etc.) caused by brake failure. As a result, the software behavior has been modified (2.1.4) to limit the brake slip detection behavior based on data from the field to decrease the number of false detections. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2018)
3968 | 137 Error (brake slip) asking to reboot the device. | new | duanee | Brake / Release / 9V Battery |  | 2018-06-04 18:50 | 2018-05-03 |  | 100378 | Unassigned |  |  |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Undetermined | Bridging Bionics ‚Äì Occurred on 5/3/2018 in Americas Region (Device Serial Number: 100378) Issue ‚Ä¢ 137 Error (brake slip ‚Äì left knee) asking to reboot the device. Will not harm a patient. Background ‚Ä¢ Occurred during initial training- training was not cancelled. There have been 2 OOB failures. Strong Points ‚Ä¢ There is an inspection built into the MPI to ensure proper engagement after the brake is installed. | (Explain actions taken to resolve this issue) |  | The analysis of brake slip errors in the field resulted in the discovery that the software routine that checks for brake slip (a) checks for slip in both the extension and flexion directions, (b) has been triggered frequently in extension despite the fact that this is not the direction of potential harm, and (c) the number of both background slip detections and those that are presented to the customer across the field has caused more inconvenience to the customer than is warranted. This is supported by the lack of catastrophic failures in the field or safety issues (falls, etc.) caused by brake failure. As a result, the software behavior has been modified (2.1.4) to limit the brake slip detection behavior based on data from the field to decrease the number of false detections. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2018)
3967 | Rotating Leg (4065) | closed | duanee | Lower Leg | fixed | 2018-05-11 17:50 | 2018-03-20 | 126154 | EKSO-0196-00000000000021502005 | Unassigned | Shantel Firpi | SFirpi@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | Shantel reported: Hello I am here at CHOA for an advanced training. They reported that the lower leg is rotating on the right, similar to a previous issue that they had on the left. The right side rotated 3 times during 1 session, and past the first line, as the leg fell into external rotation. They have already tried new screws. Additionally, the sounds on this machine are significantly more quiet than a new Ekso. Step count 126,154 If there is any trouble shooting I can do tonight, please respond ASAP so I can work on it before I leave. I am here until Wednesday afternoon. - Shantel doubled doubled up the washers, which didn't help. She tested both tools and likely one of them is bad. - We will be shipping a new set of tools. | Tools and screws were sent per SRO-3401. | Customer tool is not holding the torque required. | The issue here (rotation of the leg during walk) has been reviewed and the root cause was determined to be that the adjustment screw(s) bottomed out and could not, as a result, provide adequate clamping force to hold as expected. This has been addressed and the shank collar on this unit has been replaced with the updated design. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3966 | #266, Knee brake not disengaging (4064) | assigned | duanee | Brake / Release / 9V Battery | fixed | 2018-05-11 17:41 | 2018-03-20 | 447522 | EKSO-0266-00000000000021601005 | OrthoMedico NV - Belgium | Sebastien Vanderlinden | svdl@orthomedico.be | +32(0)54515451 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Materials | Brake switch is not disengaging the brake even after replacing the batteries in the knee housing | Knee brake switch cable and housing replaced per SRO-3400. | The solder joint between the switch and the cable became disconnected, causing an open in the circuit. | The cables on the brake switch (B100637-009) have been known to be a weak connection at the solder joint that is easily damaged.
3965 | #234, issue with hip tube (4062) | assigned | duanee | Adjustments - Hip | fixed | 2018-05-03 13:59 | 2018-03-20 | 421978 | EKSO-0234-00000000000021509001 | Unassigned | Tania Romeu | neurologia@centrosinapse.com | 34981175859 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Customer Use | Unable to adjust hip width from fully extended. | hip width tubes and latch pin replaced per SRO-3396. | Damage to hip width latch pin (102876) causing the pin to bind on the hip width tube. | Damage occurred due to a very large patient being placed in the device without locking out the hip width settings.
3964 | #243, issue with left leg irregular step (4071) | new | jwitherspoon | Pick an option |  | 2018-05-03 13:39 |  |  | EKSO-0243-00000000000021510006 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | 34932231663 |  | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor |  | Undetermined | e-mail by Joan Romero dd. 19.03.18 Sorry about the delay in my answer. we started having the same problem again when doing steps. I just wanted to wait to be sure it was happening with more than one client. I'm attaching a video recorded this morning just to show you the problem. As you will see, suddenly the unit starts doing irregular steps, especially (but not only)with the left leg. When It happens, there is an increasing risk of falling forward. We tell the client to run the loaner unit so we could see if we can reproduce the problem. No problem when using the #109. customer already exchanged ankle assemblies and foot sensors!! | (Explain actions taken to resolve this issue) | (Required for complaints)
3963 | #128, hips cracked (I-4061) | closed | duanee | Pick an option | fixed | 2018-05-03 13:30 | 2018-03-16 | 509837 | EKSO-0128-00000000000021404003 | Unassigned | Dennis Veit | dennis.veit@youwalk.de | +49 15150463507 |  | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Dennis reported cracks in the hip housings of his device. | replaced hip housings with parts via SRO-3994. | Stress concentration at crack origin. Parts have been redesigned. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) #832 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/16/2017)
3962 | thigh structure failure (I-4133) | assigned | duanee | Adjustments - Leg |  | 2018-04-23 18:24 | 2018-04-23 | 1845641 | EKSO-0233-021508004 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 (0) 479 62 99 45 | Belgium | TBD | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Major | Investigation Needed | Undetermined | We have a major problem with one of our devices during walking with a patient. The patient is quiet experienced with Ekso walking (experience of more than 3 years ‚Äì SCI AIS A Th6). During walking the right leg component broke off on level of the thigh (metal fatigue?) Our patient was in a very unconfortable situation. We checked him clinically ad didn‚Äôt notice an immediate problem. But we will also evaluate his situation next days. I will WeTransfer the pics. Details : No. : 21508004 Total Step Count : 1845641 Kind regards, Kris Goos Manager REVAlution Clinical Training Specialist Ekso Bionics REVAlution center ‚Äì To Walk Again | Knee housings replaced per SRO-3480. | (Required for complaints)
3961 | Sensor Issue (I-4056) | closed | duanee | Foot Sensor | fixed | 2018-03-14 18:28 | 2018-03-14 | 533227 | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com | 703-408-7713 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Determined Unknown | Valerie called ater a FSV and reported the below issue: Valerie called me again this morning and her Ekso is stepping in place again, it‚Äôs not advancing the step, it goes straight up and straight down. This happened with Shantel during training and once before Shantel arrived. We‚Äôve done a manual cal to fix this the previous 2 times but it continues to happen‚Ä¶ The only obvious difference is the new software and the black sensors, this unit was one of our most reliable in the fleet. Unit #240, the last session this morning is when the issue began again. Can we look into the data to see if their sensors are causing this issue? I‚Äôm having Valerie run another cal but right now it seems like a band aid. Follow up from Katie - Is this true for both legs? Just clarifying- this _only_ happened in the last session of the day? Did it happen the whole session? Ok, when it‚Äôs having issues, it looks like the R toe is never coming fully off (this may be true for the L toe too, but so far I‚Äôve seen it on the R side). Possibilities: 1. Bad calibration 2. Bad sensor construction (something is allowing it to sometimes get stuck together and not release when let go) An auto cal should work if the sensor is behaving at the time (ie fully releasing). If it‚Äôs already compressed, this will cause more problems. Figuring if it‚Äôs ok is going to be hard, but I guess you could tell if it‚Äôs stepping normally during walking. However, since this behavior seems to appear later in the walk/day/etc, I‚Äôm wondering if it‚Äôs the sensor itself somehow getting kinda stuck. ** These sensors will be sent back to HQ | Josh walked Valerie through a manual foot cal on the phone. | Possibility of software issue. 2.1.2 has been know to cause issues with step after an upgrade is performed. Adam Zoss is working to determine the root cause of this issue. | previous investigation performed - INV-0006
3960 | Broken Foot Screw - Inc 4055 | new | duanee | Pick an option |  | 2018-03-14 18:23 | 2018-03-13 |  | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org | 509.665.6156 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor |  | Undetermined | ATTACHMENT: <C:\Users\lmcneil\Desktop\IMG_6349.JPG> Today during a treatment, while performing a turn, the screws holding the foot plate to the leg broke in half. The foot is no longer connected and the ends of the screws are stuck inside the holes of the leg. I am going to send you a picture just to be clear. If you have any questions about how it happened Enrique was the aide helping and can maybe explain further. However, we still have patients to treat. Is there any possibility of someone coming and pulling the screws out or do we need to send the unit back and get a loaner? Our acute rehab facility has a patient at 8:30 a.m. tomorrow and then we have a treatment at 1. If there is a possibility of getting it fixed instead of cancelling these appointments that would be great! Thanks for working with us as fast as you can, let me know if there is anything I can do to help!
3959 | Issues with controller (I-4044) | closed | duanee | UI - Controller | fixed | 2018-03-09 16:53 | 2018-03-01 | 89681 | 281 | Franciscan Health Lafayette | Corry Voll | corry.voll@franciscanalliance.org | (765) 423-6011 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Wear | Matt Knowlton at St Elizabeth's had the below issue during his ITW: I am at Franciscan St. Elizabeth‚Äôs Hospital in Lafayette, IN. On 3 occasions during gait training, the therapists have attempted to enter turn mode using the green check button and it did not register. During one of those incidents the patient triggered an unexpected step as the therapists believed they had entered turn mode. Upon inspecting the controller, all of the buttons appear somewhat loose and the green button in particular does not seem to register consistently. I feel this could present a safety issue in the future. I don‚Äôt believe it presents an emergent safety issue for today. However, I believe it should be replaced soon to prevent any mishaps. Please advise. | Mailed Part per SRO-3425. | Wear to the buttons on the controller causing presses to not respond. | The degree to which this issue could be investigated was limited by the inability of reproducing the reported behavior (see Root Cause Analysis). The complaint database was searched for similar issues, and two complaints with symptoms that may be related were found and reviewed (3342 and 3175). The current risk assessments for 1.2 FMEA-758 and 1.2 FMEA-762 (102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet) were reviewed, and it was determined that no update was necessary based on recent field data. No further action is necessary at this time. (D. Edmonds ‚Äì 06 Nov 2017)
3957 | Ankle sliders (I-4032) | closed | duanee | Ankle | fixed | 2018-02-23 16:19 | 2018-02-20 | 11437 | 331 | TLM - Mexico | Mike Glover | mglover@eksobionics.com |  | Mexico | No injury | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor | Investigation Needed | Determined Unknown | Hey guys- Something seems to be internal in the ankle box/resting angle bolt. When loaded with a patient standing in the device, I cannot rotate from -2 to -3 regardless of standing position moving through relative Dorsi/plantarflexion in Sagittal-A/P plane. I sat the patient and attempted to rotate and felt the same block with similar ‚Äòbreak‚Äô sound to continue to rotate around and was successful. With return to stand, I am blocked again. | Mailed replacement parts to distributor for replacement per SRO-3338. | Unable to recover part for investigation. | Issue with ankle slider seizing. Unable to recover part. Will further investigate if issue occurs again.
3956 | Crutch Interface not working - Inc 4031 | closed | duanee | UI - Crutch Interface | fixed | 2018-02-23 16:14 | 2018-02-20 |  | 331 | TLM - Mexico |  |  |  | Mexico | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Determined Unknown | Mike reported at the ITW: I went to train the team at CRIT/Teleton Guerrero (Acapulco) on the Crutch interface and it is not working properly. As soon as it‚Äôs plugged in, it gives the warning signal and message ‚ÄòCrutch interface disabled‚Äô and not registering this device. Some malfunction in this cable. I will bring it home with me. In the meantime, we will need to get them a new one before their AT | Sent a new controller in it's place. (SRO-3330) | (see Investigation Summary) | This issue with the crutch interface (warning signal and error message) does not present a safety hazard (see Risk Level Explanation). The Built-In Test (BIT) system properly detects the condition and prevents the use of the device. The customer has been sent a replacement crutch interface. Investigating this is, as a result, a low priority. If, however, this issue becomes more of an issue across the fleet (higher frequency of occurrence), a more in depth general investigation will be conducted. No further action is necessary at this point. (D. Edmonds ‚Äì 19 April 2018)
3954 | AMP Board Issue - Inc 4024 | closed | dbetteridge | PCBA | fixed | 2018-02-23 15:54 | 2018-02-08 |  | 308 | University of Montreal - Canada | Martin Vermette | martin.vermette@sympatico.ca | +1 (514) 340-2111 | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Martin Vermette, the main PT at University of Montreal reached out w/ concerns of a buzzing noise in the left knee. See below from customer: The right knee motor of our Ekso is making a peculiar sound this morning and we can not put that leg in Free Leg mode. We can feel the motor vibrate when we put our hand on it, which is not the case for the other motors. I took two small videos that I will share with you from my dropbox. https://www.dropbox.com/home/Ekso%20-%20RT%20Knee%20Motor%20-%202018-02-08 | In order to resolve the buzzing break issue, a new amp board was installed into the unit. | (see RCA for #3780) | previous investigation performed (#3780) No further action is necessary. (D. Edmonds ‚Äì 19 April 2018)
3953 | AMP Board - Inc 4021 | closed | dbetteridge | PCBA | fixed | 2018-02-22 19:25 | 2018-02-06 |  | 304 | Doctors Hospital at Renaissance | Nathan Dunn | nathand22@aol.com | 956-661-7100 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Nathan Dunn, a PT at Renaissance reported the below issue: The problem with Ekso SN: 021610003 is as follows: No error code appeared but the left knee joint will become difficult to passively move in legs free and during patient use it lacks a smooth arc of motion at the knee. A loud buzzing noise occurs at the left knee motor. We have tried changing the 9v battery for the left knee, but the problem persists. If you could help us resolve these issues promptly it would be greatly appreciated as we utilize both units daily. David-For the buzzing issue, this sounds like the amp board issue we had with 300 and 303. Is there any way to see if the breaker is tripping that powers the brake for this device? Eric-There is no way to check the status of the brake circuit fuse in software. We‚Äôd need to get at the amp board to confirm that this is the source of the problem. Dan-The problem has been that the inline resettable fuse that delivers 12V to the knee boards is intermittently opening up and closing thus powering on and off the break. With the 12V dropping out, I am a little surprised that the knee board computers do not reboot but the power outage transient might be so quick or not quite low enough for this happen. The fuses (2 of them to each knee joint) were changed from 0.5A rating to 1A rating on CO-11970. Granted, the AMP board on the unit mentioned is probably not of this newest revision, but I have found that a failure might even be due to poor solder workmanship of the fuse on the board. Recommendation is one of the following: 1) Reflow the solder joints on the fuses 2) Change the fuses to 1A fuses 3) Change the AMP board * We follow up and will be sending them a loaner device. | In order to resolve the buzzing break issue, a new amp board was installed into the unit. | (see RCA for #3780) | previous investigation performed (#3780) No further action is necessary. (D. Edmonds ‚Äì 19 April 2018)
3951 | Issues with stepping (I-4010) | closed | duanee | Foot Sensor | fixed | 2018-02-22 19:03 | 2018-01-31 | 336953 | 229 | TIRR | Marcie Kern | Marcie.Kern@memorialhermann.org | (713) 799-5051 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | Darrell reported the following during SA training: Doing the Smartassist upgrade training at TIRR and it looks like we are having a sensor issue. Both trajectory steps and free trajectory steps are restricted on the left. I did a sensor auto cal and I'd passed. We had a few good steps before it started acting weird again. This is affecting the training. | Foot sensors were mailed to the customer. The sensors were replaced and the device was recalibrated. | INV-0006 | INV-0006
3950 | Issues with stand to sit - Inc 4009 | new | duanee | Pick an option |  | 2018-02-22 18:55 | 2018-01-31 |  | 134 | Chuck Slyngstad | Nicole Hecock | threeforkspt@gmail.com | (406) 570-2867 | USA | No injury | Same as contact listed | Pick an option | FALSE | Unassigned | Minor |  | Undetermined | Chuck S's PT, Nicole reported the below issue: So thanks for the advice. We changed out the sensors and that works fine, BUT.... yesterday while walking, the machine had a "hitch" in what seemed like the hip off and on and then when we went to sit down, it stopped halfway down as if it was to the chair, but it wasn't. Ended up "freeing legs" and then he could get the rest of the way down. No error code or anything. What should I do? We tried sit to stands a couple more times and it happened each time. Please call me 406-539-0639
3949 | Standing Issue | new | ostewart | System |  | 2018-02-21 21:24 | 2017-11-25 |  | 275 | Good Shepherd - IP | Alyssa Hauck | Alyssa Hauck <AHauck@gsrh.org> |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Unassigned |  | Undetermined | Issue: Good Shepherd was using 275 as a loaner device, the unit would not stand when commanded w/ the crutch interface. | (Explain actions taken to resolve this issue) | (Required for complaints)
3945 | Worn thigh straps - Inc 4004 | closed | duanee | Soft Goods | fixed | 2018-02-13 19:09 | 2017-12-07 |  | 276 | AUVA - Weisser Hof, Austria | Marianne Worisch | marianne.worisch@auva.at | 4308063 | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | Customer reached out asking for a replacement set of thigh straps due to normal wear and tear. | Mailed part | INV-0009 | previous investigation completed ‚Äì INV-0009
3944 | Noisy Foot Sensors - Inc 4001 | closed | duanee | Foot Sensor | fixed | 2018-02-13 19:00 | 2017-10-24 |  | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org | 509.665.6156 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Customer reported swing leg block when in 2free. | Customer was shipped foot sensors. | INV-0006 | previous investigation performed - INV-0006
3943 | Foot Sensors worn - Inc 3998 | closed | duanee | Foot Sensor | fixed | 2018-02-13 18:45 | 2017-10-17 |  | 245 | Ostracon Ltd - Greece | Dimitris Aslanidis | physio@animus.com.gr | +30 2410 851851 | Greece | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Customer reported tearing on the foot sensors. | Customer was shipped new sensors. | INV-0006 | previous investigation performed - INV-0006
3942 | Replace Hip Abductors (I-3996) | closed | duanee | Adjustments - Hip | fixed | 2018-02-13 16:00 | 2017-10-10 | 580282 | 165 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Design | Customer reported hip abduction settings changing during walk. | Field Service team replaced hip abductors per SRO-3021. | Occurs due to the thread pitch on the lead screw assembly. | The hip abduction/adduction mechanism components were replaced (see Resolution Description) and testing showed that the issue reported by the customer was no longer present. The complaint database was reviewed for similar field issues, and the corresponding risk analysis (1.2 FMEA-393-C) was reviewed and determined to be consistent with what has been observed in the field. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3941 | Foot Binding and Foot Sensors - Inc 3990 | closed | kmettler | Foot Sensor | fixed | 2018-02-13 15:45 | 2017-10-06 |  | 294 | Santa Clara | Tawni Ybarra | Tawni.Ybarra@hhs.sccgov.org | (408) 885-5000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Customer reported a swing leg block while in 2free. | Foot Sensors replaced and calibrated. | INV-0006 | previous investigation performed - INV-0006
3938 | Wear to shoulder strap webbing - Inc 3986 | closed | duanee | Soft Goods - Shoulder Straps | fixed | 2018-02-13 15:12 | 2018-01-28 |  | 232 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | Customer reported wear to the top shoulder strap webbing. | Mailed Part | INV-0009 | previous investigation completed ‚Äì INV-0009
3937 | Thigh straps and closures - Inc 3985 | closed | duanee | Soft Goods | fixed | 2018-02-09 19:55 | 2017-08-07 |  | 206 | Rolling Ball - Ireland | Kim Gregg | millsygregg@icloud.com | +44 7740 839736 | United Kingdom | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | Customer reported the velcro on the thigh straps and side closures was not as effective. | Mailed part | INV-0009 | previous investigation performed - INV-0009
3933 | Error #168 - Incident 3956 | closed | duanee | Foot Sensor | fixed | 2018-02-09 18:48 | 2018-01-16 |  | 134 | Chuck Slyngstad | Nicole Hecock | threeforkspt@gmail.com | (406) 570-2867 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Chuck's PT, Nicole reported the below issue: Chuck and I are repeatedly getting error #168 during our walks, three times today. Any suggestions? Typically happens during right swing. Could it have something to do with duct tape on the feet? | Replaced Foot Sensors | INV-0006 * We will not be getting the sensors back for proper RCA. | previous investigation performed - INV-0006
3922 | Issues with device walking | new | lkealoha | Pick an option |  | 2018-01-22 18:57 | 2018-01-22 |  | 175 | Barrow Neurological Institute | Nate West | Nathan.West@DignityHealth.org | 801-472-7650 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Unassigned |  | Undetermined | Issue: The PT's Barb and Nate had been complaining about several issues over the past year. One of their devices is still having the ‚Äúcatch‚Äù while walking in 2free with various patients. A new set of foot sensors did not help the issue so the device was sent back to HQ for 'PM' | (Explain actions taken to resolve this issue) | (Required for complaints)
3917 | Incident 3942 - Worn Strap Failure | closed | duanee | Soft Goods | fixed | 2018-01-10 19:26 | 2018-01-09 |  | 274 | Good Shepherd - OP | Susan Holt | ssolt@gsrh.org | 610-776-3100 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Christina Wood at GS reported the below issue: I was using our ekso and the upper leg strap keeps coming undone due to being worn down. today it came undone and a patient hips dropped down from the device and caused some irritation. I have some safety concerns with sit to/from stand transfers due to strap not sticking. Is there a way to send new upper leg straps for the ekso? | Mailed Part | INV-0009 | previous investigation performed - INV-0009
3916 | Incident 3941 - Crutch Interface | closed | duanee | Pick an option | fixed | 2018-01-10 19:21 | 2017-10-19 | 33862 | 311 | Totally Kids | Pauline Hennessey | pauline@totallykids.com | 909-796-6915 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Materials | Mike at AT reported: Heads up, the top Velcro piece of the CI stitching is coming apart after a couple uses during the advanced training. | Mailed Part per SRO-3045. | Issue with the connection of the crutch interface to the torso shroud. | This issue (crutch interface failure) was investigated and the connection between the crutch interface and the torso was found to be faulty. The complaint database was reviewed for similar issues and the previous risk assessment was reviewed based on recent field data and found to be consistent with the probability reflected in rev. W of 102437. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3915 | Incident 3940 - Issues with foot sensors | closed | duanee | Pick an option | fixed | 2018-01-10 17:29 | 2017-10-19 |  | 228 | Sister Kenny | Rachel Vos | rachel.vohs@allina.com | 612-863-6095 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Josh reported via a PT at CK the below issue: I talked to Courage Kenny today and they are still having some issue going from stand to sit. The Ekso is finishing the sit around 65 degrees and they have to roll the patient back into the chair. I know we think this is a foot sensor issue, but they aren‚Äôt having any issues while walking‚Ä¶ Are we able to look at their data and see if it may be something else? The pt identifier for 2 of the pts are 0055 and 0064, the latter has had it happen multiple times and both had it happen the last time they used the device. Unit #228. Appreciate any insight you can provide. | Mailed Part | INV-0006 | previous investigation performed ‚Äì INV-0006
3911 | Incident 3930 - Foot Binding Jammed | closed | duanee | Foot Assy | fixed | 2018-01-09 16:32 | 2017-11-14 |  | 136 | Kessler | Raza Hussain | RHusain@kesslerfoundation.org | 732-794-1225 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | N/A - Identified known wear item | Wear | Karen Nollan reached out about the below issue: The strap that is suppose to adjust on the foot binding is jammed and it is difficult to adjust the foot buckle from patient to patient. Also, the strap is sticking out of the foot buckle on the medial side therefore hitting the other foot with every step. | The parts were removed and replaced by Raza, in biomed.
3906 | Customer reported resistance in the RK motor during walk | closed | duanee | Motor | fixed | 2018-01-08 17:58 | 2017-12-21 | 149745 | 231 | University of Calgary - Canada | Kyle McIntosh | kyle.mcintosh@ahs.ca | 403-404-1660 | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Workmanship | During walk, towards the end of the session, customer reported increased resistance in the RK motor | Right knee motor replaced per SRO 3198 | Small crack on rotor magnet during installation which propagated and broke off a piece of the magnet | The core issue leading to the reported behavior has been isolated to a defect in the right knee motor on the device. The defect (damaged magnet in the motor) resulted in slightly different motor torque constant and inconsistent resistance throughout the motor rotation. The complaint database was reviewed, and this appears to be an isolated failure at this point. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3904 | Incident 3921 - Battery Charger Issues | closed | duanee | Battery - Charger | fixed | 2018-01-05 19:43 | 2017-12-07 |  | 59 | Able Bionics - Aspen | Amanda Boxtel | amanda@bridgingbionics.org | (970) 379-9260 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | Amanda Boxtel reached out saying: Our batteries are showing mismatched and resulted in us not being able to get a client up today to walk. He was really disappointed. Can we make it a rush order please so we aren‚Äôt compromising our program? | Parts mailed and received. | We will not be getting the charger back, so we cannot perform a proper RCA. | There are no safety concerns associated with this failure. The charger has been replaced, and we have been able to confirm that this resolved the issue. No further action is necessary. (D. Edmonds ‚Äì 17 May 2018)
3902 | Incident 3918 - Issues with Foot Sensors | closed | duanee | Pick an option | fixed | 2018-01-05 19:25 | 2017-12-19 |  | 271 | Reeves Rehabilitation | Samantha Zimmerman |  | (210) 358-1056 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Gabbie Canales reached out stating they were having issues in 2FRee with the device and some of their patients not being able to complete a step. We send them instructions for a re-cal for the Ekso and it didn't pass. | Parts Received | INV-0006 | previous investigation performed ‚Äì INV-0006
3898 | Incident 3913 - Clip Crack on Backpack Strap | closed | duanee | Soft Goods | fixed | 2018-01-05 17:02 | 2018-01-02 |  | 289 | Dominican Hospital | Andrea Martin | Andrea.Martin@DignityHealth.org | (831) 462-7385 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Determined Unknown | Andrea Martin reported: 1. Right shoulder clip crack- this crack appears to be superficial but would hate for it to give at the wrong moment. | Mailed parts and parts received. | It is not possible to determine root cause for this issue because the part was not returned. | previous investigation completed ‚Äì INV-0009
3897 | Incident 2 - Hip adjustment loosening during walk | closed | duanee | Pick an option | fixed | 2018-01-05 16:29 | 2017-10-04 |  | 15 | Magee | Elizabeth Watson | Ewatson@mageerehab.org | (215) 218-3900 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Hex bolt that locks the hip width in place on the right side is loosening during walk and falling out. Unit has the upgraded nut with the helicoil per CO-10949. | New adjustment bolt shipped to customer per SRO-2. | Wear to the threads of the bolt were allowing the bolt to loosen during walk. | Material wear and infrequent event.
3895 | Incident 3872 - Foot Sensors ripped | closed | duanee | Foot Sensor | fixed | 2018-01-05 16:13 | 2017-10-09 |  | 142 | Burke | Kristen Hohl | khohl@ricres.org | (312) 238-1000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Block in 2 free | Mailed part | INV-0006 | previous investigation performed ‚Äì INV-0006
3894 | Incident 3873 - Broken Foot Sensors | closed | duanee | Foot Sensor | fixed | 2018-01-05 15:59 | 2017-10-09 |  | 106 | RIC | Kristen Hohl | khohl@ricres.org | (312) 238-1000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Customer reported issues while walking, with sensor error codes. | Mailed Part | INV-0006 | previous investigation performed ‚Äì INV-0006
3893 | Incident 3874 - Foot Sensors broke | closed | duanee | Foot Sensor | fixed | 2018-01-05 15:48 | 2017-10-09 |  | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Design | Foot Sensors broke | Mailed Part | INV-0006 | previous investigation performed ‚Äì INV-0006
3892 | Minimum engagement stitch missing | closed | duanee | Soft Goods | fixed | 2018-01-05 15:37 | 2017-11-09 |  | 339 | Henry Ford Macomb Hospital |  |  |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Materials | Red stitch line to indicate minimum required engagement on thigh strap missing in first batch of parts from Agora. (3884) | Mailed Part | The supplier did not use the correct thread for one section of the thigh strap, and the defect was not caught in receiving/incoming inspection, on the production line, or during the walk test. | The thread used for the stitching on a number of thigh straps from the supplier was not correct. There is one area where red stitching is specified, however, on these parts the supplier did not switch to red thread. This is a relatively new supplier for these parts, and the supplier has been notified. There is no safety issue associated with this defect (see Risk Level Explanation). The customer has been sent thigh straps with the correct thread color. (D. Edmonds ‚Äì 29 Mar 2018)
3891 | Incident 3891 - Noise coming from right hip | closed | duanee | Actuators | fixed | 2018-01-05 15:25 | 2017-11-15 | 261060 | 274 | Good Shepherd - OP | Susan Holt | ssolt@gsrh.org | 610-776-3100 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Wear | Screeching sound from right hip. David 1/29/18 - Customer reported screeching sound, but the sound was a click. Screeching often indicates a broken hip housing, but no signs were present on this device. | Bearing was replaced per SRO-3087. | The ball bearing in the hip housing had worn to the point where there was some play between the inner and outer races of the bearing. The bearing did not appear to have sufficient amount of grease packed between the balls. | Clicking occurs at low frequency and does not affect use of the device.
3890 | Hip Abduction Slipping | closed | duanee | Adjustments - Hip | fixed | 2017-12-21 15:55 | 2017-10-01 | 18949 | 338 | Immanuel Rehab | Amanda Engler | amanda.engler@alegent.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Design | Hip abduction changing settings during walk. | Hip abductors replaced per SWO-03022. | Occurs due to the thread pitch on the lead screw assembly. | The hip abduction/adduction mechanism components were replaced (see Resolution Description) and testing showed that the issue reported by the customer was no longer present. The complaint database was reviewed for similar field issues, and the corresponding risk analysis (1.2 FMEA-393-C) was reviewed and determined to be consistent with what has been observed in the field. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3889 | Phone not Transmitting Data | closed | duanee | Pulse Data | fixed | 2017-12-21 15:07 | 2017-08-01 | 469204 | 22 | Spaulding | Kate Bennett | kebennett@partners.org | 617-573-7000 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor | Investigation Needed | Wear | Unit 14 has been changed out with unit 22 for Spaulding. Unit 22 was not transmitting data when it arrived at customer site. | phone replaced per SWO-02904. | Phone battery not holding a charge. | Battery in phone was not holding a charge.
3888 | Continuous error 27 | resolved | dbetteridge | System | fixed | 2017-12-06 20:11 | 2017-11-21 |  | 228 | Sister Kenny | Krista Henshue | Kristina.Henshue@allina.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Unassigned |  | Undetermined | Issue: We had another safe mode here at Courage Kenny at Abbott. These have been happening almost every week recently. The current code was #27. The patient is going to continue walking as it says ‚Äúreboot and contact customer relations if it occurs again‚Äù. Any suggestions on this error code? | (Explain actions taken to resolve this issue) | (Required for complaints)
3887 | Continuous Joint Cal required | new | dbetteridge | System |  | 2017-12-06 20:07 | 2017-12-06 |  | 319 | Good Shepherd - OP | Christina Wood |  | 302-598-5628 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Unassigned |  | Undetermined | Issue: GS is getting joint cal errors on their right knee. It happens during setup and said that they had three patients in it yesterday afternoon. It looks like we haven‚Äôt gotten data since 11/17, but they left the device on in case you need to call data. Phone is: 806c1bf268c3. | (Explain actions taken to resolve this issue) | (Required for complaints)
3884 | Rotating Leg | closed | duanee | Adjustments - Leg | fixed | 2017-12-05 23:54 | 2017-11-17 | 60600 | 196 | Children's Hospital of Atlanta | Darrell Musick | dmusick@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | Investigation Needed | Materials | Issue: I am at CHOA doing Smartassist training and have a few small items to report. The left lower leg is rotating during walking/turning. This is a known issue and worse with our new sensors on their highly tacky linoleum floor. They have replaced the screws a few times already. It is happening every session. | Upper shank replaced per SRO-3187. | Upper shank not holding torque due to the angle of the bores where the screws sit. | The issue here (rotation of the leg during walk) has been reviewed and the root cause was determined to be that the adjustment screw(s) bottomed out and could not, as a result, provide adequate clamping force to hold as expected. This has been addressed and the shank collar on this unit has been replaced with the updated design. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3880 | Issues with R leg | closed | duanee | Foot Sensor | fixed | 2017-11-20 21:17 | 2017-11-19 | 79933 | 190 | Mt. Sinai | Nick Manente | Nicholas.manente@mountsinai.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | Issue: Hello I received a text from the PTs at Sinai this weekend about some issues they are having with their GT. I have attached the video to this email . The report issues with advancement of the right leg, no matter if the leg is under a trajectory or if it‚Äôs free. When in left affected, and stepping with the right leg, the remote will say waiting for right step even though the step is completed and is now the weight bearing leg. They say the patient is shifting equally, and this is happening with multiple patients. Please follow up with Nick (Nicholas.manente@mountsinai.org) And William (William.weinrauch@mountsinai.org) Thank you and hope you all had a nice weekend, Shantel | Josh walked the customer through a foot calibration. | Wear to electrode in the foot sensor. | Previous investigation completed (INV-0006)
3879 | Screeching from right hip. | closed | duanee | Pick an option | fixed | 2017-11-15 19:04 | 2017-10-18 | 260557 | 274 | Good Shepherd - OP | Susan Holt | ssolt@gsrh.org | 610-776-3100 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Wear | Loud screeching sound coming from right hip during walk. | All excessive debris was cleaned per SRO-3087. (This is part of the 100-point inspection to clean the lower DDL tubes and the inside of the knee housing.) | Material build-up on the don-doff link and wear in the bearings in the hip. | This was determined to be a non-performance issue. The accumulation of material in the device drive train caused abnormal sounds to be generated by the device, causing concern. The device was serviced to remove the material build-up and was tested to confirm that this addressed the sound issue. No further action is necessary. (D. Edmonds ‚Äì 21 Dec 2017)
3877 | Continuous Joint Cal required | closed | dbetteridge | Pick an option | fixed | 2017-11-06 23:18 | 2017-11-01 | 36811 | 121 | Baylor | Chad Swank |  | (214) 820-9393 |  | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Issue: Joint calibration continues to pop up for the PT's. | Soft pot was replaced per SRO-3196. | Soft pot de-laminated from the hip housing. | The core issue was isolated to damage to one of the poteniometers used for joint position feedback. The component was replaced, and the device was tested successfully to verify normal operation. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3875 | Trinity Sports Centre | closed | duanee | Pick an option | worksforme | 2017-10-26 10:39 | 2017-09-29 |  | 192 | Unassigned | Ronan Langan |  |  | Ireland | ankle fracture | Same as contact listed | ekso 1.2 | TRUE | 0 - Not Required | Minor | Investigation Needed | Customer Training | Patient requested to participate in a clinical trial. He is a C6 complete spinal injury and had not stood for 17 years. Patient had reduced bone density on Dexa scans (-2.3) but was happy to take the risk. Without any specific cutoffs for acceptable bone mineral loss for walking or standing the clinic decided the best approach for determining eligibility for the study was to manage it in stages. The patient was taken into the rehab clinic to trial verticalisation on a tilt table to avoid a fainting incident in the exoskeleton and he tolerated it for about 5 minutes (without binder) so the Physical Therapist decided this was good enough as he would have abdominal support and be walking in the exoskeleton and this indeed proved not to be an issue when they walked him the first time. The Physical Therapists were conservative for first walk and the patient walked 200 steps without incident. The PT‚Äôs had some difficulty aligning the patient so the walk was tough and felt it would improve so they checked his skin and observed no other adverse events. The PT‚Äôs decided to try walking the patient again in a week. On the second walk the patient completed 400 steps but reported ankle swelling and feeling dysreflexic. The PT‚Äôs reviewed his foot the next day and identified swelling - advised him to ice but it did not resolve the swelling. After a week he went for x-ray and fracture was identified. The patient was placed in a boot for 6 weeks and then will be back to normal duty. The patient is fine and going to start treatment for bone density. Date of Incident: Sep 29, 2017 Date Became Aware: Oct 3, 2017 Patient Name: Phillip Patient Age: Unknown | n/a | (see Investigation Summary) | The device was evaluated and found to be operating within specifications. Walking was smooth with no unexpected behaviors. A review of the lot history records identified no manufacturing nonconformities issued to the reported lot that would have contributed to this event. Additionally, a review of the complaint history did not indicate a lot specific quality issue. There is no indication of a product quality issue with respect to the design, manufacture or labeling of the device. The patient is outside the IFU inclusion criteria so the clinic was asked to eliminate him from the study. It is suspected that the combination of the patient not meeting the inclusion criteria, having reduced bone density, and being non-ambulatory for 17 years contributed to the injury. (D. Humphrey ‚Äì 16 October 2017)
3868 | Patient lowered to ground and removed from device | closed | duanee | Unknown | fixed | 2017-09-22 19:10 |  |  | 279 | Sunnyview Rehab | Patti Valenza | Patricia.BroschartValenza@sphp.com | 518-382-4528 | USA | Orthostatic Hypotension | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor | Investigation Needed | Customer Use | It was reported that the Ekso GT exoskeleton, serial number 021602008 (EKSO-0279-021602008), was used as part of a stroke rehabilitation regime for a 63 year old male patient. The patient was placed in the device and performed sit to stand with minimum assistance and verbal cues. While the patient was ambulatory for approximately 10-15 feet he became ashen and then unresponsive. The patient immediately sat down, the device was removed, and he was transferred to a mat. He was responsive and able to answer questions. Nurse practitioner Michele Cournan made the determination to send the patient to an acute care hospital where he expired 20 days later while in the ICU. In the opinion of Ms. Cournan, the use of the device did not cause or contribute to the patient death. No additional information was provided. | CAPA-00002-Report | CAPA-00002-Report | CAPA-00002-Report
3866 | Unit #260 - BG Hamburg - Ekso does not count steps in 2free anymore | closed | mblaschke | Pick an option | fixed | 2017-09-21 16:21 | 2017-09-01 |  | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | 040/73063670 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Customer Use | According to customer the unit was walked with a patient and after changing to 2free, steps were not counted any longer. | Phone calls to the customer and support from Eric and Katie helped the team figure out the root cause. | Customer was not pressing the button on the UI to enable walk. | Customer use issue.
3864 | Errors 1058, 133 and 121 | new | dbetteridge | Pick an option |  | 2017-09-19 21:22 | 2017-09-19 |  | 334 | Star Rehab | Joe Bongiovanni | Joe@STARrehab.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issue: I was going to email you this morning to let you know that we have it all packed up; however,wWe had staff working with the unit you just returned to us and it threw the same code as it did before we sent it back and it went into safe mode. Upon reviewing the service summary, I noticed that it was serviced for an issue with the batteries disconnecting. I was under the impression that the issue we were having corrected was a hardware problem for newer components resulting an error occurring and the device going into safe mode. Are these (battery and hardware) one and the same or was there another issue? What would you like use to do as the same issue has again occurred? | (Explain actions taken to resolve this issue) | (Required for complaints)
3863 | Unit #236 - EMAC, Italy - cca defect | closed | duanee | PCBA | fixed | 2017-09-19 13:09 | 2017-09-11 | 15498 | 236 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor | No investigation needed - see Justification | Materials | customer reported an issue with the knee | mailed parts, swo 2997 (102286) | This issue was isolated a failure with the thigh PCBA. | This is a known potentital issue with the signal chain for the brake. This does not present a significant safety issue. The distributor has been sent replacement parts. No further action is necessary. (D. Edmonds ‚Äì 19 April 2018)
3861 | Unit #270 - AUVA Bad H√§ring - batteries are not loading 100% | closed | mblaschke | Pick an option | fixed | 2017-09-19 08:28 | 2017-09-13 |  | 270 | AUVA - Bad Haring, Austria | Stephan Sefqa | stephan.sefqa@auva.at | +43 5 93 93 52317 | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | According to customer, the battery 2B ist not completely loading, 88% is the most. Even after testing the power cable, still the same issue. | New batteries sent directly to customer. | INV-0010 | previous investigation performed ‚Äì INV-0010
3858 | Unit #209 - Technomex - knee motor shows error121 | closed | duanee | Lower Leg | fixed | 2017-09-18 10:58 | 2017-09-14 | 380322 | 209 | Technomex - Poland | Krzysztof Cygon | kcygon@technomex.pl | +48 32 231 66 60 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | customer reported error message during session | replaced parts, swo 2990 | The error was narrowed to an issue with the knee motor assembly that caused the Elmo motor drive to enter a fault condition. | Service was performed on this unit and the issue was isolated to the knee motor assembly. Something about the motor caused the diagnostics on the Elmo motor drive to enter a fault condition, triggering the Error 121 from the Built-In Test (BIT) routine on the Ekso GT DSP. The system responded as designed to the error condition. The knee motor assembly was replaced, and the device was tested without issue. The complaint database was searched for similar issues. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3855 | Swing Leg Block | closed | duanee | Foot Sensor | fixed | 2017-09-13 03:30 | 2017-09-12 | 26627 | 313 | Loma Linda | Lisa Zidek | LZidek@llu.edu |  | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Materials | Free leg having hard block at 35 degrees consistently. From Mike: Seems like we have some hot toe sensors on the R foot in Free. The patient needed significant elevation /unloading to minimize the foot being 'stuck' on the ground. Video included. This is on their 'outpatient Ekso'. We're getting replacement foot sensors tomorrow, for a different reason and this issue is likely resolved. I tried an Auto Calibration and the sensors read less-high and the sticking seemed less, but still present. | Both foot sensors replaced per SWO-02983. | This was determined to be due to noise coming from the foot sensor that caused the unit to believe that it had not lifted the leg off the floor. | INV-0006
3854 | Broken Ankle Screw | closed | duanee | Ankle | fixed | 2017-09-12 22:07 | 2017-06-08 | 57945 | 190 | Mt. Sinai | Rose-Marie Faotto | tose-marie.faotto@mountsinai.org | 212-241-6500 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | A loud popping sound was reported from the ankle joint. The PT that walked the unit said "it felt like he stepped in a bubble that popped under his forefoot." | Ankle assemblies replaced per SWO-02771. | Vendor issue that prevented the pre-applied Loctite patch from curing. | A technical analysis was performed (summary included below), and it was determined that the threadlocker on the fasteners had not set (cured) correctly, allowing the fasteners to back out over time after they were correctly seated and tightened. 'Added my observations to the end of this document. None of the footplate screws failed or showed signs of cracking ‚Äì the front two fasteners had backed out. The ‚Äúbubble that popped‚Äù described by the PT was most likely the ankle box shoulder re-seating in the counterbore once the device was loaded. The screw threads showed heavy signs of polishing/wear from rubbing against the inside walls of the footplate thru holes. For all 4 fasteners, it does not appear that the pre-applied Loctite (ND Microspheres) had cured. It The threadlocker was still gummy and it did not require much torque for David to remove them. The major diameter of the screw threads and the minor diameter of the housing threads are all within spec. My theory is that the front two fasteners began backing out because the threadlocker had not set. Once this progressed enough, the ankle housing was able to separate from the footplate during each step, creating the loud ‚Äúpopping‚Äù noise reported by the PT below. There does appear to be damage to one of the internal threads where the screw backed out, but not as bad as if the fastener had pulled the threads completely out.' The complaint database was searched for similar issues, and it has been determined that this specific complaint is related to the broader investigation summarized in INV-0001. No further action is necessary. (D. Edmonds ‚Äì 07 Nov 2017)
3850 | Unit #197 - EMAC, Italy - moving pin | resolved | duanee | Pick an option | fixed | 2017-09-08 10:12 | 2017-09-07 | 949775 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Undetermined | customer reported a moving pin on the right hip. | replaced pin | (Required for complaints)
3849 | left knee cracked | assigned | mblaschke | Pick an option |  | 2017-09-07 14:35 | 2017-09-07 |  | 129 | Unassigned |  |  |  |  | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Undetermined | (Describe issue, including geographic location, if appropriate) | (Explain actions taken to resolve this issue)
3847 | Foot sensors coming apart | closed | bhanson | Foot Sensor | fixed | 2017-08-30 19:40 | 2017-08-30 |  | 59 | Able Bionics - Canada | Amanda Boxtel | Able Bionics <ablebionics@gmail.com> | (970) 379-9260/ (970)927-9319 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | Issue: We are in need of a new left sole replacement for our Ekso please. We have temporarily duct taped the sole in place. | Sent new sensors via SWO-2953. | Known wear item, will not be receiving these sensors back for proper RCA | The tearing of the sole does not present a significant performance issue with the use of the device. Similar to wear on the sole of shoes, to some degree, this is expected wear. The complaint database was reviewed for similar issues. The cap sole design was updated with CO-11648 in Feb 2017. (D. Edmonds ‚Äì 04 Oct 2017)
3846 | Unit #197 - EMAC/Istituto Montecatone - problems with left hip width | closed | duanee | Pick an option | fixed | 2017-08-30 10:24 | 2017-08-23 |  | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | The PT sent the attached pictures to show the minimum value reachable from the left hip width. | Matteo from EMAC solved this issue, the small tube on the torso was loose. He opened the back and tightened the screws on the clamps. | The issue with hip width adjustment was isolated to clamp screws that had loosened over time in the torso. | The issue with hip width adjustment was isolated to clamp screws that had loosened over time in the torso. The complaint database was reviewed for similar issues, and this was added to the list of reliability topics to be discussed with the Manufacturing Engineering and Field Service teams. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3843 | Unit # 153 - Step by Step, Barcelona - foot assembly broken | closed | mblaschke | Foot Assy | fixed | 2017-08-28 14:39 | 2017-08-28 | 326520 | 153 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | 34932231663 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Determined Unknown | The customer tried to use the loaner but when they stood up for the first time with a client they heared a weird noise coming from both feet. They tried to find out what happened and found that the same piece was broken on both feet (please watch video attached). | replaced parts, swo 2300 | Several components were replaced under SWO-2300 and Field Service confirmed that the issue has been resolved. It is not possible to further narrow down the root cause. | Several components were replaced under SWO-2300 and Field Service confirmed that the issue has been resolved. It is not possible to further narrow down the root cause. No further action is necessary. (D.Edmonds ‚Äì 23 Sep 2018)
3842 | Ankle and Foot Swelling | closed | duanee | Pick an option | fixed | 2017-08-28 13:44 | 2017-06-26 | 134808 | EKSO-0033-011208002 | WakeMed |  |  |  | USA | No injury | MAUDE Adverse Event Report No. 6740206 | ekso 1.1 | TRUE | Unassigned | Minor | Investigation Needed | Customer Use | MAUDE Adverse Event Report No. 6740206 {{{Patient w/ thoracic spinal cord injury & flaccid paraplegia. Following use of the ekso bionic gate device, the patient had swelling of the right ankle & foot. This was the patient's first time using the device. A trained physical therapist & pt assistant were present during the session and reported that the patient had a good fit and did well. No skin issues were noted at the end of session. The next day, the patient reported having spasms in her leg and noted some swelling. When she removed her shoe, she noticed that her r ankle looked bruised. Physical therapy was called and the patient was assessed. The patient was found to have swelling in r foot medial ankle & distal part of lower leg and bruising below the medial malleolus. A physician was notified and x-ray was completed. X-ray reported an "avulsion of unknown chronicity over the medial talus body and lateral malleolus. No obvious acute fractures. " casting was not needed, but the patient was treated with ice, compression, and an aso brace. The patient's husband had a video of her walking in the ekso- this was reviewed and noted that her r ankle was supinating inside the foot strap of the ekso (x372 steps taken during that session). The patient's ankle moved in a frontal plane rather than a a/p plane possibly causing stress to the lateral ankle. Ultimately, it is believed that the patient's dorsiflexion was overestimated and perhaps, additional necessary adjustments to the ekso were not made. It is staff's perception that the device did not cause this injury, but perhaps, lack of visualization of the maladjustment was the problem. Typically, 2 physical therapists are present/assist new spinal cord injury patients when using the ekso. In this reported case, the patient was not a new injury, and instead of 2 therapists being present, there was 1 physical therapist & 1 pt assistant. As of filing of this report, there are 134,808 steps of use on the device.}}} | No further action is necessary. | (see Investigation Summary) | The details of the MAUDE Adverse Event Report (6740206) were reviewed and discussed with the Clinical, Quality Assurance and Regulatory Affairs departments. The conclusion of the review agrees with the perception provided in the report (based, in part, on the review of video from the session) that the injury was not caused by the device (Ekso 1.1). The injury was caused by inaccurate measuring of the range of motion (ROM) of the patient's ankle and suboptimal positioning of the foot in the foot binding of the device. The report indicates that a post-session skin inspection was performed and indicates no issues. A post-session standing visual inspection may have helped with the ankle supination issues, although it is possible that the ankle may have moved slightly during the walking session. (CAPA-00003-Report, D. Edmonds ‚Äì 05 Oct 2017)
3841 | Errors 1241 and 21 | closed | duanee | Pick an option | fixed | 2017-08-25 18:11 | 2017-08-25 | 379352 | 263 | Rehab Institute of Kansas City | Natalie Brandt | natalie.brandt@rehabkc.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Issue: Customer: RIKC Unit #: 263 RIKC called this morning and reported multiple !#1241 and !#21 errors, with several patients. Additionally, they tried to enter in patient information and the controller was not saving their settings. They then tried to walk the device and said that it felt like it was ‚Äòmoving in slow-mo‚Äô (sluggish) and that the steps were hesitating. We will either be sending a loaner and bring their unit back to HQ or do a field service visit. | David, Katie, and Eric were unable to find the root cause of the issue so we decided to replace the entire signal chain from the amp board to its connections within the knee (includes the main control board, external knee cable, and all control boards in the knee). Original motor is still on the unit. SRO-2954 Patient information is now saving when prompted with the new amp board. | failure with the communication lines between the main control board and the motor | The issue was determined to be an isolated issue that resides in the communication lines between the main control board and the motor. Details of the timeline of the investigation of this issue are captured below. 8/25/17 - The device was reported to have multiple 1241, 21, and CAN errors on the right knee. See the TFF email attached for details. 9/3/17 - The device was placed on the squat rack and failed at about 86 minutes in with the following errors: 27, 1054, 1066, and 1240. This was with all original parts. 9/11/17 - The joint board was replaced on the device and placed back on the squat fixture. The unit threw 1054, 1066, 1240, and 27 at about 50 min into the test. 9/26/17 - An older amp board was installed on the device and the device was squatted without issue for about 3 hours without issue. 9/26/17-9/27/17 - The original amp board was reinstalled and run through about 3 hours without issue. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3836 | Foot sensor issue | closed | duanee | Foot Sensor | fixed | 2017-08-23 21:47 | 2017-06-27 |  | 130 | Sunnyview Rehab | Patti Valenza | Patricia.BroschartValenza@sphp.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | Issue: PT reported that the R Knee was getting caught, interfering w/ patient sessions. These issues were coming up on the loaner device (130) | Foot sensors sent via SWO-02824. | The sensors that were replaced on SWO-02824 are not available for further analysis. | INV-0006
3835 | Foot sensor issue | closed | duanee | Foot Sensor | fixed | 2017-08-23 21:24 | 2017-07-05 |  | 196 | Children's Hospital of Atlanta | Erin Eggebrecht | erin.eggebrecht@choa.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | Issue: PT reported that the R Knee was getting caught, interfering w/ patient sessions | Foot sensors were sent. | The parts that were replaced in the field are not available for further analysis. | INV-0006
3831 | Charger not working | closed | duanee | Battery - Charger | fixed | 2017-08-23 20:50 | 2017-03-02 | Not applicable | 136 | Kessler | Raza | RHusain@kesslerfoundation.org | 732-794-1225 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: The other item I had was that one of our chargers does not charge batteries past 88%. Once it reaches that threshold, it starts to flicker as if it cannot decide whether or not the battery has enough charge. I have tested each of the batteries that we have, and this charger consistently gives us the same issue. Is there a way for me to send it in for a repair/replacement? | Sent parts via SWO-2522 | Known wear item, not receiving the charger back for proper RCA | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3830 | Controller Malfunction | closed | duanee | UI - Controller | fixed | 2017-08-23 20:48 | 2017-03-02 | 661804 | 136 | Kessler | Raza | RHusain@kesslerfoundation.org | 732-794-1225 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | Issue: FYI, the controller on our original device (the Investigational unit) malfunctioned upon standing with a participant this afternoon. I powered the device down and lowered the participant back onto the bench. When we turned the device back on, the controller worked properly and training proceeded without incident. This problem had also occurred back in March of 2015. The controller will blank out and then operate just fine after the reboot. | Part replaced via SWO-2522 | Known wear item/ design issue, not receiving the part back for proper RCA | INV-0008
3829 | Charger not charging | closed | duanee | Battery - Charger | fixed | 2017-08-23 20:27 | 2017-06-13 |  | 311 | Totally Kids | Pauline Hennessey | pauline@totallykids.com | 909-796-6915 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | Issue: PT's reported that one of their battery chargers is not charging | New battery charger sent via SWO-2808 | It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3827 | Charger not charging | closed | duanee | Battery - Charger | fixed | 2017-08-23 20:02 | 2017-06-13 | 174832 | 279 | Sunnyview Rehab | Patti Valenza | Patricia.BroschartValenza@sphp.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Determined Unknown | Issue: PT's reported that one of their battery chargers is not charging | Sent new charger via SWO-2788 | It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3826 | Battery charger issue | closed | duanee | Battery - Charger | fixed | 2017-08-23 19:47 | 2017-06-13 |  | 228 | Sister Kenny | Rachel Vos | rachel.vohs@allina.com | 612-863-6095 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: Hello all, I am reaching out to the people at Ekso regarding an issue with our battery charger and including our Manager and Supervisor in case a new one needs to be purchased. One of our battery chargers is not charging; is there an easy fix? Should we send it in? Do these need replacing on a certain scheduled basis? We have been just using the one charger but then the turn-around time is much slower for being able to have a set ready to go. Thanks for any help in this manner, | Sent parts via SWO-2789 | It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3823 | Batteries not charging | closed | duanee | Battery | fixed | 2017-08-23 19:16 | 2017-06-27 |  | 281 | Franciscan Health Lafayette | Corry Voll | corry.voll@franciscanalliance.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Determined Unknown | Issue: PT reported that a set of their batteries were not charging to a full percentage | New batteries sent. | Route cause is unknown as we cannot received 'damaged' batteries back due to the DG regulations | This set of batteries no longer charged to full capacity. To some degree, this is expected. It is not clear, however, if this set experienced a faster than normal decline in capacity based on how they were used. We do not have a great understanding of how they were used, and the batteries that were replced are not available for further analysis. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3822 | Foot sensor issues | closed | duanee | Foot Sensor | fixed | 2017-08-23 19:05 | 2017-08-23 |  | 175 | Barrow Neurological Institute | Nathan West | Nathan.West@DignityHealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: We are having problems with our first Ekso. The L knee is catching and at times not letting the knee to bend. I tried it out and it was definitely a knee issue and not a hip issue. It also seemed to be partly velocity dependent, but not always. I am not sure if there is any troubleshooting that can be done, or if we need to send it in again. | Parts were replaced via SWO-2861 | The parts that were replaced are not available for further analysis. | INV-0006
3820 | Faulty controller | closed | bhanson | UI - Controller | fixed | 2017-08-23 17:46 | 2017-06-27 |  | 161 | Glenrose - Canada | Kristin Tarnowski | Kristin.Tarnowski@albertahealthservices.ca |  | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Materials | Issue: Yesterday the screen scrambled when the physios were setting the machine up. This happened several times over a course of a few weeks. - The PT's reached out to get a new controller | Sent new parts via SWO-2823. | This issue was caused by a susceptibility of the LCD PCA used in the control pad assembly to electrostatic discharge (ESD), which at times causes the LCD display to scramble requiring a power cycle to correct. | The complaint database was searched for similar issues, and two complaints with symptoms that may be related were found and reviewed (3342 and 3175). The current risk assessments for 1.2 FMEA-758 and 1.2 FMEA-762 (102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet) were reviewed, and it was determined that no update was necessary based on recent field data. No further action is necessary at this time. (D. Edmonds ‚Äì 06 Nov 2017)
3819 | Charger not charging batteries | closed | duanee | Battery - Charger | fixed | 2017-08-23 17:22 | 2017-08-07 |  | 281 | Franciscan Health Lafayette | Corry Voll | corry.voll@franciscanalliance.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: Lily- I am thinking the charger needs replaced or looked at as none of our batteries including our new set are not holding a full charge and we are having to switch out batteries in the middle of treatments. Thanks for any help, Corry | Sent new parts via SWO-2903 | It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3815 | No Data Transfer From Phone | closed | duanee | Pulse Data | fixed | 2017-08-22 15:49 | 2017-08-10 | 188082 | 205 | Barrow Neurological Institute | Nathan West | nathan.west@dignityhealth.org | 602-406-3000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Wear | Customer not receiving any data in pulse for this unit. | Phone upgraded to Moto-E per SWO-02916. | Phone battery was completely dead at time of service. Would not take a charge. | This issue was narrowed to a failure of the cellular phone in the device. The phone would not boot and, as a result, could not transmit the device data to the cloud. The phone was replaced. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3814 | Swing Leg Block | closed | duanee | Foot Sensor | fixed | 2017-08-22 15:04 | 2017-08-10 | 341965 | 175 | Barrow Neurological Institute | Nathan West | nathan.west@dignityhealth.org | 602-406-3000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Mishandling | Free leg having hard block at 35 degrees consistently. | Cable replaced per SWO-02917. | Customer damaged the yellow cable on 103472 during replacement of the foot sensor and the unit was unable to read one of the pads. | The device behavior reported was isolated to damage to the shank cable that carries the foot sensor signal. The cable was damaged during the replacement of the foot sensors. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 04 Oct 2017)
3812 | Missing Battery | closed | duanee | Torso | fixed | 2017-08-18 19:02 | 2017-08-18 | 4243 | 334 | Star Rehab | Joe Bongiovanni | Joe@STARrehab.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Specification Not Correct | Issue: The battery on one side of the Ekso is showing missing, if you take that same battery and change sides the Ekso works fine. It appears we have one battery if put on one side of the device it's not connecting properly and will show missing.... Matt our PT has narrowed this down to 1 battery and 1 particular side. | Replaced dowel pins with those of correct size per SWO-02944. | Engineering investigated battery engagement and discovered that dowel pins of the incorrect size had been specified when developing the GT build process. | The investigation into this uncovered a slight discrepancy between the design intent and implementation of the battery interface mechanics. In rare circumstances, the tolerance stack up could result in a condition in which the battery connection to the device did not engage fully. In this case, the connection and, resulting communication between the device and the battery, could fail. This triggers the Built-In Test (BIT) routines on the device and prevents the continued use of the device. CO-11942 released a change to the Device Master Record to shorten the dowel pin (C101646) in the torso assembly to remove the potential interference fit, and the dowel pins on this device were replaced to agree with the change. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3811 | Unit #235 - Rollingball/Gerry McCabe - black rubber peeled off | closed | duanee | Foot Sensor | fixed | 2017-08-18 12:46 | 2017-08-18 |  | 235 | Rolling Ball - Ireland | Jane Evans | jane@rollingball.ie | +353 45882600 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | customer reported the black rubber came off completly from foot sensor | replaced foot sensors, swo 2853 | This is a known wear component. The cap soles that were replaced were not returned for further analysis. | The tearing of the sole does not present a significant performance issue with the use of the device. Similar to wear on the sole of shoes, to some degree, this is expected wear. The complaint database was reviewed for similar issues. The cap sole design was updated with CO-11648 in Feb 2017. (D. Edmonds ‚Äì 04 Oct 2017)
3808 | Unit #171 Loaner - RKU Ulm - movement of the lower left leg not smooth | closed | mblaschke | Lower Leg | wontfix | 2017-08-18 08:09 | 2017-08-17 | 457267 | 171 | ULM - Germany | Sebastian Lux | sebastian.lux@rku.de | +49 731 1775530 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Undetermined | The customer found that the movement of the lower left leg is not running smoothly. Exchange of battery was done, it did not solve the problem but the customer can still use the loaner. | (Explain actions taken to resolve this issue) |  | This complaint is not a safety issue. The resolution of this has been added to the service maintenance log to be updated the next time the unit is serviced, and the complaint ticket is being closed.
3807 | Foot Sensor Issue | closed | duanee | Foot Sensor | fixed | 2017-08-17 20:29 | 2017-08-17 |  | 309 | Swedish Covenant | Cody Fisher | CFisher@schosp.org | 773.878.8200 ext. 4918 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | Issue: I got the new Velcro installed and took the Ekso for a test spin this afternoon. By looking at the affected side compared to the unaffected side I could still see a little difference with the toe of the sensor not being quite as flush with the foot plate as the other side. With the step height set at 0, if I went along with the machine and gave a little effort it there was no dragging. If I tried not to go along with the machine and gave a little downward force (as a very weak patient might do), there was still a little catching. Then what I did was put a piece of tape at the toe of the sensor to keep it flush onto the foot plate and that took care of it. What I'm thinking is that if I use a little hot glue or something to give it just a little extra stick at the foot plate onto the sensor, that should take care of it... thoughts? | Sent parts via SWO-2925 | The parts that were replaced in the field are not available for further analysis. | INV-0006
3806 | Unit #0316 - Rheinhessen Fachklinik Alzey - no charging and interfeace defect | closed | duanee | Battery - Charger | fixed | 2017-08-17 10:34 | 2017-08-17 |  | 316 | Rheinhessen-Fachklinik Alzey - Germany | Jochen Gradwohl | j.gradwohl@rfk.landeskrankenhaus.de | 49670749 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Customer Use | During his visit Michael found that the charger has a strongly bent pin and can not charge. Moreover the buttons of the interface do not work. | New charger shipped per SWO-02923. | Customer damaged the pins on the battery charger. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3803 | Unit #162 -Hobbs Rehab - rotating of lower left leg | closed | duanee | Lower Leg | fixed | 2017-08-16 10:18 | 2017-07-27 | 131722 | 162 | Hobbs - UK | Louis Martinelli | louis.martinelli@icloud.com | +49 40 800 4049 20 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Customer Use | The customer found that the left lower leg of the Ekso has been rotating. Initially he thought this was only happening on occasion when turning but it's happening very frequently now. | cleaned tubes on lower shanks | The clamping force for the shank decreased over time as the device was used. This appears to have been caused by the shank tube and/or the shank clamp getting contaminated by dirt, etc. | Cleaning the shank tubes resolved the turning issue and returned the shank assembly to sufficient holding force. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3802 | Poor battery engagement from oversized dowel pins. | closed | duanee | Torso | fixed | 2017-08-15 18:04 | 2017-08-17 | 760925 | 220 | Good Shepherd 3 | Alyssa Hauck | AHauck@gsrh.org | 484-523-0756 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Specification Not Correct | Unit was experiencing battery connectivity issues. Engineering discovered the issue was due to over-sized dowel pins being used for the installation of the battery mount. | Replaced dowel pins with those of correct size. | Engineering investigated battery engagement and discovered that dowel pins of the incorrect size had been specified when developing the GT build process. | The investigation into this uncovered a slight discrepancy between the design intent and implementation of the battery interface mechanics. In rare circumstances, the tolerance stack up could result in a condition in which the battery connection to the device did not engage fully. In this case, the connection and, resulting communication between the device and the battery, could fail. This triggers the Built-In Test (BIT) routines on the device and prevents the continued use of the device. CO-11942 released a change to the Device Master Record to shorten the dowel pin (C101646) in the torso assembly to remove the potential interference fit, and the dowel pins on this device were replaced to agree with the change. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3800 | Unit #131 - RKU ULM - rattling noise at the left knee | closed | duanee | Lower Leg | fixed | 2017-08-09 14:42 | 2017-08-08 | 213638 | 131 | ULM - Germany | Dominik Mayer | dominik.mayer@rku.de | +49 731 1775530 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Design | The customer suddenly heard rattling noise in the left knee. After closing down, starting a new with leg test and exchange of batteries the noise still did not disappear. | replaced cca 102286 | The printed circuit board assembly that controls the current to the knee brakes (102286) was found to be faulty resulting in undesired fluctuations in the current to the brake assembly. This has a side effect of creating an audible noise. | The reported issue was found to be caused by a faulty circuit board (102286 ‚Äì CCA, THIGH, EKSO 1.2 ‚Äì see Root Cause Analysis). The circuit board was replaced, resolving the issue. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3799 | Unit #270 - AUVA Bad Haering - replace crutch interface | closed | duanee | Cane / Walker / Crutch | fixed | 2017-08-09 12:49 | 2017-08-02 | 27031 | 270 | AUVA - Bad Haring, Austria | Mrs. Heidrun Laner | heidrun.laner@auva.at | +43 53327901001 | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Materials | Our PT noticed that the crutch interface worked for 1 day but not for the 2nd day during training. The error was #164 showing "deactivation of the interface". | New crutch interface sent per SWO-2908. | Issue with the connection of the crutch interface to the torso shroud. | This issue (crutch interface failure) was investigated and the connection between the crutch interface and the torso was found to be faulty. The complaint database was reviewed for similar issues and the previous risk assessment was reviewed based on recent field data and found to be consistent with the probability reflected in rev. W of 102437. No further action is necessary. (D. Edmonds ‚Äì 18 Jan 2018)
3795 | Amp Board Issue | closed | duanee | System | fixed | 2017-08-07 21:48 | 2017-07-26 | 7771 | 333 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org | 509.665.6156 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | Investigation Needed | Design | Issue: Per usual, as I give the all clear alert we have something weird happen. Grace and an aide had a patient in the new device (ID:60 I think). Within the first 5 steps, they had a strange error, labeled ‚ÄúMotor Error.‚Äù It required a restart, which they completed. Then after taking a few more steps they got the same error. We finished the session with our original unit and I got in the new device to see if I could elicit the response. I walked for almost 200 steps and never got the error message. | Reworked amp board installed into unit per SWO-02911. | New revision of amp board is causing the breaker to trip at a much lower value that designed. Rev K amp boards are tripping between 10-15 amps instead of the designed 27 amps. | The errors reported (Error 121 and Error 133 ‚Äì "Motor Error Ekso in Safe Mode") indicate fault conditions detected by the motor amplifiers. Through the Built-In Test (BIT) motoring routines, the device responds to the fault conditions by transitioning to a programmed Safe Mode, disabling the motor drives and engaging the knee brakes. This is the expected behavior in this condition and does not present a significant safety hazard to the patient or operator of the device. The core cause of this behavior was isolated by the R&D, Field Service, and Customer Experience teams to a change in the amplifier board that had the uncommon and unfortunate effect of causing the amplifier bus to cut out under conditions under which it should not. The amplifier board on this device was replaced with a revision that corrected this issue. No further action is necessary. (D. Edmonds ‚Äì 26 Oct 2017)
3794 | Missing Battery Error | closed | duanee | Pick an option | fixed | 2017-08-07 21:43 | 2017-07-26 | 6049 | 336 | Mercy Rehab Hospital | Candace Pella | candace.pella@allina.com | (763) 236-6000 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Specification Not Correct | Issue: Hi team - At Mercy today we have gotten three missing battery errors - 2 right and 1 left. Even after switching to second set of batteries. Each time, the batteries appear to have been seated properly and fully engaged (silver tab down). | Both battery side bars were replaced, as well as a new AMP board on SWO-02938. | Engineering investigated battery engagement and discovered that dowel pins of the incorrect size had been specified when developing the GT build process. | The investigation into this uncovered a slight discrepancy between the design intent and implementation of the battery interface mechanics. In rare circumstances, the tolerance stack up could result in a condition in which the battery connection to the device did not engage fully. In this case, the connection and, resulting communication between the device and the battery, could fail. This triggers the Built-In Test (BIT) routines on the device and prevents the continued use of the device. CO-11942 released a change to the Device Master Record to shorten the dowel pin (C101646) in the torso assembly to remove the potential interference fit, and the dowel pins on this device were replaced to agree with the change. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3793 | Amp Board Issue | closed | duanee | System | duplicate | 2017-08-07 21:20 | 2017-07-20 |  | 303 | Texas Rehabilitation Hospital | Matt Knowlton | MKnowlton@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Unassigned |  | Undetermined | Issue: Unfortunately the loaner we sent to support Matt and his ITW began having issue‚Äôs late this afternoon. Unfortunately, I was not able to repair the loaner unit that was having issues yesterday. This is another bad amp board and the third one this week (both the loaner and the customer device at Texas Rehab as well as the device at Renaissance in Edinburg, Texas). Renaissance and the loaner at Texas Rehab appear to have the same issue with intermittent power to the knee brake. Since we cannot currently qualify amp boards in the field, we will have to ship all of these units back to HQ. Another loaner will be arriving at the hotel around 9-10am today and I will be driving it straight to the hospital. I will let everyone know when the loaner is at the hospital and the first session with the new loaner is completed. | (Explain actions taken to resolve this issue) | (Required for complaints)
3792 | Amp Board Issue | closed | duanee | System | duplicate | 2017-08-07 21:15 | 2017-07-21 |  | 335 | Texas Rehabilitation Hospital | Matt Knowlton | MKnowlton@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Unassigned |  | Undetermined | duplicate of #3781 | (Explain actions taken to resolve this issue) | (Required for complaints)
3791 | unit #046 - Instituto Guttmann - battery not recognized | closed | duanee | Battery | fixed | 2017-08-03 12:16 | 2017-08-02 | 485139 | 46 | Guttmann - Spain | Silvana Andrade | gestiodeserveis@guttmann.com | +34 93.497.77.00 ext. 3199 | Spain | No injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | The battery does not get recognized from the ekso anymore, not matter if they connect it left or right. (error message: "Falta bat. Izquzierda" and "Falta bat. Derecha"). Battery serial: 1B serial-013356201500008F. | Batteries shipped per SWO-02891. | The communication between the battery and the device has failed in a way that has been isolated to this specific battery itself. The same behavior is shown with different devices and different battery chargers. The battery is not available for deeper analysis. | This issue has been isolated to a failure in the communication of this battery to the device (see Root Cause Analysis). The issue has been isolated to the battery itself, and there is no indication that there is any issue with the device it was used with. The battery has been replaced. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3785 | Hip abduction lock turning with adjustment. | closed | duanee | Adjustments - Hip | fixed | 2017-07-24 10:22 | 2017-06-15 | 146043 | 300 | RIM | Phoung Vu | pvu@dmc.org | 313-745-9745 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Design | During field investigation, technician noticed that the hip abduction lock (i.e. the control bar) would rotate sometimes while changing hip abduction adjustment. | Plunger re-seated per SWO-02785. | Plunger that holds the hip abduction lock was not seated properly. | A component in the hip abduction/adduction assembly was replaced (see Resolution Description) and testing showed that the issue reported by the customer was no longer present. The complaint database was reviewed for similar reports, and the reliability of the hip abduction/adduction adjustment and lock assembly has been added as a topic for discussion with the Manufacturing Engineering and Field Service teams. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3781 | Multiple errors #133 and #121 | closed | duanee | Pick an option | fixed | 2017-07-20 22:52 | 2017-07-19 |  | 335 | Texas Rehabilitation Hospital | Donna Robacker | drobacker@texasrehabhospital.com | (817) 820-3400 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major | Investigation Needed | Design | Issue: I am on Day 2 of an ITW at the Texas Rehabilitation Hospital of Fort Worth. Starting this morning, we have experienced 6 errors that caused Safe Modes. The errors seemed to become more frequent as the day progressed. These errors occurred under different circumstances with 4 different participants using the Ekso. The errors are listed below: 1) #121 LH error during PT in Ekso session [Day 2] (PT attempting forceful step, took an exaggerated step R and then LH error occurred; long leg sit to chair) 2) #133 R Knee during Participant 3 session [Day 2] (Pt took right step which landed early resulting in error; acknowledged error, returned to stand, and proceeded with walking) 3.) #121 R hip error during Participant 1 session [Day 2] (Pt took a right step and foot caught the ground causing error; acknowledged error, returned to stand, and proceeded with walking) 4.) #133 L Knee during Participant 2 session [Day 2] 5.) #121 R hip error during Participant 2 session [Day 2] 6.) #133 R Knee during Participant 2 session [Day 2] The device serial number is 021706003 Step Count is 6413 steps Please let me know if you can troubleshoot this from your end. | Reworked amp board installed into unit per SWO-02880. | New revision of amp board is causing the breaker to trip at a much lower value that designed. Rev K amp boards are tripping between 10-15 amps instead of the designed 27 amps. | The errors reported (Error 121 and Error 133 ‚Äì "Motor Error Ekso in Safe Mode") indicate fault conditions detected by the motor amplifiers. Through the Built-In Test (BIT) motoring routines, the device responds to the fault conditions by transitioning to a programmed Safe Mode, disabling the motor drives and engaging the knee brakes. This is the expected behavior in this condition and does not present a significant safety hazard to the patient or operator of the device. The core cause of this behavior was isolated by the R&D, Field Service, and Customer Experience teams to a change in the amplifier board that had the uncommon and unfortunate effect of causing the amplifier bus to cut out under conditions under which it should not. The amplifier board on this device was replaced with a revision that corrected this issue. No further action is necessary. (D. Edmonds ‚Äì 26 Oct 2017)
3780 | Issues w/ R Knee | closed | duanee | Pick an option | fixed | 2017-07-20 22:47 | 2017-07-20 | 43865 | 303 | Texas Rehabilitation Hospital | Donna Robacker | drobacker@texasrehabhospital.com | (817) 820-3400 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Design | This issue has been report to Josh Witherspoon who developed a plan of action to remedy the following situation. An engineer and loaner device will be sent out to our location in Fort Worth. No further action needed at this time. During our last patient session of today (Day 3 of 4), the Ekso's right knee started making a buzzing sound that progressively increased in volume. The patient had only taken a few steps in the device. We had the patient sit and the sound seemed to go away when I pressed the right knee brake release button. I restarted the Ekso and freed the Ekso legs. When attempting to move the legs (knee extension and flexion) I experienced a cogwheel-like, intermittent resistance. I replaced the 9V battery to the brake release button thinking that the brake release was intermittently turning on and off. However, this did not seem to fix the problem. After troubleshooting with Josh on the phone, he and the rest of the team decided it would be best to have an engineer travel out and take a look. Please let me know if you have any questions. | Amp board replaced per SWO-02915. | Dan determined that the fuse that controls the power to the knee brakes was undersized on the amp board. The fuse was tripping on and off so quickly that it was creating a "buzzing" sound coming from the brake. | The increased resistance in the motion of knee was found to be caused by a current-limiting component on the amplifier printed circuit board assembly (PCA) behaving in a way that caused the knee brake not to fully disengage as desired, resulting in a situation where the motion of the knee was resisted by the partially engaged brake. The complaint database was searched, and it has been determined that this behavior is not isolated to this occurrence. The amplifier PCA (102284) was updated with CO-11970 to address this issue. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3778 | L Leg motor | closed | duanee | Pick an option | fixed | 2017-07-19 16:52 | 2017-07-18 | 28596 | 305 | Doctors Hospital at Renaissance | Nathan Dunn | nathand22@aol.com | 956-661-7100 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Issue: My name is Nathan Dunn and I am a physical therapist at Doctors Hospital at Renaissance and have experienced a problem with our Ekso unit. The problem started earlier this morning with a loud buzzing sound on the left motor that is close to the patient's thigh. Later when trying to use the Ekso again, the noise returned but now the left leg was not able to freely move when ranging the patient before coming to a stand. I have attached a video of the problem we encountered. The SN of the Ekso is: 021610004 Thank you for you help | Amp board replaced per SWO-02915. | Dan determined that the fuse that controls the power to the knee brakes was undersized on the amp board. The fuse was tripping on and off so quickly that it was creating a "buzzing" sound coming from the brake. | INV-0011
3776 | Issues with controller | closed | lmcneil | UI - Controller | fixed | 2017-07-18 23:01 | 2017-07-06 |  | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: Yesterday at about 4:30 we had a patient (ID:13) walking in the EKSO. During the swing phase of the left leg, we encountered an error, which locked that leg in a bent positon. The caused us patient to stand there like a flamingo. Right knee locked in extension, left knee locked in about 45 degrees of flexion. The EKSO was making the ‚Äúerror‚Äù sound, but when I looked at the controller, it was completely blank. The backlight was on but no text was displaying. I hit the yellow button which stopped the ‚Äúerror‚Äù sound. The screen continued to be blank. I hit the double back which took me to a screen that told me to acknowledge the error but it would not let me progress to a different screen. We did a power-down long sit, got everything ready for another stand, and continued to walk for another 18 minutes with no errors. I looked in the diagnostics to see what the error code was, but it said no error was detected. The particular patient that we had in the EKSO is fairly experienced and we just happened to be close to a wall so we avoided a fall. But, other patients with less experience the results may not be a great. I‚Äôm wondering if this is an issue we need to address immediately, or if you can tell me what I can do on my end to make sure this does not happen again? We have patients today from 11-3. Let me know if you need any more details. | Parts sent via SWO-2843 | The symptoms of this issue appear to be isolated to the user interface controller assembly. The issue has not recurred since the controller was replaced. However, it is not possible to fully confirm that the controller caused the reported behavior because the original issue could not be repeated. The controller that was replaced is not available for further analysis. No further action is necessary. (D. Edmonds ‚Äì 06 Nov 2017) | The degree to which this issue could be investigated was limited by the inability of reproducing the reported behavior (see Root Cause Analysis). The complaint database was searched for similar issues, and two complaints with symptoms that may be related were found and reviewed (3342 and 3175). The current risk assessments for 1.2 FMEA-758 and 1.2 FMEA-762 (102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet) were reviewed, and it was determined that no update was necessary based on recent field data. No further action is necessary at this time. (D. Edmonds ‚Äì 06 Nov 2017)
3775 | R Knee get's caught | closed | duanee | Foot Sensor | fixed | 2017-07-18 19:48 | 2017-07-13 | 28889 | 309 | Swedish Covenant | Cody Fisher | CFisher@schosp.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Issue: Hi I am at Swedish covenant SN 021611004 And when In 2 free, the right knee gets caught at about 35 degrees when walking. I have Included a video. It happens intermittently. | Sent a set of sensors via SWO-2868 | The foot sensors have either worn in to a point where they require re-calibration or one of the sensing pads is not reading correctly. | INV-0006
3774 | unit #243 - Step by Step/Barcelona - Error RK- 27 | closed | duanee | Unknown | fixed | 2017-07-13 15:25 | 2017-07-13 |  | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | Error RK- 27. Problem sensor on right knee. I did the "leg test" and "analog sensors" and they are ok, but when we try to stand up, the right knee is block and reset is needed. | replaced parts, swo 2926 | Multiple components (foot sensors, knee assemblies, cabling, etc.) were replaced on this unit to address the issue reported by the customer and return the unit to operation quickly. It is no longer possible to isolate a true root cause for this issue. | Multiple components (foot sensors, knee assemblies, cabling, etc.) were replaced on this unit to address the issue reported by the customer and return the unit to operation quickly. The unit was successfully tested and returned to the customer. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2018)
3773 | Dead joint board | closed | duanee | PCBA | fixed | 2017-07-12 22:54 | 2017-07-12 | 518822 | 161 | Glenrose - Canada | Vickie Buttar | Vickie Buttar <Vickie.Buttar@albertahealthservices.ca> |  | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | Investigation Needed | Materials | Knee CCA was malfunctioning and was causing knee brake to lock up intermittently while knee was actuating. see attached videos | Replaced knee CCA per currently accepted practices. | Knee CCA was malfunctioning and was causing knee brake to lock up intermittently while knee was actuating. | The issue reported was narrowed to a failure of one of the knee joint boards, which caused the knee brake to engage unexpectedly. The knee joint board was replaced, and Service was able to verify that the issue was resolved by both confirming that the issue could not be reproduced with the new joint board and by also putting the original board back into the unit and confirming that the original issue came back. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3772 | Battery charger not charging | closed | duanee | Battery - Charger | fixed | 2017-07-12 19:48 | 2017-07-12 | n/a | 279 | Sunnyview Rehab | Patti Valenza | Patricia.BroschartValenza@sphp.com | 518-382-4528 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | Issue: Hi guys, We are still having trouble with one of the chargers. One of the pins is bent, and sometimes is not charging the battery. Allison was able to straighten the bent pin, but the next time the charger was used, it was bent again. In addition, when they put the batteries into the Ekso, it is only reading 87 or 88% charged, even though it is green on the charger. Patti | Sent new parts via SWO-2856 | It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3770 | unit #235 - Rolling Ball/Garry McCabe - ankle adjustment loose | closed | duanee | Adjustments - Ankle | fixed | 2017-07-12 14:27 | 2017-07-12 | 100566 | 235 | Rolling Ball - Ireland | Jane Evans | jane@rollingball.ie | +353 87 224 6287 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | clinical reported a loose ankle adjustment knob | replace ankle assemblies, swo 2853 | It is no longer possible to further isolate the root cause of this issue. Due to the priority of addressing this for the customer and the low-risk nature of this behavior, the assemblies were replaced and the initial issue can no longer be reproduced. | The knob in the ankle adjustment assembly as reported to be loose during the initial training week. The knob did continue to perform its function. The ankle assemblies were replaced. No further action is necessary. (D. Edmonds ‚Äì 15 March 2018)
3769 | Error 50 LK | closed | mblaschke | Encoder / Potentiometer | fixed | 2017-07-12 12:00 | 2017-07-12 | 695255 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Determined Unknown | Customer reported error 50 LK during setup of patient. | Leg test, no parts | The root cause for this issue cannot be fully determined because upon retest and calibration of the unit, the reported behavior could no longer be reproduced and there have been no additional complaints for this unit. | The error (Error 50 ‚Äì "Pot out of range") thrown indicates that the potentiometer that is used for redundant feedback on the left knee reported a value that is outside the operating range of the sensor. The Built-In Test (BIT) routine on the device correctly detected the condition and prevented the continued use of the device while this condition was present. The investigation into this issue is limited by the fact that after the leg test/calibration routine was run, the unit no longer demonstrated the reported behavior (Error 50). No further action is necessary. (D. Edmonds ‚Äì 07 Nov 2017)
3768 | Repeated #137 | closed | duanee | Pick an option | fixed | 2017-07-11 23:03 | 2017-07-03 | 3203 | 281 | Franciscan Health Lafayette | Corry Voll | corry.voll@franciscanalliance.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Wear | Issue: One of our outpatient therapists has provided me the following info: The 137 error happened to me twice in op recently. both times were when we had put the device in standstill (resulting in split stance position) and I was helping the patient move a leg back to be even with the other for stable stand rest position. I've been told the #137 error occurred once in the rehab unit and at least 2-3 times in OP. Thanks, Corry | Right knee motor was replaced entirely according to currently accepted practices. | Brake slipping was a repeated issue with this unit, and at this point it was suspected that the shaft was badly worn and marred from attempts to reseat brake set screw. | The error reported (137 ‚Äì "Motor Error" "Brake failed to hold position") indicates that the software check to ensure the knee brakes are operating correctly failed. The motor and brake assembly was replaced on this unit, and we were able to verify that the brake error could no longer be reproduced. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3767 | Hip Abduction Slipping | closed | duanee | Adjustments - Hip | fixed | 2017-07-07 17:52 | 2017-06-12 | 460146 | 106 | RIC | Kristen Hohl | khohl@ricres.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Design | Issue: Customer reported hip abduction settings to be slipping several increments during walking. | Control bars have been replaced per SWO-02396. | Design issue with the control bar being back driven during use. | The control bars on the device were replaced (see Resolution Description) and testing showed that the issue reported by the customer was no longer present. The complaint database was reviewed for similar reports, and there do not appear to be significant numbers of occurrences of this issue in the field. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3766 | Hip width adjustment | closed | duanee | Adjustments - Hip | fixed | 2017-07-07 17:42 | 2017-06-07 | 493743 | 161 | Glenrose - Canada | KristinTarnowski | Kristin.Tarnowski@albertahealthservices.ca |  | Canada | No injury | Same as contact listed | Pick an option | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Materials | Issue: After a recent field service visit customer reported still having issues with the hip width adjustment, especially on the L side. During prior visit, issue was only reported with R side and L hip width was fine when tested by technician. Hey Lily, thanks so much! The abduction is fixed it‚Äôs the hip width that‚Äôs the problem ÔÅä its really really hard to turn the screw driver first but once you turn it back and forth a few times it does become easier. | Replaced both hip abduction assemblies according to currently accepted practices. | Hip abduction assemblies were replaced and it was confirmed that the issue could no longer be reproduced. | The difficulty in adjusting the hip width was addressed by replacing the hip adjustment assemblies on both sides of the device. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3765 | Left knee resistance | closed | duanee | Adjustments - Leg | fixed | 2017-07-07 17:35 | 2017-06-19 | 146043 | 300 | RIM | Phuong Vu | pvu@dmc.org | (313) 966-8309 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Design | Issue: Customer reported that the left knee was sometimes very difficult to move when left leg is set to "2 free" | Amp board replaced per SWO-2785. | Dan confirmed that the fused that is used to supply the knee brakes was tripping. Fuse is suspected to be undersized. | The increased resistance in the motion of knee was found to be caused by a current-limiting component on the amplifier printed circuit board assembly (PCA) behaving in a way that caused the knee brake not to fully disengage as desired, resulting in a situation where the motion of the knee was resisted by the partially engaged brake. The complaint database was searched, and it has been determined that this behavior is not isolated to this occurrence. The amplifier PCA (102284) was updated with CO-11970 to address this issue. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3764 | Pulled sensor wire | closed | duanee | Foot Sensor | fixed | 2017-07-07 17:24 | 2017-06-12 | 2436 | 307 | RIC | Kristen Hohl | khohl@ricres.org | 312-238-6875/773-255-5800 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Mishandling | Issue: We had a chance to use the loaner for a patient this week and had a bit of trouble with it that we wanted to run by you. We had issues with the first step on the R - only way I was able to get it was changing his step height from 0.4 to 1.0. Also we noticed asymmetry in the motor sounds between legs. We only worked briefly with the patient due to these concerns and then retrialed it today with sara and I in the ekso. With the "lighter" sound on the right, it also feels weaker in extension. We were walking in bilateral adapt. I attached a video but don't know how well it will capture the noise. Sent out foot sensors to replace this issue via SWO-2782, however during attempted replacement the customer pulled out the ankle cable foot sensor plug - resulting in needing a field service visit. | Replaced shank wire. | Customer accidentally pulled wire out of plug when attempting to swap foot sensor. | The issue reported was caused by damage to the shank cable during the replacement of the foot sensors. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 04 Oct 2017)
3763 | Pulled sensor wire | closed | hrayner | Foot Sensor | duplicate | 2017-07-07 17:24 | 2017-06-12 | 2436 | 307 | RIC | Kristen Hohl | khohl@ricres.org | 312-238-6875/773-255-5800 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Undetermined | Issue: We had a chance to use the loaner for a patient this week and had a bit of trouble with it that we wanted to run by you. We had issues with the first step on the R - only way I was able to get it was changing his step height from 0.4 to 1.0. Also we noticed asymmetry in the motor sounds between legs. We only worked briefly with the patient due to these concerns and then retrialed it today with sara and I in the ekso. With the "lighter" sound on the right, it also feels weaker in extension. We were walking in bilateral adapt. I attached a video but don't know how well it will capture the noise. Sent out foot sensors to replace this issue via SWO-2841, however during attempted replacement the customer pulled out the ankle cable foot sensor plug - resulting in needing a field service visit. | (Explain actions taken to resolve this issue) | (Required for complaints)
3762 | unit #316 - Rheinhessen Klinik Alzey - Pulse report showing wrong start time | closed | bantonio_schroll | Pulse Data | wontfix | 2017-07-06 15:40 | 2017-07-06 |  | 316 | Rheinhessen-Fachklinik Alzey - Germany | Jochen Gradwohl | 49670749 | j.gradwohl@rfk.landeskrankenhaus.de | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Undetermined | At first the customer reported that the start times are wrongly displayed (by two hours) in the Pulse portal. He then corrected the time zone in the Pulse settings and the Pulse portal then started to show correct times. But when he then downloaded the report from Pulse it did not show the same start times than the Pulse portal but still the wrong ones from before changing the settings (see screenshot attached). | There is no safety hazard presented by this error the Ekso Pulse data. This issue will be addressed with the latest Revison of Pulse.
3761 | unit #243 - Step by Step/Barcelona - broken joint screw in left knee | closed | duanee | Actuators | fixed | 2017-07-06 15:15 | 2017-07-06 | 649.282 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Design | customer reported a broken joint screw in the left knee of our unit n¬∫ 243. | customer replaced joint caps, and screw | The loading on the joint caps is different than assumed for some customers, and the component may need to be replaced prior to the full design life of the device. | The risk evaluation (Failure Modes and Effects Analysis, 102437AB ‚Äì Ekso 1.2 Risk Management Spreadsheet, rev. AB) in the Risk Management File for this failure was reviewed. The probability of failure and severity of potential harm agree with feedback from the field. The customer issue has been addressed by the replacement of the joint caps. No further action is necessary at this point. (D. Edmonds ‚Äì 17 May 2018)
3747 | UI Screen Scrambling | closed | ckealoha | UI - Controller | fixed | 2017-06-28 20:05 | 2017-06-27 | 493743 | 161 | Glenrose - Canada | Kristin Tarnowski | Kristin.Tarnowski@albertahealthservices.ca |  | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | During service, technician noticed UI screen scrambled several times. While customer did not report having seen the scrambling, UI should be replaced for good measure. | New UI shipped to customer | This issue was caused by a susceptibility of the LCD PCA used in the control pad assembly to electrostatic discharge (ESD), which at times causes the LCD display to scramble requiring a power cycle to correct. | The complaint database was searched for similar issues, and two complaints with symptoms that may be related were found and reviewed (3342 and 3175). The current risk assessments for 1.2 FMEA-758 and 1.2 FMEA-762 (102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet) were reviewed, and it was determined that no update was necessary based on recent field data. No further action is necessary at this time. (D. Edmonds ‚Äì 06 Nov 2017)
3745 | Errors 27 LH | closed | duanee | Foot Sensor | fixed | 2017-06-27 23:46 | 2017-06-17 |  | 181 | VA Bronx | Hong, EunKyoung | EunKyoung.Hong@va.gov |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | Issue: We (Bronx VA) have issues with Ekso. Last Thursday, Ekso(right leg) was kicked backward so quickly. After that was happened, we recorded walking to catch any weird moments. I will attach one video which is slowly kicked backward. With kicking again, Ekso made noise little bit different than usual. Therefore, we stopped training. Today, we started walking with a patient. Suddenly, an error was occurred. That error was ‚Äú#27-LH‚Äù. Please see pictures. Also, I am attaching one more picture of Ekso which shows left hip was fixed some point. This situation was happened when I tried getting the serial number. Please see pictures and video and let us know. | Sensor error - external cable replaced per SWO-02797. Leg kick - foot sensor replaced per SWO-02797. | Intermittent issue with the external cable on the left hip. Leg kick - Occurred due to noise in the right foot sensor toe. | The cause of the error thrown by the BIT routine was narrowed to an external cable on the device. The BIT routine correctly prevented the use of the device in this condition. The cable and foot sensors were replaced, and the device successfully passed the Service Walk Test. No further action is necessary. (D. Edmonds ‚Äì 04 Oct 2017)
3744 | unit #178 - Technomex/Kamien Pomorski - battery not working anymore | closed | duanee | Battery | fixed | 2017-06-26 14:49 | 2017-05-29 | 179803 | 178 | Unassigned | Jakub Maliszewski | jmaliszewski@technomex.pl | +48 32 40 10 350 w. 38 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | One of batteries is not recognized by device anymore and even the chargers do not see it. The distributor tested the battery in different device and with different chargers and result stays this same. | Batteries replaced per SWO-02818. | The communication between the battery and the device has failed in a way that has been isolated to this specific battery itself. The same behavior is shown with different devices and different battery chargers. The battery is not available for deeper analysis. | This issue has been isolated to a failure in the communication of this battery to the device (see Root Cause Analysis). The issue has been isolated to the battery itself, and there is no indication that there is any issue with the device it was used with. The battery has been replaced. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3742 | unit #232 - OrthoMedico/To Walk Again - battery looses during walking | closed | duanee | Battery - Connector | fixed | 2017-06-21 14:49 | 2017-06-21 | 754825 | 232 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Design | right battery looses during walking (several times, also after double-check) | replaced battery interface, swo 2827 | The core cause of this issue was due to an intermittent connection between the battery and the device. This could have been caused by a problematic connector on the battery interface panel, however, it is possible that this was caused or exacerbated by a recently discovered tolerance stack up issue related to the length of a dowel pin in the battery interface assembly on the torso. | This issue was resolved for this customer by replacing the battery interface on the unit. The complaint database was searched for similar issues, and this does not appear to be isolated. Upon further investigation, it was determined that the original release of the Ekso GT battery interface in the torso assembly included an error in which a dowel pin was incorrectly sized, causing less than intended engagement between the battery and device connectors. The dowel pin (C101646) length was corrected with CO-11942. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3741 | unit #266 - Orthomedico - motor error #1057-LH | closed | duanee | Motor | fixed | 2017-06-21 14:44 | 2017-06-21 | 426113 | 266 | OrthoMedico NV - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 (0)479 62 99 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | device shows motor error !#1057-LH. Today I had a demo and during demo this error occured. | replaced hip housings, SWO-02842 | After service was performed on this unit (SWO-02842), the error could not be reproduced. It is not clear, however, that the service performed is directly related to the error thrown. It is possible that the communication error was a rare occurrence. Without being able to reproduce the behavior reported, it is not possible to isolate the root cause of this issue further. | The error shown (1057 ‚Äì "Software Error Ekso in Safe Mode CAN PDO received with error") indicates that there was an error with the CAN bus that is used for communication between the main processor and the amplifiers that control current to each of the hip and knee motors. When a communication error is detected, the device is designed to prevent the use while this condition is present. We were not able to reproduce the error further following the service performed on SWO-02842. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3740 | Unit #187 - Care Technique - replace foot sensors | closed | duanee | Foot Sensor | fixed | 2017-06-21 14:44 | 2017-06-13 | 40811 | 187 | Care Techniques - Saudi Arabia | Dr.Moataz Hamada | moatazcaretechnique@gmail.com | 966547082930 | Saudi Arabia | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Materials | According to customer, the black rubber tread is coming off so we send new foot sensors to replace those. | New foot sensors sent per SWO-02809. | Tread glue was no longer adhered to the sole. | INV-0006
3739 | Unit #151 - Righospitalet Glostrup - replace of soft goods | closed | duanee | Soft Goods | fixed | 2017-06-20 13:55 | 2017-06-19 | 234161 | 151 | Glostrup - Denmark | Mette Skov Henriksen | mette.skov.henriksen@regionh.dk | +45 38631947 | Denmark | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | The Ekso was updated to the latest soft and hardware by Michael. We need to send new t-wrenches and a set of soft goods: side closures, thigh straps, tibial straps and torso pads. | Soft goods replaced per SRO-2805. | Velcro loop is known to wear during repeated use. | Known wear item.
3738 | Unit #252 - Ostracon/Animus - replace of soft goods | closed | duanee | Soft Goods | fixed | 2017-06-19 16:19 | 2017-06-14 | 84750 | 252 | Ostracon Ltd - Greece | Dimitris Aslanidis | physio@animus.com.gr | +30 2410 851851 | GREECE | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | During training, PT noticed that : 1.Their side rivet Velcro is worn out 2.Their shin straps Velcro is also worn out | New soft goods shipped to customer per SRO-2801. | Velcro loop is known to become frayed over use. | Known wear items.
3737 | unit #133 - Uppsala - foot plate loose | closed | duanee | Foot Sole | fixed | 2017-06-16 14:44 | 2017-06-16 | 352557 | 133 | Uppsala - Sweden | Peter Svensson | peter.svensson@akademiska.se | +46 18-611 28 44 | Sweden | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Wear | We have experienced one problem with the Ekso since you were here. There as been a cracking sound during some walking sessions and when that have occured we have noticed that some of the screws "under" the right foot are loose. We have tried to tighten them but the problem is coming back. | customer's technician thightened and glued the screws according to the info provided from Michael | Known design failure (see #3431) | Investigation performed previously (see #3431)
3736 | unit #127 - EMAC/Padre Pio - defective charger | closed | duanee | Battery - Charger | fixed | 2017-06-14 10:52 | 2017-06-14 | 531109 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Determined Unknown | this Ekso with Ekso Care has a problem with a battery charger that is not charge any more the batteries. | New charger sent per SWO-2791. | Connection issue with the battery. It is not possible to determine root cause because the battery charger is not available for analysis. | The complaint file database was reviewed for other instances of similar reports in the field, and, although the root cause may not be identical, this does not appear to be the only occurrence. The root cause, however, cannot be determined in this case because we do not have the battery charger in question for analysis. No further action is necessary. (D. Edmonds ‚Äì 28 Sep 2017)
3735 | unit #127 - EMAC/Padre Pio - sensor error #25 LH | closed | duanee | Encoder / Potentiometer | fixed | 2017-06-14 10:47 | 2017-06-14 | 531109 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Wear | Yesterday, when they have switch on the device it required a Leg Test. I followed by phone the customer to do it, as we had done a month ago more or less. But this time, after the left hip calibration, the error message "Potentiometer Error on Left Hip" is showed. Update: During checking the function of the hip potentiometer, the customer got ‚ÄúSensor Error‚Äù #25 LH | Replaced soft pot. | While we unable to reproduce the issue in-house via leg test, manual actuation, or extensive walk testing, we did discover that the soft pot had a groove worn in it. | The cause of this issue is very likely the potentiometer used for position feedback on the left hip. This is consistent with the reported behavior, and a faulty potentiometer would cause the BIT routine to generate a warning by triggering an error. The potentiometer was found to have a groove worn into it upon visual inspection, but we were not able to fully reproduce a failure. The visually-damaged potentiometer in question was replaced. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3734 | unit #146 - Regibase - left knee blocked | closed | duanee | Pick an option | fixed | 2017-06-13 15:47 | 2017-06-13 | 254033 | 146 | Regi Base - Czech Republic | Jakub Petioky | Jakub Pƒõtiok√Ω <jakub.petioky@rehabilitace.cz> | (+420) 317 881 111 | Czech Republic | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Left knee motor is locked, during stand up protocol blocks Lf knee (remains in 90 dg)= Ekso is not able to stand up. Mode Free leg in sitting sometimes release Lf- knee, after while completely block knee (clicking sound). | Replaced joint board per SWO-2497 | Faulty component on joint board 102286 caused the brake to intermittently turn off and on. | The cause of this issue was narrowed to a faulty left joint printed circuit board assembly. The faulty board assembly was replaced and the issue has not occurred again. No further action is necessary.
3733 | unit #140 - EMAC/Villa Beretta - no up-to-date data in Pulse | closed | duanee | Pulse Data | fixed | 2017-06-12 14:27 | 2017-05-31 | 1610468 | 140 | Distributor - Italy | Matteo Carelli | Matteo Carelli <m.carelli@emac.it> | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | Investigation Needed | Wear | no up-to-date data is displayed in Pulse Update Chad (May 31 2017): The issue must therefore be on the actual device and/or phone. When looking at the last time that the phone sent data we can see that the phone battery level hit 10%. The software on the phone causes all operations to be shutdown after 20%. We have seen phones where the phone gets stuck into a loop because of the low battery and how it interfaces with the software. I think we should inspect the device and be ready to swap out the phone. | Phone battery replaced per SWO-02816. | Phone battery known to degrade over time. | This issue was narrowed to a failure of the battery of the cellular phone in the device. The phone would not transmit the device data to the cloud. The phone was replaced. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3732 | unit #213 - EMAC/Istituto Agazzi - no up-to-date data in Pulse | closed | duanee | Pulse Data | worksforme | 2017-06-12 14:21 | 2017-05-31 | 423255 | 213 | Distributor - Italy | Matteo Carelli | matteo.carelli@mcmedtech.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor | No investigation needed - see Justification | Determined Unknown | In Pulse the last shown data is from February 10th. And in the M2M tool no data transfer is shown for the last 30 days. update Chad: all of the data suggests that the phone never turned on or the battery level on the phone was so low that none of the software on the phone could properly run...we should go out and be ready to replace that phone | No items where changed. Confirmed that the phone was functioning as expected. | Data was checked when device was in house and the most recent data is shown in pulse. | An investigation is not necessary because no issue with the device was found. (Duane Edmonds ‚Äì 10 Sep 2017)
3730 | Unit #276 - AUVA Wei√üer Hof - issue with foot sensors | closed | duanee | Foot Sensor | fixed | 2017-06-09 14:20 | 2017-06-09 |  | 276 | AUVA - Weisser Hof, Austria | Marianne Worisch | marianne.worisch@auva.at | 4308063 | Austria | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | The foot sensors need to be replaced | Foot Sensors were shipped via SWO-02774. Customer did foot sensor calibration. | There is not enough information to complete an RCA. | INV-0006
3728 | Right Leg Rotating During Walk | closed | duanee | Lower Leg | fixed | 2017-06-07 17:29 | 2017-05-22 | 73298 | 274 | Good Shepherd - OP | Alyssa Hauck | AHauck@gsrh.org | 484-523-0756 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Design | Right leg rotates inward during walk. | Collar replaced with most recent version per SWO-02762. | Adjustment screw bottoming out on collar. | The issue here (rotation of the leg during walk) has been reviewed and the root cause was determined to be that the adjustment screw(s) bottomed out and could not, as a result, provide adequate clamping force to hold as expected. This has been addressed and the shank collar on this unit has been replaced with the updated design. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3727 | Clicking in Left Hip | closed | duanee | Actuators | fixed | 2017-06-07 17:17 | 2017-05-22 | 73298 | 274 | Good Shepherd - OP | Alyssa Hauck | AHauck@gsrh.org | 484-523-0756 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Wear | Customer described hearing a click coming from the left hip during turning. | Spacer replaced per SWO-02762. | The spacer in the left hip was worn down. | The clicking sound observed when the device is loaded heavily is a perception issue and has been resolved by replacing a component on the device. No further action is needed. (D. Edmonds ‚Äì 21 Sep 2017)
3725 | Unit #170 - FLM Medular/Spain - strange walking | closed | duanee | Foot Sensor | fixed | 2017-06-07 14:09 | 2017-06-06 | 297523 | 170 | FLM - Spain | Sara C√°ceres | rehabilitacion@medular.org | +34 917775544 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Materials | The right leg is kind of stopping during the step so it does not make a flowing movement and at the end of the video the left leg is doing it, too. | customer did foot test which solved the issue | The foot sensor feedback appears to have drifted with use causing the need to recalibrate the device to the sensors. | INV-0006
3724 | unit #267 - Fundacion Nipace/Spain - knee break button broken | closed | duanee | Brake / Release / 9V Battery | worksforme | 2017-06-06 15:00 | 2017-06-06 | 401942 | 267 | Jannah Armstrong - UK | Mayte Vega | mayte@fundacionnipace.org | 34949254140 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Determined Unknown | internal PT reported that the knee brake button on the right leg does not work, replacing batteries did not solve the issue | checked wires and connectors, no issue. Reason probably loose connector. | (n/a ‚Äì unable to reproduce reported behavior) | (n/a ‚Äì unable to reproduce reported behavior)
3723 | Shoulder strap broke | closed | duanee | Soft Goods - Shoulder Straps | fixed | 2017-06-02 18:02 | 2017-06-02 | 1065880 | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | Issue: Good morning, Last night we had a buckle break when attempting to get a patient in the EKSO. It was the male part of the upper buckle on the shoulder strap. It broke right along where you thread through the strap. We ended up ‚Äúmacgyver-ing‚Äù a solution so we could still walk but for safety reasons on our patients that require more trunk support, the buckle is definitely needed. Let me know when you can get that buckle shipped out! Thank you! | Sent new parts via SWO-2763 | Known wear item | An assessment of risk was performed (see Risk Level Explanation) and current and previous complaint files were reviewed to understand the frequency with which this has occurred in the field. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3722 | unit #213 - EMAC/Istituto Agazzi - feet touch while walking | closed | duanee | Actuators | fixed | 2017-06-02 10:22 | 2017-06-01 | 732306 | 213 | Distributor - Italy | Matteo Carelli | matteo.carelli@mcmedtech.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Design | distributor report Like showed in the video, with different patients and different hip width and hip ABD, during the step the right leg hit the left leg, also with an ABD 4. Moreover they told me that in all the patients the left step is always longer than the right one and that the patients feel the right leg harder to move. | Replaced hip housings. | Hip housings were cracked which allowed for the feet to touch during walking | The root cause of this issue with the feet touching during walking was isolated to the cracks discovered in the hip housings of the unit. Both hip housings on this unit were replaced. No further action is necessary. (D. Edmonds ‚Äì 26 Oct 2017)
3716 | unit #170 - FLM Medular/Spain - battery not recognized | closed | duanee | Battery | fixed | 2017-06-01 10:06 | 2017-05-30 | 325886 | 170 | FLM - Spain | Sara C√°ceres | rehabilitacion@medular.org | +34 917775544 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | One of our batteries is not recognized by the ekso and it appears at the screen a message of incompatibility. And it doesn‚Äôt charge either. The serial number of the defective battery is: 01F3ECEA16000024. Part 102040B. We have tried in the four chargers we have and with the other batteries the chargers work well, not with the defective one. | customer received new set of batteries per SWO-02769. | The defective battery is not available for further analysis. | The issue has been isolated to a specific battery serial number, and the battery is no longer in use. The battery is not available for further technical investigation. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3715 | unit #213 - EMAC/Istituto Agazzi - green button on UI not working | closed | duanee | UI - Controller | fixed | 2017-06-01 09:47 | 2017-05-29 | 423255 | 213 | Distributor - Italy | Matteo Carelli | matteo.carelli@mcmedtech.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | distributor reported that this Ekso has problem with the UI, the green button doesn‚Äôt work well all the time. | send parts to distributor - SWO-02758 | Known wear component | The safety risk was discussed with a cross-functional team on 20 June 2017, and the risk level and explanation have been updated based on the discussion. The complaint database has been searched for similar issues. This appears to be an isolated incident. No further action is necessary. (Duane Edmonds ‚Äì 26 June 2017)
3714 | unit #141 - EMAC/Villa Adria - batteries discharging to quick | closed | duanee | Battery | fixed | 2017-06-01 09:42 | 2017-05-29 | 854155 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Wear | distributor reported that the batteries couple 2A and 2B are not working properly, with full charge the percentage come down quickly and the customer can do just one patient with a fully charge. | Batteries replaced per SWO-02768. | Batteries are known to degrade over time. | This set of batteries discharged faster than expected. This is expected over the life of the batteries. It is not clear, however, if this set experienced a faster than normal decline in capacity based on how they were used in the field. The batteries have been replaced, and the batteries that were replaced are not available for further analysis. No further action is necessary. (D. Edmonds ‚Äì 17 Oct 2017)
3711 | Repeated #1241 Errors | closed | duanee | Other | fixed | 2017-05-30 21:35 | 2017-05-30 | 888558 | 142 | Burke | Helma Zanders | wzanders@Burke.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | Issue: Burke called this morning, they had an error 1241 with their first patient today. Upon reboot, the same error repeated and it seems as though they can't use the device. They don't have any patients scheduled for the rest of the day. Can we look into data? | Installed replacement CCA in accordance with currently accepted practices per SWO-02747. | Right knee CCA was failing to even turn on when Ekso powered up. | The cause of the error thrown by the BIT routine was narrowed to a faulty joint printed circuit board assembly. The BIT routine correctly prevented the use of the device in this condition. The board assembly was replaced. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3708 | unit #234 - Neurojoint Centro Sinapse - strange noise | closed | duanee | Pick an option | fixed | 2017-05-30 07:47 | 2017-05-30 | 123792 | 234 | Centro Sinapse - Spain | Tania Romeu | neurologia@centrosinapse.com | 34981175859 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Workmanship | customer reported that they had a problem with the Ekso a week ago and since then it¬¥s doing a strange sound. update from customer June6th: The first time we heard the sound we were doing squats. It only appeared at that moment, we did not hear it while we walked. However, later we could hear it when we were making a turn, but we do not notice anything different with walking. Right now we do not hear the sound while we make the turns; however, we hear a new type of sound (as a friction) while walking, as you can hear in the video. Besides, one day we noticed burnt out smell (something exceptional anyway). We are not sure but we think the sound may come from the left hip. When we want to place hip width, we notice a blocking on the left hip at the time we move it to flexion, so we have to lift it up by hand to set the values. In the same way, this limits the other hip's flexion. We cannot perform the manual inclination to stand up; it has to be done automatically. The values Ekso gives us when we make squats are not correct, since they don't match to what we observe. Also, we can see how the iron rod moves when we make the separation or approximation of the width hip (in the place where the lock to block the hip is located). | tightened screws on rail | The screws holding the linear rail loosened slightly with use leading to an unusual sound when operated. | The source of the reported noise was isolated to loose screws holding one of the linear rails in one of the actuator assemblies. The complaint database was scanned for similar issues and two were found (#1676 and #2627) from 2014 and 2015, respectively. There is no evidence in the database of any systematic workmanship issue. No further action is necessary. (D. Edmonds ‚Äì 24 Jan 2018)
3703 | Lower Shank | closed | duanee | Adjustments - Leg | fixed | 2017-05-24 21:52 | 2017-05-12 | 1041590 | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Wear | Customer reported a clicking sound in the ankle. Issue: - Observed by Hunter while on a FSV - Damage to the lower shank from the loose rocker roller. "Discovered housing of right lower shank had been damaged due to jamming up against broken rocker roller (see photo). Housing (or lower shank) will require replacement." Hunter Rayner 5/12/17 | Rocker and slider assembly replaced per SWO-02697. | The pin that connects the two rollers in the rocker assembly sheared off at the base of the roller on the side closest to the cover. | The clicking from the ankle assembly reported by the customer has been isolated to wear on a pin in the assembly. The rocker and slider assembly has been replaced on this unit, and it successfully passed the testing following service. Previous complaints have been searched indicating that this is an uncommon occurrence. No further action is necessary. (D. Edmonds ‚Äì 30 Jan 2018)
3702 | Step Aborts | closed | duanee | Foot Sensor | fixed | 2017-05-24 21:37 | 2017-05-11 | 1041590 | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org | 509-662-1511 | usa | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | Issue: The customer reported some aborted steps after replacing the foot sensor. | Walked PT through a manual foot cal and the issue hasn't been present since cal was passed. | Customer indicated that, after replacing the foot sensors, there were aborted initial steps for the right foot with a right affected patient. He also indicated that, when doing manual foot sensor calibration, he has had to squeeze very hard with both hands on the right toe sensor to get it to register properly, both for the newly installed sensors and the previous pair. When I manually calibrated the foot sensors, all sensors registered fine with a simple, one-handed squeeze. Neither I nor the PT who tried the Ekso were able to reproduce the step abort when testing the Ekso ourselves. | Investigation performed previously (see #3431)
3699 | unit #260 - BG Hamburg - Pulse/unknown error | closed | duanee | Customer Use | fixed | 2017-05-24 13:12 | 2017-05-24 | 214172 | 260 | BG Klinikum Hamburg - Germany | Sarah Scheerer | s.scheerer@bgk-hamburg.de | +49 40 7306 ‚Äì 3671 / 3673 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | Investigation Needed | Customer Use | customer tried to create PW with first Pulse Login, but always received an 'unknown error' (see attached) update customer uses Internet Explorer Version 11.0.9600.18665 And she received an certificate error for the pulse login page, see attached, but could enter the page after clicking continue and now wondered if there might be a problem with their firewall? | Juliana resolved the issue with the customer with a phone call. | Customer had a technical issue on their computer. | The cause of this issue was isolated to an IT issue on the customer's side. There is no associated issue with the device or Ekso Pulse. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3698 | unit #152 - BGU Duisburg - noise in right knee | closed | duanee | Pick an option | fixed | 2017-05-24 10:27 | 2017-05-24 | 645008 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Workmanship | customer reported a rattle noise in the right knee | soldered pins on connector, see pics. | Amp board was delivered to Ekso without properly soldering the phase connections onto the board. | A technical assessment was performed the Engineering team isolating the cause of the issue to soldering workmanship on the amplifier PCA. Other amplifier boards in inventory were inspected for similar issues. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3693 | unit #112 - Heliomare - broken walker | closed | duanee | Cane / Walker / Crutch | fixed | 2017-05-23 11:33 | 2017-05-05 |  | 112 | Heliomare - Netherlands | Ronald Blok | rb@orthomedico.nl | +31(0)6 119 57 508 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | customer reported that walker is instable and lock opens while walking | Customer received parts. Fedex Tracking: 779214809840 Delivery: 779214809840 signed by: V.Elde | With continued use, the load from the pin that locks the walker in the open position on the slot in the plastic ring is able to cause the ring to fail or cause the pin to not seat properly in the slot. | n/a
3691 | unit #125 - Technomex - issue with walk angle | closed | mblaschke | Adjustments - Ankle | fixed | 2017-05-22 15:35 | 2017-05-22 | 1529821 | 125 | Technomex - Poland | Jakub Maliszewski | Jakub Maliszewski PHU Technomex Sp. z o.o. <jmaliszewski@technomex.pl> | +48 32 40 10 350 w. 38 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Determined Unknown | distributor reported with 0 on ankle it‚Äôs like they start write relative angle -3. Then when they set it to negative 3 it looks like 0. Torso test did not help to solve the issue | replaced amp board | The amplifier board was replaced resolving the issue, however, the board that was removed is not available for further analysis. | A risk assessment was performed after discussing the issue, and the amplifier board in the device was replaced. Current and past complaint files were searched for similar issues. This is an uncommon device issue. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3689 | Multiple Errors #1241 RH | closed | duanee | Software | fixed | 2017-05-18 23:55 | 2017-05-18 |  | 307 | RIC | Kristen Hohl | khohl@ricres.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | Issue: Just wanted to follow up with my voicemail regarding our Ekso safe mode this morning. I got the message software issue, reboot required and contact ekso bionics with an error !#1241-RH. The reboot didn't solve it and we are still getting the message after turning on the ekso. We were able to safely transfer out of the device but certainly would like to get this fixed. | Amp board replaced per SWO-02723. | Problem with the communication lines on the amp board. | The error reported (Error 1241 ‚Äì "Software Error Call Ekso Bionics", "Joint board offline") indicates a fault condition in which the joint board (PCA) is not communicating correctly. The Built-In Test (BIT) routine responds to this event by transitioning to Safe Mode, sending the warning message to the user, and preventing the continued use of the device while the communication error is present. The amplifier board was replaced on this unit, and the unit was successfully put through the qualification test routines. The complaint database was reviewed for other instances of similar issues. Based upon that review, joint board failures will be monitored closely. No further action is necessary to address the issue experienced by this customer. (D. Edmonds ‚Äì 06 Nov 2017)
3682 | Unit #265 - Centro de Medicina de Rehabilita√ß√£o Altcoit√£o (SCML) - charger makes resistance | closed | duanee | Battery - Charger | fixed | 2017-05-16 12:04 | 2017-05-16 |  | 265 | SCML - Alcabideche | Tiago Teixeira | tiago.teixeira@scml.pt | +351 930411889 | Portugal | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | The problem of the charger that makes resistance in the placement of the battery seems to be related to the pins. | SWO-02715 replaces battery charger SWO-02811 sending correct EU power cords for new chargers | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 7/19/2017)
3681 | UNIT #185 - Technomex/Ergo Hestia - strange walking | closed | duanee | Foot Sensor | fixed | 2017-05-15 15:39 | 2017-05-15 |  | 185 | Technomex - Poland | Jakub Maliszewski | Jakub Maliszewski PHU Tecjmaliszewski@technomex.pl | +48 698 609 922 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Michal reported an issue with the unit today, "strange walking". Michael advised doing the foot sensor calibration which solved the issue (for the moment). | Distributor received the parts and replaced SWO-02708 | Known wear component | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Brian Hanson on 5/24/2017)
3677 | unit #235 - Rolling Ball/Garry McCabe -hip abduction changing | closed | duanee | Adjustments - Hip | fixed | 2017-05-15 13:37 | 2017-05-05 |  | 235 | Rolling Ball - Ireland | Jane Evans | jane@rollingball.ie | +353 87 224 6287 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | customer reported that the they have to re-adjust the abduction several time because it is changing while walking. | replaced parts, swo 2899 | It is not possible to isolate the root cause of this issue at this point. The components that were replaced in the hip abduction/adduction assembly on the device are not available for further analysis. | The customer issue was addressed in this case by replacing components in the hip abduction/adduction assembly and testing the unit to verify the reported behavior did not return. It is not possible to narrow the root cause further at this point. No further action is necessary. (D. Edmonds ‚Äì 18 Oct 2017)
3676 | unit #176 - Technomex/Konstancin - broken joint screw | closed | duanee | Actuators | fixed | 2017-05-12 13:55 | 2017-05-12 |  | 176 | Technomex - Poland | Micha≈Ç Janus | <mjanus@technomex.pl> | +48 694 482 726 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor | Investigation Needed | Wear | technomex reported a broken joints screw on the left hip | Sent parts to distributor, SWO-02706 | Known fatigue failure in joint cap screw. | An assessment of risk was performed and the probability of occurrence recorded in the previous assessment ( 1.2 FMEA-768, 102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet, Rev. V) was re-evaluated based on field data. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3675 | Error 136 and continuous errors | closed | duanee | Pick an option | fixed | 2017-05-11 21:34 | 2017-05-11 | 761378 | 220 | Good Shepherd 3 | Alyssa Hauck | AHauck@gsrh.org | 484-523-0756 | USA | Slight redness at tibias and mild indentations on shins | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Issue: I am emailing to discuss the incident that occurred with one of our ekso units today. On 5/11/17 at 1:45, one of our most experienced ekso users was ambulating using the ekso when the ekso machine went into error mode with his left leg fully locked into extension and right leg locked into flexion. The error code read !#136 LH "ekso motor hot". I acknowledged the error code and then it started shooting consistent errors about the foot sensor. I tried to hit the yellow button many times to lock into full stand still but the machine did not go into stand still and continued to remain in error mode. The patient nearly fell and luckily one of my coworkers was able to bring a chair in time to slowly lower him to the chair with his legs locked in that position. The patient has good upper extremity strength and was able to assist in maintaining his balance before we could get a chair but I fear if it had been a less experienced user that the patient likely would have fallen in this situation. The patient denied injury during the incident but I did notice some slight redness at both of his tibias and mild indentations on his shins due to the pressure from the ekso tibial pads when locked into that position while losing his balance. From my understanding, Sue has already contacted you regarding this issue. I feel that this is an urgent issue regarding patient and therapist safety because other patients witnessed this event and are now fearful to use the device and therapists are also anxious to use the device. While using the loaner device over the past two weeks, our patients walked very well and we did not have issues. Additionally, our other device appears to continue to have inconsistent performance during L swing phase with the L LE exhibiting "catches" during swing with multiple different patients which affects quality of gait and patient care. Please advise how you would like to resolve these issues as many patients are complaining about the inconsistency in performance and overall quality of the devices over the past few weeks. | Fan cable replaced per SWO-02742. | The fan cable jumper that connects the fan to the hip CCA was sound to have an intermittent short. | The cause of this issue was isolated to a fault in the motor fan cable. The fan provides airflow for convective cooling of the motor. This is consistent with the reported behavior and led to BIT routine to generate a warning and triggering an error as the detected motor temperature began to reach an unacceptable level. The motor fan cable was replaced, and the unit was successfully put through the standard system tests. The reported behavior with respect to the (yellow) Stop Button is consistent with the expected behavior of the device, although different than the expectations of the customer. The behavior of the device in response to pressing the Stop Button in an error case is necessarily different than when the Stop Button is pressed during normal operation. No further action is necessary. (D. Edmonds ‚Äì 03 Oct 2017)
3673 | unit #251 - RollingBall/EliteGym - broken foot sensor | closed | mblaschke | Foot Sensor | fixed | 2017-05-11 14:20 | 2017-05-11 |  | 251 | Rolling Ball - Ireland | Jane Evans | Jane@rollingball.ie <jane@rollingball.ie> | +353 87 2246287 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Every 3/4 steps with the left leg, it doesn't complete a full step and lands just forward of the standing position. This has been happening for a few days now. | distributor replaced parts - SWO-02705 | Known wear issue | Investigation performed previously (see #3431)
3672 | unit #145 - Royal Bucks/UK | closed | duanee | Foot Sole | fixed | 2017-05-10 14:14 | 2017-04-25 | 231215 | 145 | Royal Hospital - UK | Eva Sobonova | Certified Physiotherapist eva.sobonova@royalbucks.co.uk | +44 1296 678880 | United Kingdom | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | No investigation needed - see Justification | Wear | customer reported that the rubber on both insoles is peeling off. | replaced foot sensors, swo 2962 | Known wear item. | INV-0006
3671 | unit #178 - Technomex/Kamien Pomorski - clicking noise in knees | closed | duanee | Pick an option | fixed | 2017-05-10 10:14 | 2017-05-04 |  | 178 | Technomex - Poland | Lech Garasimowicz | lgarasimowicz@technomex.pl | +48 694 483 126 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | Investigation Needed | Wear | distributor reported a clicking noise in the knees | added grease to bearings | The lubrication of the bearings changed over time. | The noise reported by the customer/distributor was found to come from bearings in the knee assembly. Grease was added to the bearings. This resolved the noise issue. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3670 | unit #177 - Technomex/Kamien Pomorski- left foot rotating | closed | mblaschke | Foot Assy | fixed | 2017-05-10 10:08 | 2017-04-14 |  | 177 | Technomex - Poland | Lech Garasimowicz | lgarasimowicz@technomex.pl | +48 694 483 126 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | distributor reported that left foot was rotating on device | distributor replaced parts - SWO-02637 | It is not possible to narrow down the root cause without the original parts that were replaced by the distributor. | An investigation is been performed previously for this issue (see #3507).
3669 | unit #046 - Instituto Guttmann - battery not recognized | closed | duanee | Battery | fixed | 2017-05-10 09:43 | 2017-05-08 |  | 46 | Guttmann - Spain | Silvana Andrade | gestiodeserveis@guttmann.com | +34 93.497.77.00 ext. 3199 | Spain | No injury | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Customer Use | customer reported that one battery is not working anymore. It does not get recognized from the ekso anymore, not matter if they connect it left or right. (error message: "Falta bat. Izquzierda" and "Falta bat. Derecha"). Battery serial: 1B serial-013356201500008F. | Directed the customer to straighten pin on battery connector | Customer accidentally bent pins on battery connector | Investigation performed previously (see #3566)
3668 | unit #046 - Instituto Guttmann - issue with left battery latch | closed | mblaschke | Other | fixed | 2017-05-10 08:43 | 2017-05-08 |  | 46 | Guttmann - Spain | Silvana Andrade | gestiodeserveis@guttmann.com | +34 93.497.77.00 ext. 3199 | Spain | No injury | Same as contact listed | ekso 1.1 | FALSE | Unassigned | Minor | Investigation Needed | Determined Unknown | customer reported that it is hard to unlatch the battery on the left. It is only possible with help of a screw driver. | Disassembled and reassembled battery latch | The latch was rebuilt to address the customer's issue. It is not possible to narrow down the root cause of this issue further at this point. | The latch on this battery was disassembled and rebuilt, which resolved this issue. The complaint file database was also reviewed for similar issues. This is an uncommon issue. No further action is necessary.
3667 | unit #185 - Technomex/Ergo Hestia - broken screw on ankle | closed | mblaschke | Ankle | fixed | 2017-05-10 08:22 | 2017-05-09 |  | 185 | Technomex - Poland | Micha≈Ç Janus | <mjanus@technomex.pl> | +48 32 40 10 350 w. 38 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Materials | distributor reported a broken screw on the sole of the left ankle | distributor replaced parts - SWO-02695 |  | Previous investigation completed (INV-0001)
3664 | unit #144 - BGU T√ºbingen - error #21 LH | closed | mblaschke | Other | fixed | 2017-05-08 09:48 | 2017-05-04 |  | 144 | BGU Tubingen - Germany | Annette Werz | awerz@bgu-tuebingen.de | +49 7071 6063237 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Determined Unknown | customer received sensor error !#21 twice while walking with patient. | calibrated joint | The issue appears to have been resolved by recalibrating the joint. We have not been able to reproduce the issue. | The risk assessment has been completed and the appears to have been resolved by recalibrating the joint. The complaint database was reviewed for similar issues in the field. No further action is necessary. (Duane Edmonds ‚Äì 05 July 2017)
3663 | Ankle rocker roller broken | closed | duanee | Adjustments - Ankle | fixed | 2017-05-04 18:32 | 2017-05-03 | 1041590 | 154 | Confluence Health | Colton Frazier | Colton.Frazier@confluencehealth.org | Attn.: (509) 665-6156 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Undetermined | Issue: Unfortunately there is another issue. When replacing the footplate sensors, on the right side, this little guy fell out (see ‚ÄúAnkle2‚Äù and ‚ÄúAnkle3‚Äù). While doing a bilateral comparison I was able to find where it belongs (see ‚ÄúAnkle1‚Äù). It looks like it was sheared off so I cannot just place it back in its corresponding location. It does not seem to be giving us any issues and to be honest I‚Äôm not sure what its function is. Appears to be something with the spring but I‚Äôm not entirely sure. Is this something that needs to be addressed ASAP or should we wait until the next scheduled maintenance check? | Replaced rocker according to currently accepted best practices. Work completed per SWO-2697. | Roller broke off of Ankle Rocker Assembly. Cause Unknown. | This issue appears to be an isolated event. The complaint database was searched for similar issues, and this is the only occurrence of this failure in the database. The rocker was replaced on this unit, and there have been no further issues of this type. No further action is necessary. (D. Edmonds ‚Äì 24 Jan 2018)
3662 | Cracked Hip Housing | closed | duanee | Actuators | fixed | 2017-05-04 17:42 | 2017-04-20 | 727669 | 207 | Good Shepherd - OP | Alyssa Hauck | AHauck@gsrh.org | 6107763247 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Design | Crack in hip housing, found at customer site. | Replaced hip housings per SWO-02666. | Stress concentration at crack origin. | ] No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) #832 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/16/2017)
3661 | Errors 168, 121 and 132 | closed | duanee | System | fixed | 2017-05-03 23:21 | 2017-05-01 | 100135 | 49 | Able Bionics - Canada | Taylor McLuhan | ablebionics@gmail.com | 1(519)852-2253 | Canada | No injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Issue: Could you check data for Able Bionics - Ontario (Unit 49) They had several errors and seems like they may be potentially down. Yesterday with one patient they had errors 168, 121 LK and 132. This morning they walked a PTA and had the same errors as well as a fourth error, she couldn't remember what the fourth one was. They have a patient at 2 PM today. Let me know if there is any data! | Knee motor was replaced. | Failure in the left knee motor. | The errors reported (Error 121 "Motor Error Ekso in Safe Mode" "Elmo motor failure bit set", Error 132 "Motor Error" "Motor on timeout", and Error 168 "Step Aborted Ekso in StandStill" "invalid weight shft") are thrown by the Built-In Test (BIT) system self-check routine on the device. The response of the BIT routine is to place the device in Safe Mode, in which the motors are disabled and the knee brakes are locked to prevent undesired behavior due to errors with the motor or motor amplifiers. The BIT routine performed as designed in this case, and alerted the operator of the problem with the left knee motor and correctly prevented the continued use of the device. The left knee motor assembly was replaced, and the unit passed exit tests successfully. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3660 | unit #235 - Rolling Ball/Garry McCabe - issue with right hip | closed | mblaschke | Pick an option | fixed | 2017-05-03 14:28 | 2017-05-03 |  | 235 | Rolling Ball - Ireland | Jane Evans | jane@rollingball.ie | +353 87 224 6287 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Design | distributor reported There is oil coming from the hip motor on the right leg (ie the users right leg side) of the Ekso. There is also a black oily substance coming out of the airvents. This is accompanied by a smell of burning as though the device is over heating. | no issue found, customer had only noted grease at joint cap | It is possible for a small amount of grease to come out device during use. This does not affect the functionality of the device. | Discussed source of grease/oil and associated risk at with 20 June 2017 meeting (D. Edmonds, C. Kealoha, K. Strausser, L. Salska-McNeil). The grease/oil does not affect the functionality of the device. Reviewed complaint file database for similar or related issues. No further action is required.
3658 | clicking noise from left knee | closed | kmettler | Actuators | fixed | 2017-05-03 08:58 | 2017-05-03 | 470075 | 123 | Ekso Bionics EU | Michael Blaschke | mblaschke@eksobionics.com | +49 173 155 9837 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | Reported during a demo that a clicking noise was coming from the left knee | replaced joint board 102286, swo 2674. | Faulty component on joint board 102286. The joint board 102286 had a faulty component which caused the knee brake to not engage/disengage in a continuous manner. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437V ‚Äì Risk Management Spreadsheet (FMEA Tab) #1-C for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 5/16/2017).
3656 | Swing Leg Block | closed | duanee | Foot Sensor | fixed | 2017-05-02 22:33 | 2017-04-20 | 307730 | 175 | Barrow Neurological Institute | Nathan West | Nathan.West@DignityHealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Issue: Initial issue reported during SA traing - Hi team, We are training group #2 on smartassist at Barrow's. Similar to last week, we are seeing the swing leg get blocked on the right (as if it is in stance) on a frequent basis. I did a foot calibration last Friday before I left. ID number is 30 and we are walking now. Can we get data on this? Thank you, Darrell | Both foot sensors replaced per SWO-02707. | This was determined to be due to noise coming from the foot sensor that caused the unit to believe that it had not lifted the leg off the floor. | The risk associated with the failure of the foot sensors (this particular failure mode and others) has been evaluated. With the current mitigations in place, there is no significant safety risk associated with these failures. The complaint file database was reviewed to re-evaluate the frequency of occurrence based on current field data. No further action is necessary. (D. Edmonds ‚Äì 21 Sep 2017)
3655 | Leg Rotating | closed | duanee | Adjustments - Leg | fixed | 2017-05-02 20:57 | 2017-05-02 | 318857 | 175 | Barrow Neurological Institute | Luis Escalante | lues0424@yahoo.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Design | Issue: When our device returns, while they are at it could you ask them to check the left leg as well. I'm just wrapping up a treatment where the left leg was internally rotating on its own during ambulation. We stopped reset the leg and it did it again. | Collar replaced per SWO-02707. | Caused by an older revision of the collar. | This issue (rotation of the leg during walk) has been reviewed and the root cause was determined to be that the adjustment screw(s) bottomed out and could not, as a result, provide adequate clamping force to hold as expected. This has been addressed and the shank collar on this unit has been replaced with the updated design. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 02 Nov 2017)
3654 | Leg Buckling | closed | ckealoha | Pick an option | duplicate | 2017-05-01 21:53 | 2017-05-01 |  | 142 | Burke | Helma Zanders | wzanders@Burke.org | Phone #: 914-548-2857 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Unassigned |  | Undetermined | Issue: We were having some troubles today with EKSO. The robot kept buckling when we freed one leg (on both L and R aff) and the swing assist was higher than usual and the leg kept bouncing on the right. This happened with our patient number 64. We sat her down to shut the robot off and turn it back on, but the problem remained. Also, the battery on the left leg is low, which seems pretty fast to us since we just got it back a few months ago. I hope you can help us out! | (Explain actions taken to resolve this issue) | (Required for complaints)
3653 | Possible ESD Shock | closed | duanee | System | worksforme | 2017-04-28 17:42 | 2017-04-28 | 10099 | 306 | United Hospital | Josh Tanner | joshua.tanner@allina.com | (651) 241-8000 | USA | Electric shock to PT | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Determined Unknown | Issue: I just wanted to discuss an event that happened during my training today. I was participating in a lunch and learn with a PT that missed due to illness yesterday. I was in the Ekso demonstrating various gait deviations while another PT held onto the controller. We were almost finished up when the controller sent a severe shock through the PT holding on.She stated that it felt like she was electrocuted versus any type of static shock. It was clearly very painful for her and she stated that it went through her entire arm up to her shoulder. This shock created a device error that forced us to perform a safe mode sit and reboot. The conditions (regular walking) and environment (tile flooring) were normal. I spoke with Josh about this and he explained that it was electrical static discharge build up from the Ekso. It sounds as though this has happened on occasion in the past but this was more severe. The therapist was unable to participate in the following session reporting that her arm was still bothering her. She participated in the final session of the day but was very apprehensive about touching the device or controller. I am wondering if any of you have experienced this in the past and how you have responded. | UI was replaced. | Unable to find or replicate issue. | The device involved in this event (serial number: EKSO-0306-021611001) was put through a detailed inspection (work order: SWO-02670) looking for possible causes of a shock hazard to an operator or patient. No issues with the device (damaged or incorrectly wired electrical connections, damaged or improperly assembled or reworked printed circuit boards, etc.) were found. We were also unable to reproduce the reported behavior after discussions with the customer about the scenario in which the event occurred. The complaint database was searched for similar issue. There are some instances of the control pad screen (see INV-0008) responding to static electricity build up, but none of shock to a user by other causes. Based on this, the control pad on this unit was replaced proactively. No further action is necessary. (D. Edmonds ‚Äì 24 Jan 2018) inspection photographs are available on the network drive at \\exo-file02\Berkeley Bionics\Customer Relations Tools\TRAC Pictures and Videos\3653\download.zip
3651 | right hip cracked | closed | duanee | Actuators | fixed | 2017-04-28 09:03 | 2017-04-28 | 1283529 | 233 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | I found a cracked right hip during service | replaced hip housings, swo 2668 | material fatigue at a rate higher than expected based on device load model | No new investigation needed as previously investigated and deemed not a safety rick or a reportable event. See FMEA [1.2 FMEA-390-C] for original investigation information and mitigations put in place. (Evaluation done by Brian Hanson on 3/6/2017) (copied from #3484)
3648 | Unit #198 - EMAC/Negrar - right toe pressure | closed | duanee | Foot Sensor | fixed | 2017-04-27 12:24 | 2017-04-26 | 743287 | 198 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39.347.54.34.417 | ITALY | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | This Ekso shows the error #41. Emac tried to do a foot calibration but it doesn‚Äôt works. Is it necessary to replace now the foot sensors. | Distributor replaced foot sensors per SWO-2661. | Foot sensors are known wear component | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. Error codes 40 - 43 indicate foot sensors out of range. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Brian Hanson on 5/24/2017)
3647 | unit is walking strange | closed | duanee | Pick an option | fixed | 2017-04-27 08:10 | 2017-04-26 | 323152 | 209 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 4010350 x 54 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Determined Unknown | customer reported the patients feel a bit bent backwards while walking | Micheal performed torso calibration per SWO-02660. | Torso requires re-calibration. | The behavior reported was found to be caused by a change in the torso calibration that caused the device to control to a non-ideal torso angle. The unit was run through the (re)calibration routine, which resolved the issue. Unfortunately, it is not possible to determine what caused the device to lose the calibration originally set at the factory. No further action is necessary. (D. Edmonds ‚Äì 16 Nov 2017)
3645 | Issues w/ 1st step | closed | kmettler | Foot Sensor | fixed | 2017-04-26 23:09 | 2017-03-21 | 392536 | 161 | Glenrose - Canada | Vickie Buttar | Vickie Buttar <Vickie.Buttar@albertahealthservices.ca> | +1 (780) 735-7999 | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Issue: They were having similar issues w/ the foot sensors as they had one month before at an ITW w/ our Ekso PT. the left foot was on the ground and we wanted a right step - controller read "left stepping" They continued to have issues taking the right step | Send out replacement parts | Known wear item. Will not be getting this part back for proper RCA | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437V ‚Äì Risk Management Spreadsheet (FMEA Tab) #364-C for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 4/27/2017).
3643 | Issues w/ 1st step | closed | kmettler | Foot Sensor | fixed | 2017-04-26 22:38 | 2017-04-26 | 328483 | 130 | Sunnyview Rehab | Patti Valenza | 518-382-4528 |  | USA | No injury | Same as contact listed | Pick an option | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | Issue: This customer was having issues initiating the first step w/ the customer loaner device. | Parts replaced via SWO-2477 | Known wear item, not receiving these parts back for proper RCA | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 4/27/2017)
3642 | Faulty UI | closed | duanee | UI - Controller | fixed | 2017-04-26 21:43 | 2017-04-26 | 3135 | 306 | United Hospital | Matt Knowlton | MKnowlton@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: Ekso PT at a training sent through that the controller was faulty, they had a couple of stand stills during walking. Errors: 1111 and 1112 and are indicative of a UI failure. | Sent new parts via SWO-2659 | The control pad was replaced by the customer resolving the issue. However, the faulty control pad is not available for further analysis. | Errors 1111 and 1112 indicate issues with the device serial bus (break detect and framing error, respectively) used to communicate from the control pad to the main device controller. The errors are thrown as part of the device Built-In Test (BIT) routine which responded as designed and stopped the use of the device. (D. Edmonds ‚Äì 14 Sep 2017)
3639 | cracking noise from right knee | closed | duanee | Other | fixed | 2017-04-26 11:30 | 2017-04-26 | 1582940 | 129 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 4010350 x 54 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | I found a knee making a cracking noise while walking. The reason seems to be a gap between both parts of the knee housing. It looks like there is a gap of 1mm on each side so the housing is moving while walking. | Knee housing replaced per SWO-02654. | The pins that hold the thigh structure to the knee housing loosen on certain units. | The issue was isolated to specific pins in the knee assembly that became loose over time (see Root Cause Analysis). The assembly was serviced (see Resolution Description for details), and the device was tested prior to the return of the unit to the customer. No further action is necessary. (D. Edmonds ‚Äì 26 Oct 2017)
3637 | Foot Sensor issues | closed | kmettler | Foot Sensor | fixed | 2017-04-25 23:33 | 2017-04-25 | 995745 | 154 | Confluence Health | Frazier, Colton | Colton.Frazier@confluencehealth.org |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | N/A - Identified known wear item | Wear | Issue: I‚Äôm emailing you today because of an issue we are having with the Ekso. As a result of the issue, we are having to manually re-calibrate the foot plate sensor every day. We are doing this because of lag-steps while walking with patients in a variety of different settings. I‚Äôve found that re-calibrating the foot plates manually resolves the problem for a few hours but it has become time consuming and cutting into patient-care times. To me it seems that the right forefoot plate sensor is not working. When attempting to re-calibrate, this particular sensor requires quite a bit of force, respectively, in order to pass the test. Let me know if you have any questions! | Sent new parts via SWO-2653 | Known wear item | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 4/27/2017)
3633 | Device going into standstill | closed | kmettler | UI - Controller | fixed | 2017-04-25 18:11 | 2017-04-25 | 342779 | 173 | Warm Springs | Salina Morgan |  | 210-216-4537 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Issue: Hey Guys, Can we pull data on unit 173? PT is telling me the device will go into standstill without being commanded‚Ä¶ I‚Äôm thinking the only thing that can make this happen is a faulty UI‚Ä¶ Thoughts? | Sent new UI via SWO-2649 | UI known to wear over time. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 5/5/2017)
3632 | DuplicateUnit #233 - To Walk Again - left leg turns too far | closed | kmettler | Adjustments - Leg | duplicate | 2017-04-24 15:00 | 2017-04-22 | 1275106 | 232 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | Left leg turns too far in adduction while setting ABD are more than +3..., reboot required after missing right batterie while batterie is attached properly... | Duplicate to #3631 | Duplicate to #3631 | Duplicate to #3631
3631 | Unit #232 - To Walk Again - left leg turns too far | closed | mblaschke | Adjustments - Leg | fixed | 2017-04-24 14:52 | 2017-04-22 | 704973 | 232 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Left leg turns too far in adduction while setting ABD are more than +3..., reboot required after missing right batterie while batterie is attached properly... | replaced hips (SWO-02669) | hip housings broken, see ticket 3652 | Investigation performed previously (see #3662)
3629 | Extra Lateral movement in hips | closed | kmettler | Adjustments - Hip | fixed | 2017-04-20 23:14 | 2017-04-20 | 797088 | 220 | Good Shepherd - IP | Alyssa Hauck | AHauck@gsrh.org | (610) 776-3247 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Design | Issue: I am emailing to let you know that both of our devices continue to have issues with the hip abduction holding its value during the sessions. One of our patients nearly fell today due to severe clipping at the feet despite multiple adjustments and another very experienced walker was having a lot of difficulty walking efficiently today due to the hip issue and was disappointed in his session due to the poor quality of his gait mechanics. I think it would be best if we ship both units out for maintenance at this point because there seems to be an underlying issue that did not get resolved in the maintenance we received yesterday for the devices. I have CCed Christina Wood on this email as well who was one of the therapists who noticed a continued issue with the device and a feeling that one of the hips was "dropping down" when walking as if it was loose. Please advise on how you would like to proceed with this issue. | Issue was originally found per SWO-2623. Replaced hip housings per SWO-2667 | Hip housing cracked as per normal design failure. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437V ‚Äì Risk Management Spreadsheet (FMEA Tab) 390-C for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 5/16/2017).
3627 | Unit #141 - EMAC/Istituto S.Stefano- T-wrench not working | closed | kmettler | Tools | fixed | 2017-04-19 12:48 | 2017-03-20 |  | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39.347.54.34.417 | ITALY | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | One of the T-Tool is not working properly. When you use it to loosen the screw, it is not able to do it. | Mailed replacement items per SWO-2648 | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler on 4/24/2017)
3625 | Plug of the manual UI is faulty | closed | kmettler | UI - Crutch Interface | fixed | 2017-04-19 10:25 | 2017-03-09 |  | 48 | Bad Aibling - Germany | Denny Muller | dmueller@schoen-kliniken.de | +49 8061 903 1679 | Germany | No injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | the cable that is to say the plug of the manual control module is faulty | Sent replacement items per SWO-2633 | Known wear component - plug and cable gets worn due to flexing over time. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 5/5/2017)
3622 | Repeating Error #51 | closed | duanee | Pick an option | fixed | 2017-04-17 21:57 | 2017-04-11 | 3000 | 309 | Swedish Covenant | Matt Knowlton | MKnowlton@eksobionics.com |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | On the 2nd day of ITW the device is throwing error !#51 LH (repeating). This is requiring a joint calibration, the Ekso will pass the calibration but the error continues to return after each successful cal. | Entire hip housing with new soft pot was used to replace the customers soft pot in the field. | A contamination in the adhesive between the soft pot and the hip housing caused the soft pot to delaminate. | The potentiometer was removed and examined. Contamination between the potentiometer and the housing led to the potentiometer reading incorrectly. The BIT safety checks performed as designed and threw the appropriate system error. This has been discussed with the manufacturing and service teams. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3621 | Knee stuck | closed | duanee | Adjustments - Leg | fixed | 2017-04-17 17:38 | 2017-04-14 | 527211 | 115 | NRH - Med Star | Katie Seward |  | 908-770-3809 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Materials | Issue: While walking a patient today the Knee froze and prevented the patient from taking any further steps. The PT was able to sit the patient without issue, while the device is powered on the knee is extremely hard to move, the PT was finally able to straighten, when the device is off the Brake Release doesn‚Äôt work. Sounds like the issue we had @ Burke and Confluence. Loaner is going out today and will arrive on Monday. Customer device will be back late next week. | Knee assembly replaced per SWO-02639. Old knee assembly will be used for further testing. | Issue with the joint board in the knee. Causing the brake to turn on and off continuously. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3619 | Error 25: LH Pot/Encoder Mismatch during training | closed | duanee | Encoder / Potentiometer | fixed | 2017-04-14 14:44 | 2017-04-11 | 3792 | 309 | Swedish Covenant | Mark Esper |  | 773-878-8200 ext. 6778 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | During training at customer site (Chicago), it was found that unit would throw Error 25 when left hip was moved without anyone wearing unit. | Replaced entire left hip and don-doff assembly with new parts except for motor and covers. | Particulate was found between the softpot and adhesive backing. | Particulate was found in the potentiometer was causing readings that were at odds with the motion of the device. This is what caused the BIT routine to generate the warning by triggering an error. The BIT safety checks performed as designed. (D. Edmonds ‚Äì 14 Sep 2017)
3618 | Unit #258 - Pasori - hip mechanism blocked | closed | duanee | Adjustments - Hip | fixed | 2017-04-13 13:31 | 2017-04-13 |  | 258 | CRF Pasori - France | Stephane Antonin | kine@pasori.fr | +33 386 28 4389 | FRANCE | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Customer Use | They cannot use the exo device anymore. In fact sometimes there is a woman for whom they need to enlarge the mechanism a lot until 15-20 max. Last days, they noticed that it was getting even harder to make it smaller after her using it. Now the hip mechanism is blocked and cannot get smaller anymore. | Sanded latch pin to remove any burrs caused by customer use issue. The latch pin conformed to specifications after service. | Customer attempted to move the hips larger than what is physically possible for the device. When this occurred, the hip width limit pin was damaged. This damaged pin jammed in the assembly and prevented the hip width to be changed. | Completed a review of 102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet (rev. V) and updated Risk Level and Risk Level Explanation, and reviewed other complaint database for similar tickets. Assessment is consistent with details in previous tickets, and Essential Performance and Basic Safety are maintained. (Updated on 15 June 2017 by Duane Edmonds)
3616 | Unit #252 - Ostracon/Animus - problem left leg | closed | kmettler | Adjustments - Leg | fixed | 2017-04-12 14:32 | 2017-04-08 | 66197 | 252 | Ostracon Ltd - Greece | Dimitris Aslanidis | physio@animus.com.gr | +30 2410 851851 | GREECE | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Customer Training | When (adaptive mode, pro step) the (left only)limb is at swing phase right before heel strike, the knee remains flex a bit (knee flex at zero position from settings menu of course). Only when the limb comes to stand phase and right before push of, comes to full extention. | Directed customer to use torso spacer | Customer was not using torso spacer due to patient not being well aligned in the ekso. See comments for additional info. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 5/5/2017)
3615 | Unit #252 - Ostracon/Animus - crutch interface out of work | closed | kmettler | UI - Crutch Interface | fixed | 2017-04-12 14:20 | 2017-04-08 | 66197 | 252 | Ostracon Ltd - Greece | Dimitris Aslanidis | physio@animus.com.gr | +30 2410 851851 | GREECE | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Crutch interface does not function, work | Sent replacement item per SWO-2618 | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler on 4/24/2017)
3614 | Joint board failure | closed | kmettler | Electrical Interconnect | fixed | 2017-04-12 10:01 | 2015-10-29 | 374420 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | 4920376883613 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | Joint board would spontaneously turn off during typical use | Replaced joint board per SWO-1519 | Found that a component on the board was damaged | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/21/2017)
3613 | Clicking in left hip | closed | kmettler | Actuators | fixed | 2017-04-11 16:43 | 2017-04-02 | 523521 | 181 | VA Bronx | Steven Knezevic |  | 718-584-9000 x3130 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Customer described a clicking sound coming from the left hip of the Ekso. Occurs most often with two specific patients. | Spacer and bearing sleeves replaced on both hips per SWO-02582. | Tolerance issue between the fit of the spacer onto the bearing sleeve. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 4/21/2017)
3611 | unit #243 - Step by Step/Barcelona - screws on ankle broken | closed | kmettler | Foot Assy | fixed | 2017-04-10 14:52 | 2017-04-10 | 397532 | 243 | Step-by-Step - Spain | Joan Romero | Joan Romero <jromero@fundacionstepbystep.org> | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Design | The screws of the left ankle are broken. I have attached to pictures to let you know the situation. I'm unable to change them and it seems that the sensor cables could be damaged. | replaced ankle assemblies, swo 2677 | Design issue. Refer to TRAC # 2568 for more details. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 785 and # 786 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3609 | Ankle axel flange failure | closed | kmettler | Ankle | fixed | 2017-04-10 10:55 | 2017-04-07 | 1554267 | 129 | Technomex - Poland | Jakub Maliszewski | Jakub Maliszewski PHU Technomex Sp. z o.o. <jmaliszewski@technomex.pl> | +48 32 40 10 350 w. 38 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | No investigation needed - see Justification | Design | Customer reported "left ankle gets pin out" | sent parts to distributor, swo 2600 | Flange on ankle axle cracked and broke off. Issue resolved by design change associated with CO-11750. | No new investigation needed as similar end effects previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 785 and # 786 for similar investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/5/2017)
3608 | thigh length changing | closed | duanee | Adjustments - Hip | fixed | 2017-04-10 10:47 | 2017-04-10 | 1529821 | 125 | Technomex - Poland | Jakub Maliszewski | Tecjmaliszewski@technomex.pl | +48 698 609 922 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Determined Unknown | "therapist report that tight length can expand during walking" | sent parts to distributor, swo 2601 | Distributor did not return part for proper RCA. Installation of new clamp, part number 102089, resolved the issue. | Reviewed 1.2 FMEA-513-C (102437 ‚Äì Ekso 1.2 Risk Management Spreadsheet, rev. V) and searched complaint file database for similar issues. The assessment does not need to be changed based on the field data we have to date. No further action is necessary.
3607 | Device initiated step | closed | duanee | Pick an option | fixed | 2017-04-07 00:17 | 2017-04-06 | 46502 | 286 | Providence | Jennifer Cowan | jennifer.cowan2@providence.org | 310-832-3311 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor | Investigation Needed | Customer Use | Issue: PT called from San Pedro, They were walking yesterday from 2:30-3:30 in ProStep + and the patients leg flexed and then it initiated a unwanted step. They said they weren‚Äôt able to put the device in standstill but error 159 popped up. session was between 2:30-3:30 | Multiple phone calls to customer resolved the issue. | Determined to be a customer use issue. | The reported issue was determined to be due to a slight misunderstanding of the use of the device in ProStep+ and the information conveyed by the error report from the device (Error 159). This error ("Move feet together while in Stand Still") is a warning that indicates that the feet must be moved closer together to prevent the device from eventually reaching a condition that would require the use of the device to cease temporarily. (The continuous current required to the hip motors in this condition is high because the motors are attempting in vain to pull the feet together in Stand Still against the friction from the contact of the feet with the floor.) The communication with the customer clarified the misunderstanding. No further action is necessary. (D. Edmonds ‚Äì 26 Oct 2017)
3606 | Issues w/ R Swing | closed | kmettler | Customer Use | fixed | 2017-04-06 20:31 | 2017-01-13 | 261111 | 23 | NeuroFit360 | Guy Romain | Guy Romain <guypt@neurofit360.com> |  | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor | N/A - Identified known wear item | Customer Use | Issue: The problem is that right leg does not complete swing and stomps down prematurely. This occurs in bilateral and right affected mode in straight paths indoors. We do not perform turns or walk about in curved paths with the device in walking mode with any client, keep in mind the video sent was an able bodied individual walking around in circles trying to perform enough steps for the calibration test. The right leg has never given any issues before, we have had issues with the left leg and replaced the foot sensors twice already. I'm not really sure what clinical or technical changes could fix this issue. Please advise as soon as you can. | Recalibration of foot sensors and legs with SWO-2597 | Foot Sensor calibration appeared fine upon arrival, test walked, no issue found under normal operating conditions. Re-calibrated foot sensors. No errors found on log for foot sensors. [August and Sept 2015 Error 25 LH PTEC; error 168 Step Abort]. Recalibrated legs. No issues. | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See D100117 - Risk Management Spreadsheet (FMEA Tab) # 269 for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/5/2017)
3605 | Knee varus in Ekso | closed | duanee | Customer Use | fixed | 2017-04-06 17:56 | 2017-04-07 | 261111 | 63 | NeuroFit360 | Guy Romain | guypt@neurofit360.com | 954-252-8020 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor | Investigation Needed | Customer Use | Issue: The below issue is with their loaner device - Customer unit 23 We have found with three clients that the unit somehow creates increased knee varus, and have not used the unit because of this. We have tried to adjust the rotation of the lower limb with one client, but realized by the third person that the increased rotation and adduction was present at all times. I found two of three clients on video, hopefully you can see what I'm referring to. | Clinical team contacted the customer to discuss the clinical issues present and how to resolve the issues. | Patient which was using the device during this issue did not meet the inclusion criteria. Due to this, issues became noticeable which would suggest that there is something wrong with the device. It was concluded that there was nothing wrong with the device but the patient not meeting the inclusion criteria was causing the issue. | (see root cause analysis)
3602 | unit #144 - BGU T√ºbingen - left leg swinging into right leg | closed | kmettler | Actuators | fixed | 2017-04-06 17:05 | 2017-04-06 | 492835 | 144 | BGU Tubingen - Germany | Annette Werz | awerz@bgu-tuebingen.de | +49 7071 6063237 | Germany | bruised ankle | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | No investigation needed - see Justification | Design | customer reported that while walking the left leg startet to swing heavily into the right leg, patient even bruised his ankle lightly. Patient is used to walking with the ekso and does not have spasticity and there was not change in the abduction settings. | Replaced hip housings per SWO-2619 | Found a failure within the hip actuator housing. Known design flaw. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 832 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/21/2017)
3595 | UNIT #177 - Technomex/Kamien Pomorski- Right toe guard broken | closed | duanee | Foot Assy | fixed | 2017-04-05 09:44 | 2017-04-04 |  | 177 | Technomex - Poland | Lech Garasimowicz | lgarasimowicz@technomex.pl | +48 694 483 126 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Design | The distributor found during service review that right toe guard is broken | distributor replaced toe guards, swo 2593 | Toe guards failed during normal use due to stress concentration due to a design related issue. Current revision of the toe guards have been installed which remove this failure mode. See CO-11239. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. Toe guards are installed to lessen wear on foot sensors and periodically replaced at the request of the customer. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/10/2017)
3594 | Error 1180 | closed | kmettler | UI - Controller | fixed | 2017-04-04 22:56 | 2017-04-04 | 995745 | 154 | Confluence Health | Jeremy Coble | Coble, Jeremy <Jeremy.Coble@confluencehealth.org> |  | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Materials | Issue: With their most recent patient Confluence was receiving multiple 1180 error codes. This was happening continually whether or not they have a patient in the device. | Sent parts via SWO-2591 | UI cables failure cause error codes. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. Error code presented to user to reset device. If device reset does not work, replacement of wear item resolves errors. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/5/2017)
3591 | UNIT #178 - Technomex/Kamien Pomorski - Left ankle joint damaged, screw broken | closed | duanee | Ankle | fixed | 2017-04-04 16:27 | 2017-04-04 | 113674 | 178 | Technomex - Poland | Lech Garasimowicz | lgarasimowicz@technomex.pl | +48 694 483 126 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Materials | Left ankle joint is damaged, screw is broken. | Mailed replacement ankle assemblies to distributor per SWO-2592. | The shoulder screw lacked a fillet which decreases total stress during use. Screws currently in use have this fillet. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 785 and # 786 for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 5/24/2017)
3590 | UNIT #109 - EU Loaner at EMAC/Ospedale Sacro Cuore - LEFT KNEE STUCK | closed | kmettler | Actuators | fixed | 2017-04-04 10:00 | 2017-03-21 | 949157 | 109 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39.347.54.34.417 | ITALY | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | No investigation needed - see Justification | Materials | This loaner in Negrar has some problems with left knee: it stuck during the movement, in free legs, in walking or pushing the brake release button is the same. The device cannot be used anymore. Problem with knee brake? | replaced parts, swo 2603 | faulty cca 102286 | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/21/2017)
3586 | Issues with stand to sit | closed | duanee | System | fixed | 2017-04-03 18:44 | 2017-03-28 | 18957 | 290 | University of Brunswick - Canada | Andrew Saxton | asexton@unb.ca | +1 (506) 453-4966 | Canada | No injury | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Wear | Issue: Shantel is doing SA training and while sitting a PT after a walk the Ekso finished the sit 15-20 degree‚Äôs before the person was seated, this left them in an awkward position with the knee‚Äôs and hips locked. Can we look at the data and determine what may be causing this? Please advise. | Customer was walked through how to calibrate the foot sensors. Issue was no longer present after recalibration. | Occurs when there is bad readings from the foot sensors. Requires re-calibration or replacement if foot sensors are damaged. | INV-0006
3584 | unit #278 - Beka/Moscow - problem with right knee | closed | duanee | Actuators | fixed | 2017-04-03 16:25 | 2017-03-31 | 26824 | 278 | Beka RUS - Russia | Sergey Zimin | zimin@beka.ru | +7 495 666-3323, ext. 231 | Russian Federation | No injury | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor | Investigation Needed | Materials | The device has problems with the right knee. There is a noise in the right knee. It runs hot during the workout. Sometimes smells like a burned wire. update from Chad/Michael: after checking the transferred data and videos, it seems to be a problem with the circuit board. Michael will send one to the distributor to fix the issue update from Michael / Dennis Veit: No sound while unit is switched on only. Unit is in legs free: Left knee: Brake makes sound like knee brake batter is empty. Right knee: Knee makes no sound but can be bent hardly only. Brake seems to be free. The customer also reported smoke came from the right knee without any error message. Update from Michael - field service visit on May 23: I replaced both knees but this didn't solve the issue. It is even worse. Now it makes the same sound (like known from the know board) on both sides. The sound occurs after the legs are set to be free. This is from the walk menu and also in the leg test. There is no error shown. | replaced AMP board, swo 2741 | This issue is caused by a failure on the amplifier board (102915). See verification. | The root cause of customer issue has been narrowed to a failure on the device amplifier board. The amplifier board was replaced, and the unit no longer exhibits the reported behavior.
3582 | hip abduction changing during walk | closed | duanee | Adjustments - Hip | fixed | 2017-04-03 14:39 | 2017-04-01 | 45466 | 259 | Zentralklinik Bad Berka - Germany | Fr√ºhauf, Ina | ina.fruehauf@zentralklinik.de | +4936458 52086 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | customer reported hip abduction changing during walk. they carry the tool all the time to set the abduction when walking a patient. | Lead screw threading into control bar was reset per SWO-02447. | Thread pitch on the lead screw assembly is known to allow the hip abduction to change during walk. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3581 | broken right DDL | closed | duanee | Don-Doff Link | fixed | 2017-04-03 10:37 | 2017-04-03 | 234465 | 193 | Technomex - Poland | Jakub Maliszewski | Jakub Maliszewski PHU Technomex Sp. z o.o. <jmaliszewski@technomex.pl> | +48 32 40 10 350 w. 38 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Design | customer reported a broken right DDL | replaced ddl's and hips, swo 2744 | Design requirements did not anticipate use in terms of aggressive split stances during turning (see #2489). | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 832 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/3/2017)
3579 | Issues w/ R step | closed | kmettler | Foot Sensor | fixed | 2017-03-29 00:24 | 2017-03-28 | 76741 | 279 | Sunnyview Rehab | Bridget Rourke | Bridget.Rourke@sphp.com | 518-461-0367 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | Issue: PT called from Sunnyview while a patient was in the device. They were in ProStep + and while walking the device would not initiate the R step. They said this happened in Pro Step and FirstStep as well. We walked them through a foot calibration and it took a few steps and then would no longer initiate any steps except in stand still. | Additional sensors sent via SWO-2599 | Known wear item | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/29/2017)
3578 | unit #177 - Technomex/KAmien Pomorski - replace foot soles | closed | kmettler | Foot Sole | fixed | 2017-03-27 16:03 | 2017-03-27 |  | 177 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 40 10 350 w. 38 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | customer reported that their ekso has a broken foot sole. | send parts to Technomex/Poland - SWO-02571 Parts to be installed by distributor. | Known wear items | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3577 | unit #178 - Technomex/KAmien Pomorski - replace foot soles | closed | kmettler | Foot Sole | fixed | 2017-03-27 15:55 | 2017-03-15 |  | 178 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 40 10 350 w. 38 | POLAND | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | customer reported both soles needs to be replaced. One is torn and other is losing black rubber | send parts to Technomex/Poland - SWO-02571 Parts to be installed by distributor | Known wear items | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3576 | Adjustment Issue | closed | duanee | Adjustments - Leg | fixed | 2017-03-24 23:53 | 2017-03-24 | 217344 | 222 | Kennedy Krieger Institute | Shantel Firpi | SFirpi@eksobionics.com | (510) 984-1761 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Materials | Issue: Very Hard adjustment for left lower leg. There are small white line or scratches all over the inner tube | Collar replaced per SWO-02578. | Compatibility issue (poor fit) between the collar and the upper shank. | This issue does not present a risk to patient or operator safety (see Risk Level Explanation). While inconvenient in this case, the frequency of this issue has not been significant. No further action is necessary. (D. Edmonds ‚Äì 14 Sep 2017)
3575 | unit #141 - EMAC/Villa Adria - errors #121 RK and #133 | closed | kmettler | Motor | fixed | 2017-03-24 15:00 | 2017-03-24 | 823183 | 141 | Distributor - Italy | Matteo Carelli | matteo.carelli@mcmedtech.it | 390103290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | This Ekso yesterday noon started to has problems on the knee that stuck during walking, showing error #133 and 121 RK. This morning they tried again but like the attached video the problem persist. With the device off, pushing the release brake button the knees move smooth without problems or noise. This happen when it move from the sitting position to the standing position, during walking and when it seats down. When it is seat in free legs, if you move the left knee slowly not happen, if you move it quickly it stuck a bit. If you repeat this operation with the Ekso off using the release brake button, it move smooth. | replaced cca, swo 2572 | cca (102286)in left knee was defect, in combination with heavy patient the errors 121, 133 were shown. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for similar investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3573 | unit #48 - Schon Klinik Bad Aibling - foot sensor issue | closed | kmettler | Foot Sensor | fixed | 2017-03-23 16:37 | 2017-03-23 |  | 48 | Bad Aibling - Germany | Denny Muller | DMueller@Schoen-Kliniken.de | +49 8061 903 1679 | Germany | No injury | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor | No investigation needed - see Justification | Wear | customer reported problem with foot sensors, did manual foot test and always received a failed error for the left heel | Customer replaced per SWO-2594 | Known wear component | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See D100117 - Risk Management Spreadsheet (FMEA Tab) # 269 for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3570 | Pulse connectivity issue | closed | kmettler | Pulse Data | fixed | 2017-03-21 21:02 | 2017-03-21 | 48342 | 289 | Dominican Hospital | Andrea Martin | Martin, Andrea - DSC <Andrea.Martin@DignityHealth.org> | (831) 462-7385 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Issue: I have created my log in and done a search for data. It appears there is no data as I can get nothing to come up. Last data 12/12??? | ebd2 software upgraded to 1.1.20 per SWO-02610. | Phone had defective ebd2 software (version 1.1.19). | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 4/19/2017)
3568 | units #233 - To Walk Again/Belgium - missing data in Pulse | closed | kmettler | Pulse Data | fixed | 2017-03-17 10:01 | 2017-03-16 |  | 233 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 479 6299 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Design | By analysing the data on Ekso Pulse we can see that not all our therapy sessions are collected. For example patient ID number 26 we can see that after date 16th of August 2016 (installation SA) we can‚Äôt find 7 sessions in Ekso Pulse. Before this date the error is even bigger : August 2015 : missing sessions May-June 2016 : a lot of missing data‚Ä¶ I just mentionned one patient ID but it seems that this problem occurs with all our patients. We have never connected a session to a serial number of the Ekso used. On the date form there‚Äôs no indication for that so we only filled in the data that are asked to fill in. But I think that it is spread through all our devices (so must not be with unit #233, but could be also with #53 or #232) | Issue resolved through over the air updates of the EBD software | For data sessions in August 2015, all sessions were captured however the wrong customer (0000 - indicating unassigned customer number) was programmed into the device. This is why the customer is unable to see this data. For data sessions from May-June 2016 which are missing as described by the customer, all sessions are available. The date created time stamp is incorrect in the data however the RTC on the Ekso was fully functional and shows the correct date. All sessions were labeled as 8/11/2016. This date coincides where the phone's EBD software was updated to fix a handshake issue. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 105807 - Pulse Risk Management Spreadsheet (FMEA Tab) # 5 for original investigation information. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/21/2017)
3567 | Block during swing | closed | kmettler | Foot Sensor | fixed | 2017-03-16 18:33 | 2017-03-16 | 39275 | 227 | Children's Hospital of Los Angeles | Lauren Dunn | Dunn, Lauren <ldunn@chla.usc.edu> | 949-412-0314 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Design | Issue: Issue: PT's from site called and said when walking patients in 2 free they feel like there is a mechanical block and isn't allowing their patients to finish/start the swing. This setup you see in the video has high resistance and low stance support... They noticed their patients (who used to be able to walk in this mode/setting) not being able to walk with 2 free due to this issue. Previously this wasn't an issue and they were able to use with this same setup. The PT's originally thought it was a patient issue, but then they got a PT in the device and they say there is definitely a hard block that is difficult to power through. Darrell, any idea's from a clinical view on what may be happening? Please send through your thoughts. ----- Hey Josh it's Lauren from CHLA. Just wanted to follow up with the videos Kim sent (she's with a patient now). We had it in 2 free and the videos were in high resistance, low stance support. It happened with our patient in neutral for swing earlier this week so that's why I got in it today to test it. I could trigger a knee flexion block in initiation of swing on both legs, was happening on the left leg more with our patient (his stronger leg). I initially was able to reproduce with a very fast knee flexion moment; however it's not a completely fixed block and if I relax I can then flex past it. It kind of feels like where a trajectory stop might be when in prostep/+ | One of our tech's walked them through foot sensor checks, we re-calibrated the sensors and everything passed. | Device functioned as designed. | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364 for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3566 | unit #266 - Orthomedico - broken/damaged pins on battery connector | closed | bhanson | Battery - Connector | fixed | 2017-03-14 17:12 | 2017-03-14 | 329972 | 266 | OrthoMedico NV - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 (0)479 62 99 45 | Belgium | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Customer Use | The pins to connect the battery to the device are broken/damaged. | replaced battery connector, swo 2551 | possibly touched pin while moving the unit | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 95-C for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/23/2017)
3565 | unit #163 - Mark Pollock - different issues with standing,sitting and walking | closed | duanee | Pick an option | fixed | 2017-03-14 17:06 | 2017-03-14 | 320949 | 163 | Mark Pollock - Ireland | Mark Pollock | <mark@markpollock.com> | +35 387 2233372 | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor | Investigation Needed | Customer Training | customer reported : Over the last couple of days we have been having the following problems with my Ekso GT. * During standing the hips are not fully extending and we are getting a 'standing timed out' error message. We sat down and stood up a couple of times to fix it. * During sitting the knees are not bending fully. They are maybe only bending about two thirds of normal bending. * During walking the device sometimes feels strange. Yesterday it was really weird for the whole session with a felling of softness and misfiring. Today it was mainly fine. However, there was about 2 mins which seemed very unusual. | The customer sent videos and after a check from clinical it came out that the patient was leaned to much backwards. This caused the issue they claimed. Since than no further issues occured. | (see Investigation Summary) | An assessment of risk was performed and the Ekso clinical team was consulted and shown the video sent by the customer. Feedback was provided regarding the positioning of the patient in the Ekso for proper operation. At this point, the use of the device in this particular manner (too much rear lean) appears to be isolated and not related to the training provided by Ekso Bionics. No further action is necessary. (Duane Edmonds ‚Äì 14 Sep 2017)
3564 | Battery draining quickly | closed | kmettler | Battery | fixed | 2017-03-10 21:40 | 2017-03-10 | 576220 | 15 | Magee | Elizabeth Watson | Elizabeth Watson <Ewatson@mageerehab.org> | (215) 218-3900 | usa | No injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | Issue: The batteries are running down very quickly with the patient. He usually gets about 24 minutes of up time (22 stepping) and I am below 35% battery life. When I reach 34-32%, I almost always get a 122 RK error. It happens almost every sessions around the same point in battery life. I know that error is not related to the battery, but it doesn‚Äôt seem like it could be related to the ROM because we are so far in the session. I usually sit the patient and then I change the batteries and standing back to continue the session. This usually works, but today we got the same error after another 11 minutes. | Send parts via SWO-2535 | Known wear item | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3562 | Pot error found during service | closed | kmettler | Encoder / Potentiometer | fixed | 2017-03-10 15:53 | 2017-03-09 | 489358 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Materials | distributor reported: after the first right step, during the swing phase of the left leg, the left knee remain flexed and the device showed the following errors #121 ‚Äì RK and then #133. They did the safe sit, switch off, switch on and the problem happen again, with the same modality. Then they switch it off, try to move the left knee with the hand brake button but it doesn‚Äôt move. So they replaced the battery and after the battery replace, during the standing phase the Ekso remain sit with a continuous left knee noise. No issues found with errors #121 ‚Äì RK and then #133 and instead found pot error. | Replaced defective components per SWO-2570 | Found defects in joint board and soft pot | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 614 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/5/2017)
3561 | unit #169 - EMAC/Porto Potenza Picena - battery connector pin issue | closed | kmettler | Battery - Connector | fixed | 2017-03-10 15:34 | 2017-03-10 |  | 169 | Distributor - Italy | Matteo Carelli | Matteo Carelli <m.carelli@emac.it> | +39 347 5434417 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | distributor reported: this Ekso stop working due to a battery interface pin problem. I'm waiting pictures but should be not fixed without replace it. | distributor replaced parts - SWO-02543 | Customer accidentally docked the battery upside down and damaged the battery interface pins. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 526-C for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/29/2017)
3559 | unit #259 - Zentralklinik Bad Berka - device is switching off | closed | kmettler | Torso | fixed | 2017-03-09 15:42 | 2017-03-09 |  | 259 | Zentralklinik Bad Berka - Germany | Lars Kettwig | lars.kettwig@zentralklinik.de | +4936458 52086 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | The customer contacted us, because it happened several times during the last days that the device ‚Äòkind of‚Äô switched off, i.e. sometimes while a patient was walking and sometime when the device was simply switched on empty, suddenly the UI was black and the customer said that the power was off. On the other hand the knees and hips of the device were still soft, so it seems it was not really a complete switch off. | replaced power switch, swo 2552 | Internal spring was damaged due to repetitive use. Internal spring is rated to withstand repetitive use. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 531-C for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3558 | Error 122 | closed | duanee | Actuators | worksforme | 2017-03-09 00:27 | 2017-03-07 | 576220 | 15 | Magee | Elizabeth Watson | Elizabeth Watson <Ewatson@mageerehab.org> | (215) 218-3900 | USA | No injury | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Customer Use | Issue: They were having patient specific issues, error 122 w/ one of their patients. They used the GT device and did not have the same errors. The batteries are running down very quickly with the patient. He usually gets about 24 minutes of up time (22 stepping) and I am below 35% battery life. When I reach 34-32%, I almost always get a 122 RK error. It happens almost every sessions around the same point in battery life. I know that error is not related to the battery, but it doesn‚Äôt seem like it could be related to the ROM because we are so far in the session. I usually sit the patient and then I change the batteries and standing back to continue the session. This usually works, but today we got the same error after another 11 minutes. (since replaced the battery chargers & sending new batteries) This is all happening on the same patient. He does have some tone initially in the session, but we are in adapt mode and he is running in the 70‚Äôs assist level wise, so I am not sure why they are draining so fast. And I am really confused why the error happens around the same time and so far into the session? I have not been using it recently with anyone else, but I was about 2 months ago and I was not having the same issues although I was not walking as far with those patients. We used the GT model at our other facility yesterday. (3/7) We walked about the same distance- 879 steps with 27 minutes of stepping time and did not get the 122 error. What did happen is something that occurred the last time we used the GT model due to a scheduling issue- his adduction was set the same (1), | Followed up with PT to have them reach out should this issue reoccur. The PT's did not follow up. |  | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. Error code presented to user to reset device. If device reset does not work, replacement of wear item resolves errors. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/5/2017)
3555 | Multiple errors #137 | closed | duanee | System | fixed | 2017-03-07 20:41 | 2017-03-07 | 3574 | 286 | Providence | Jenn | jennifer.cowan2@providence.org | 310-832-3311 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Customer Use | Issue: Issue: Mike is doing advanced training and is getting repeated #137 errors with one of their patients. Can we look into the data to see if there is anything we can do/learn from what‚Äôs occurring? Please advise, device is on, Mike is at lunch at the moment. | Data from the unit was pulled on 8/14/2017 for the past months since the incident. There have not been a repeat of the error since the date of the incident. | Issue with customer use. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Brian Hanson on 5/17/2017)
3552 | Foot Binding Peeling | closed | duanee | Foot Assy | fixed | 2017-03-06 21:03 | 2017-03-06 | 3574 | 286 | Providence | Jen Cowan | jennifer.cowan2@providence.org | 614-296-2881 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Determined Unknown | Issue: A couple cosmetic things at Providence San Pedro. The foot binding pad Velcro piece has started to peel away from the foot binding. Any specific glue necessary to gain a better adherence? | Sent parts to fix issue. Parts replaced by Mike Glover in the field. | Will not be getting the part back so unknown if this is a wear or build issue. | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 5/16/2017)
3550 | unit #259 - Zentralklinik Bad Berka - battery is discharging fast | closed | kmettler | Battery | fixed | 2017-03-03 10:59 | 2017-03-02 |  | 259 | Zentralklinik Bad Berka - Germany | Ina Fruhauf | ina.fruehauf@zentralklinik.de | +4936458 52086 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Wear | customer noted that one of the batteries is discharing very fast, i.e. when they remove them after one patient has walked and only re-use them one or two hours later they get an error that the battery loads are not matching anymore. They even re-checked this on the charger where they could see that one battery was been loading fully in a couple of minutes and the other needed half an hour or more. Update: Udo confirmed from service visit that it is not possbible to use the two batteries together because of the discharging issue. | send two new batteries - SWO-02560 | Batteries are known wear component | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 4/6/2017)
3549 | unit #243 - Step by Step/Barcelona - right hip blocked | closed | kmettler | Actuators | fixed | 2017-03-03 10:18 | 2017-03-02 | 718447 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major | No investigation needed - see Justification | Materials | I'm sorry to say that the right leg of our unit is giving problems again. At this time, the right hip is totally blocked in full extension position and it is impossible to move. I tried to do the leg test but it was impossible. | replaced hip assemblies, swo 2518 | broken flexure | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 17 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/23/2017)
3548 | Torn Sensors | closed | kmettler | Foot Sole | fixed | 2017-03-02 23:44 | 2017-03-02 | 76741 | 130 | Sunnyview Rehab | Greg Devault | GDavault@eksobionics.com | 518-382-4528 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | No investigation needed - see Justification | Customer Use | Issue: Foot sensors ripped on loaner device @ Sunnyview as they are walking on carpet. - No picture | New foot sensors installed on Sunnyview device unit 279 per SWO-2516 | Wear item/ walking on carpet was the main cause to the sensors not working | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3544 | unit #268 - EMAC/IRCCS Centro Neurolesi Bonino Pulejo - foot sensors | closed | kmettler | Foot Sensor | fixed | 2017-03-01 15:19 | 2017-03-01 | 112710 | 268 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | N/A - Identified known wear item | Wear | distributor reported that the left knee click when is in charge and fully extended and the right leg is not able to reach the set length. This happen with more than one patients and the patient on the video has not spasticity. Yesterday I follow them on the foot auto calibration and it passed, but today the problem appear again. Update from Michael: I had a unit in my workshop last week which walked exactly like this. I replaced the foot sensors and it solved the issue. | distributor replaced foot sensor - SWO-02515 | Foot sensors are known wear items that periodically require recalibration or replacement. | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3537 | Missing Index Error L Knee | closed | kmettler | Encoder / Potentiometer | fixed | 2017-02-21 16:18 | 2017-02-21 | 718447 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | customer reported an LK-21 sensor error. He will do a leg test and inform us about the result. | replaced knee motor, swo 2506 | Halls issue on PCBA within motor was causing the issue. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/16/2017)
3535 | unit #235 - Rolling Ball/Garry McCabe - replace softgoods | closed | bhanson | Soft Goods - Torso / Sacral | fixed | 2017-02-21 13:22 | 2017-02-21 |  | 235 | Rolling Ball - Ireland | Gerry McCabe | gerrytrans@gmail.com |  | Ireland | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | customer called and needs new torso pads because velcro is not sticking anymore | send parts with SWO-02485 | Known wear item | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Brian Hanson on 3/10/2017)
3530 | Issue with right knee | closed | duanee | Other | worksforme | 2017-02-17 14:00 | 2017-02-16 | 156051 | 139 | Asepeyo - Madrid | Josefina Mart√≠n Osuna | jmartinosuna@asepeyo.es | +34 91 627 62 89 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | Investigation Needed | Determined Unknown | customer reported: On only the left knee works in "free legs" mode and the right sometimes fails. | Michael made full check on device and found no issues |  | A full check of the unit was performed and no issues were found. No further action is necessary. I was unable to find another instance of the reported behavior in the complaint database. No further action is necessary. (Duane Edmonds ‚Äì 29 June 2017)
3529 | unit #45 - Justin Jeffrey/South Africa - worn soles and shark fins | closed | kmettler | Other | fixed | 2017-02-14 16:24 | 2017-02-08 | 736491 | 45 | Justin Smith - So. Africa | Justin Jeffery | Justin Jeffery <justbiokinetics@gmail.com> | 072 952 4824 | South Africa | No injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor | N/A - Identified known wear item | Wear | the customer reported worn soles (black rubber is partly missing) and shark fins (See pics) | replaced parts, swo 2481 | known wear items | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 3/15/2017)
3528 | Broken hip housing | closed | kmettler | Actuators | fixed | 2017-02-14 16:16 | 2017-02-08 | 736491 | 45 | Justin Smith - So. Africa | Justin Jeffery | Justin Jeffery <justbiokinetics@gmail.com> | 072 952 4824 | South Africa | No injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Design | on request from us the customer checked the housing for cracks and found a broken hip housing | repalced hip assy, swo 2481 | Stress concentration at crack origin. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See D100117 - Risk Management Spreadsheet (FMEA Tab) # 175 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3527 | Screw on legs hard to adjust | closed | kmettler | Adjustments - Leg | fixed | 2017-02-13 22:56 | 2017-02-10 | 69232 | 264 | Cottage Rehab | Noah Gaines | Noah Gaines <ngaines@sbch.org> | 805-680-7295 | USA | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Wear | Issue: We have had a continuing problem with the two lower screws on the both legs getting sticky and not coming undone. Chad- Sounds like they have the old version of the collars so please bring a replacement set. | SWO-2466. New shank collars installed on the unit. (NOTE: SWO Shows shipped, however actually serviced in field) | Hardened bolts used to adjust the legs and lock them in place are causing wear on the shank collars. These will occasionally slide over the bolt head due to wear, entrapping the bolt so removal is not possible until the collar is lifted out of the way. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 4/21/2017)
3526 | unit #243 - Step by Step/Barcelona - foot sensors | closed | kmettler | Foot Sensor | fixed | 2017-02-13 16:36 | 2017-02-13 | 718447 | 243 | Step-by-Step - Spain | Joan Romero | Joan Romero <jromero@fundacionstepbystep.org> | +34 93 223 16 63 | Spain | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | the customer reported problems with the foot sensors | customer replaced foot sensor - SWO-02464 | Known wear
3524 | Motor error 133 and 121 | closed | kmettler | Motor | wontfix | 2017-02-10 13:53 | 2017-02-10 | 104009 | 260 | BG Klinikum Hamburg - Germany | Catharina Ohlmeier | c.ohlmeier@bgk-hamburg.de | +49 40 7306 3673 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Design | customer reported that they got twice a motor error #133 and then #121 with a patient walking in 2Free with low and high swing phase and high and medium support in stand phase. After confirming in the menu they could continue to walk without further issues. | problem did not occur again | Known system limitation. The comment from Katie "It could be just due to using it in 2free and backdriving the motors in combination with a pretty full battery. I don‚Äôt think it‚Äôs any actual issue if it doesn‚Äôt recur." was found to be correct. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 4/10/2017)
3522 | L Knee is in extension | closed | lmcneil | Adjustments - Leg | fixed | 2017-02-09 23:59 | 2017-02-09 | 6303 | 300 | RIM | Diane Patzer | Patzer, Diane E <DPatzer@dmc.org> | (313) 966-8309 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Materials | Issue: The L knee is stuck in extension | Joint board replaced per SWO-02455. | The joint board in the knee was not able to release the brake. The board contacts were cleaned and the board was tested during operational qualification squat in unit 111 without replicating the issue. Part was given to engineering for further investigation. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3521 | Issues w/ First Step | closed | duanee | Other | fixed | 2017-02-09 23:49 | 2017-02-09 | 111825 | 107 | ICORD - Canada | Alison Williams | awilliams@icord.org | 604-675-8852 | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Determined Unknown | Hey Lily, Thanks again to you and Chad for your help yesterday! Unfortunately we are still having some issues with that first step. I've attached a couple of videos that might help explain it better. The files were too big to upload in an email so I've copied them into a dropbox link, let me know if you have any issues accessing them: https://www.dropbox.com/sh/boanox4r10cds36/AADNLIZOiRPkvvQxbL5V8xiFa?dl=0 Video 1 - in this one you can see the main issue. When we try and deliver the first step, it's as if the foot doesn't lift off of the ground properly and gets stuck. When this happens, we can't complete the step (the right foot only moves forward about 2-3") so we have to go to standstill. Video 2 - this is what happens if we ask the participant to lean really far to the left so we can try and get clearance on the right foot. In this one, you can see the foot clearing, but the swing seems slow and catches a bit in the middle in comparison to the following right steps. Video 3 - once we get the participant walking, the right leg makes a lot of clicking noises. If you turn up the volume in this video, you can hear all of the clicking. Let us know if you have any thoughts on other things we could try! Thanks, Ali | Replaced components per SWO-02542. | Unable to accurately what caused this issue. The entire signal chain has been removed and replaced from the device. | This is a minor issue with the first step of the device. Otherwise, the unit performs as expected. The first step issue has now been resolved by replacing most of the signal chain and motor assembly of the device (SWO-02542). No further action is necessary. (Duane Edmonds ‚Äì 29 June 2017)
3520 | Error codes 133 and 1057 causing straight leg sit | closed | kmettler | Other | fixed | 2017-02-09 23:19 | 2017-02-09 | 76741 | 279 | Sunnyview Rehab | Alison Tallon | Allison Tallon <Allison.Tallon@sphp.com> | 518-386-3585 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Materials | Issue: I have not gotten an email to try to setup a time to speak with someone from the clinical staff RE: our patient who has gotten a software error code 2 sessions in a row. Please have someone try to get ahold of me ASAP because the patient is returning tomorrow for a session. They have had a 133 error & then a 1057 error in multiple sessions It is happening 45-50 min into the session. Several sessions it happened every 13 minutes. The patient has 12 degrees knee flexion. when the error comes up the device goes into safe mode but it doesn't let them sit in free legs. | The amp board was replaced per SWO-02516. | Eric checked the serial numbers for the ELMO's and found that the issue could have been caused by a bad batch of ELMO's. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3519 | hip abduction changing during walk | closed | duanee | Adjustments - Hip | fixed | 2017-02-09 16:23 | 2017-02-09 | 283265 | 266 | OrthoMedico NV - Belgium | Kris Goos | Kris Goos <kris.goos@towalkagain.be> | +32 (0)479 62 99 45 | Belgium | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Customer reported the hip abduction is changing during walk | disassembled hip abduction assy, cleaned threads on screw control bar, reassembled the hip abduction assy no parts needed, no swo | Design issue, as reported by engineering. More pronounced in some units, no real screening other than standard test method. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3518 | unit #179 - CareTechnique - empty knee batteries | closed | duanee | Brake / Release / 9V Battery | fixed | 2017-02-09 15:21 | 2017-02-07 | 20681 | 179 | Care Techniques - Saudi Arabia | Moustafa Kamal | Moustafa M.Kamal <eng.moustafa.kamal@hotmail.com> | +966 53 6003071 | Saudi Arabia | No injury |  | ekso 1.2 | FALSE | 3 - Next Service | Minor | N/A - Identified known wear item | Wear | our PT reported from training that the knee batteries are empty | replaced knee batteries, swo 2740 | Known wear item.
3516 | unit #278 - Beka/NNIITO-Nizhny Novgorod - loose foot sensor | closed | kmettler | Foot Sensor | fixed | 2017-02-09 12:05 | 2017-02-09 |  | 278 | Beka RUS - Russia | Sergey Zimin | zimin@beka.ru | +7 495 666-3323, ext. 231 | Russia | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | distributor reported that their customer's device lost heel part of the foot sensor - see pic attached | distributor replaced foot sensors - SWO-02457 | known wear component | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/13/2017)
3515 | unit #248 - Beka/Moscow - strange noise | closed | bhanson | Adjustments - Leg | fixed | 2017-02-08 15:56 | 2017-02-08 |  | 248 | Beka RUS - Russia | Sergey Zimin | zimin@beka.ru | 007 495 666-3323, ext. 231 | Russia | no injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Wear | Distributor reported that the device makes a strange noise. See attached video. They think in might be coming from the left knee motor. | distributor made leg test successfully | Over time the joints of the Ekso will require re-calibration. If the re-calibration is not done at a regular interval it is possible for the brakes to engage earlier than expected. When this occurs there is no risk to the patient however the device will make the sound as heard in the video. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/15/2017)
3514 | Pulse display is off | closed | kmettler | Software | fixed | 2017-02-07 23:40 | 2017-01-30 | 281493 | 164 | VA Minneapolis | Elizabeth Koch | elizabeth.koch3@va.gov | (612) 467-4845 |  | no injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Design | Issue: Minn VA Issue: signed up for pulse but the display appears to be off. Attached is a pic, it's not great but you can see where it actually says "drop down" instead of the pull down menu. Please let me know if a better pic is needed. PT did say they are running off old comps. | PT did say they are running off old comps, once they used a different browser this issue did not reoccur. | Pulse can't run off old browsers. It is only supported off of: We support firefox, safari, chrome, IE | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 4/19/2017)
3513 | Pulse Errors | closed | kmettler | Software | fixed | 2017-02-07 20:53 | 2017-01-31 | 75206 | 49 | Able Bionics - Canada | Mitch Brogan | ablebionics@gmail.com |  | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Design | Customer: Able Bionics/ Mitch Brogan Unit #: 49 Issue: When trying to update the password a box comes up saying "invitation_token:is invalid" and we are unable to update the password. Rochelle, can you please advise? We‚Äôd like to get the customer up & running as soon as we can. Thanks! | Customer was able to login when using a different browser | Customer was using outdated software on their browser which can interfere w/ the Pulse login | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/9/2017)
3512 | Problem with right hip | closed | kmettler | Fasteners | fixed | 2017-02-07 17:39 | 2017-02-07 | 736491 | 45 | Justin Smith - So. Africa | Justin Jeffery | Justin Jeffery <justbiokinetics@gmail.com> | 072 952 4824 | South Africa | no injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Design | customer contacted us because right hip is blocked, screw is missing (see pic attached) | replaced hip assy, swo 2481 | loose and broken screws on flexure. Screw length was not sufficient for design and was corrected. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See D100117 - Risk Management Spreadsheet (FMEA Tab) # 162 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3511 | unit #179 - CareTechnique - error #25-LK | closed | duanee | Pick an option | fixed | 2017-02-07 17:26 | 2017-02-07 | 20681 | 179 | Care Techniques - Saudi Arabia | Moustafa Kamal | Moustafa M.Kamal <eng.moustafa.kamal@hotmail.com> | +966 53 6003071 | Saudi Arabia | No injury |  | ekso 1.2 | FALSE | 9 - Loaner | Minor | Investigation Needed | Materials | Our physiotherapist reported from training that device gave error # 25-LK. Distributor confirmed that this problem happend several time with and without patients. Our physiotherapist tested again today to make sure, but still the same result: ---- Please find below our procedure/history of our retest with the device today: - We switched the device on (without patient) ‚Ä¶‚Ä¶..device required leg test immediately - We went through the leg tests and all tests passed. - Immediately after the leg test the error came up again when we moved the legs to sit the Ekso back on the chair - We switched the device off ‚Ä¶put the PT inside‚Ä¶.switched it on ‚Ä¶again it required a leg test. - Additionally the PTs said that error 25# LK also occurred, when they were walking a patient So as a result ‚Äì there seems to be a problem regarding the left knee sensor - no matter if the device is used with or without a patients. | replaced soft pot, swo2740 | There is a location on the soft pot that is not reading correctly when the plunger passes over it. | Upon investigation by the Field Service team, it was discovered that the potentiometer was not operating correctly in a specific location along the travel of the wiper. This is caused the BIT routine to generate the warning by triggering an error. The BIT safety checks performed as designed. The faulty potentiometer was replaced. No further action is necessary. (D. Edmonds ‚Äì 16 Oct 2017)
3510 | Controller Malfunctioned | closed | bhanson | UI - Controller | fixed | 2017-02-06 23:59 | 2017-02-06 | 628966 | 136 | Kessler Study | Erica Garbarini | egarbarini@kesslerfoundation.org | (973) 731-3600 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Determined Unknown | ISSUE: Hi, Josh, I hope all is well with you! FYI, the controller on our original device (the Investigational unit) malfunctioned upon standing with a participant this afternoon. I powered the device down and lowered the participant back onto the bench. When we turned the device back on, the controller worked properly and training proceeded without incident. I have attached a pic/flick of the controller screen. The User ID was 0031; the time was approx. 1:15 ET. This problem had also occurred back in March of 2015 (we sent you a pic/flick at that time, as well). Thanks, ~Erica | Sent new parts via SWO-2444 | Did not receive malfunctioning component | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 110 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3509 | unit #127 - EMAC/Padre Pio - lost toe guard | closed | jhanff | Foot Sole | fixed | 2017-02-06 16:58 | 2017-02-06 | 520688 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor | No investigation needed - see Justification | Wear | distributor reported a lost toe guard | Work completed per SWO-2441. | known wear component | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. Toe guards are installed to lessen wear on foot sensors and periodically replaced at the request of the customer. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/9/2017)
3508 | Error 51 right hip | closed | bhanson | Electrical Interconnect | fixed | 2017-02-06 14:33 | 2017-02-06 | 535880 | 152 | Duisburg - Germany | Ralf Lanzmich | Ralf.Lanzmich@bg-klinikum-duisburg.de | +49 203 76883613 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | No investigation needed - see Justification | Wear | customer contacted us because his device gave an error # 51 for the right hip | Replaced hip cable, swo 2417 | External cables were bent during normal use and over time a conductor failed within the cable causing the customer facing issue. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/14/2017)
3507 | Lower Legs rotating | closed | duanee | Adjustments - Leg | fixed | 2017-02-03 23:38 | 2017-02-03 | 57018 | 231 | University of Calgary - Canada | Kyle McIntosh | Kyle.McIntosh@albertahealthservices.ca |  | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Wear | Issue: We have also had some issues with the lower leg adjustment clamps not holding firmly. As a result, the lower legs rotate at times while walking and today the left lower leg shortened about a quarter inch mid session. We would love to have that issue looked at as well. We are torqueing each screw head twice as prescribed and our T-handle torque wrenches have been calibrated. | Ordered replacement T-handles to be sent to customer. Work completed per SWO-2689. | Tested (using field service tools) and determined both of customer‚Äôs calibrated T-handles were under-torquing. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 28 for similar investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3505 | Leg Rotating | closed | kmettler | Adjustments - Leg | fixed | 2017-02-02 22:54 | 2017-01-30 |  | 247 | RIM | Diane Patzer | Patzer, Diane E <DPatzer@dmc.org> | 313-682-6249 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Materials | Issue: Thanks so much for sending out the additional screws and washers! We have received them. We replaced the two on the lower leg and still had some rotation with one of our clients. We will continue to monitor and I'll let you know how it goes :o) | Collar replaced with most current rev per SWO-02436. | Adjustment bolts bottoming out on the collar on the upper shank. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 41 & 42 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3504 | Problem with hip width/abduction | closed | kmettler | Adjustments - Hip | fixed | 2017-01-31 14:03 | 2017-01-31 | 1117207 | 233 | To Walk Again - Belgium | Kris Goos | kris.goos@towalkagain.be | +32 (0)479 62 99 45 | Belgium | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Wear | customer reported that there device has problems with opening hip width right side and hip abduction right side impossible to move | disassembled hip abduction assy, cleaned tubes and lock, reassembled the hip abduction assy. Work completed per SWO-2541. | Known wear related issue | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3503 | unit #131 - RKU Ulm - error 43 | closed | kmettler | Foot Sensor | fixed | 2017-01-31 13:57 | 2017-01-31 | 105388 | 131 | ULM - Germany | Lux, Sebastian | sebastian.lux@rku.de | +49 731 1775530 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | N/A - Identified known wear item | Wear | customer reported that the device shows error 43. they made a foot test, which was successful, but when they walked with a patient again afterwards the same error showed up again. | Mailed foot sensors per SWO-2577. | Failiure 43 during use indicates foot sensor recalibration needed. | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. Error codes 40 - 43 indicate foot sensors out of range. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/7/2017)
3502 | Issue with left leg/hip | closed | kmettler | Electrical Interconnect | fixed | 2017-01-27 15:05 | 2017-01-23 | 249721 | 26 | Guttmann - Spain | Silvana Andrade | electromedicina@guttmann.com | +34 93.497.77.00 ext. 3199 | Spain | no injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Workmanship | customer reported a problem with the left hip/leg of their loaner. | Reinstalled cable clamp per SWO-2418 | Found that the cable clamp which secures the knee external cable was not properly installed and caused the cables to disconnect over time. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/6/2017)
3501 | unit #152 - BGU Duisburg - error 121 RK | closed | kmettler | Motor | fixed | 2017-01-25 16:14 | 2017-01-25 | 594615 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | 4920376883613 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Wear | customer reported a humming sound when standing up a patient and then direcly afterwards they got error 121 RK Error Code 4546 | replaced parts, swo 2417 | External cables failed over time due to wear | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/14/2017)
3499 | L Leg won't step | closed | kmettler | Foot Sensor | fixed | 2017-01-23 22:09 | 2017-01-23 | 39275 | 227 | Children's Hospital of Los Angeles | Lauren Dunn | Dunn, Lauren <ldunn@chla.usc.edu> | 949-412-0314 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: CHLA reported they are having issues w/ their foot sensors They aren‚Äôt able to step w/ the left leg, even with the targets being made on the R. It was happening 3-4 times in 20 steps and now it is happening around every step. They are in Prostep Max | Sent foot sensors via SWO-2404 | Known wear item
3498 | Duplicate:Step Count reset | closed | kmettler | Software | duplicate | 2017-01-23 19:14 | 2017-01-20 | 23556 | 231 | University of Calgary - Canada | Kyle McIntosh | Kyle.McIntosh@albertahealthservices.ca | 403-944-5358 | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor |  | Undetermined | Issue: The device today has a step total of 33,298 while we have over 100,000 documented steps from our study participants alone. I haven‚Äôt calculated the ‚Äúup‚Äù and ‚Äúwalk‚Äù times but they also appear proportionally out of whack. I can understand that the pulse data might be subject to a lag but shouldn‚Äôt the device itself be accurate? Thanks Lily ‚Äì and Happy New Year to you! | (Explain actions taken to resolve this issue) | (Required for complaints)
3497 | Customer not receiving data | closed | kmettler | Software | fixed | 2017-01-23 19:12 | 2017-01-22 | 57018 | 231 | University of Calgary - Canada | Kyle McIntosh | Kyle.McIntosh@albertahealthservices.ca | 403-944-5358 | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Issue: Initial complaint from customer - When looking into Pulse, they aren't able to access any of their recent data (only from March of this past year) Josh - I believe we found a bug within Pulse. I've been trying to connect Calgary to Pulse for months but it wasn't showing any of their current data ( they have the Vodafone SIM). Taking a deeper dive, we are receiving their data through our data dump system.... Back in March they had a loaner and for some reason this is the only data they can see, this loaner no longer has their customer #, nor has it for almost a year, but continues to be associated with their device. I can show you exactly what's occurring tomorrow. Update 1/24 - RRea - Customer data is not getting to the cloud. The data you see on the cloud is from an internal device whose amp board was scrapped by mfg into engineering and the S/N and customer number were not wiped. This is the data you see on AWS but not on pulse. You do not see this data on pulse because there are no steps taken which is a requirement of displaying data in Pulse. This issue is capture in TRAC #3500 | Replaced phone according to currently accepted best practices. Work completed per SWO-2621. | Known phone issue for Canadian customers. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 5/19/2017)
3496 | right knee motor stiff | closed | kmettler | Actuators | fixed | 2017-01-20 14:24 | 2017-01-19 | 594615 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | customer reported issues with the unit, different patients, issues occure sometimes only. Issues: noise (unlocated), unit "stiff", error messages (cannot remember the numbers) | replaced right knee motor, elmo calibration, swo 2399 | Error 121 was present. During investigation it was found that the 121 error referenced an internal short within the motor. Unable to accurately determine where the short is within the motor.
3495 | Replacement for don doff skirts | closed | kmettler | Other | fixed | 2017-01-20 14:17 | 2017-01-16 |  | 120 | Spandau - Germany | Bettina Quentin | Bettina.Quentin@vivantes.de |  | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | customer needs replacement for don doff skirts | Work completed per SWO-2400 | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 3/13/2017)
3494 | unit #113 - Vivantes Spandau - replacement for don doff skirts | closed | kmettler | Other | fixed | 2017-01-20 14:14 | 2017-01-16 | 11204 | 113 | Spandau - Germany | Bettina Quentin | Bettina.Quentin@vivantes.de |  | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | cutomer needs replacement for don doff skirts | Work completed per SWO-2400 | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 3/13/2017)
3493 | unit #113 - Vivantes Spandau - leg test not successful | closed | kmettler | Other | fixed | 2017-01-20 13:53 | 2017-01-16 | 11204 | 113 | Spandau - Germany | Bettina Quentin | Bettina.Quentin@vivantes.de |  | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | No investigation needed - see Justification | Customer Use | device asked for leg test which could not be finished successful | Completed leg calibration per SWO-2400 | Customer did not correctly perform leg calibration. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/6/2017)
3492 | Hip is blocked again | closed | kmettler | Ball Screw | fixed | 2017-01-20 13:35 | 2017-01-12 | 718447 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | No investigation needed - see Justification | Materials | customer reported that hip is blocked again | Work completed per SWO-2345 and verified as complete per SWO-2518 | Broken flexure | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 17 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3491 | unit #140 - EMAC/Villa Beretta - Adjustment tool damaged | closed | bhanson | Adjustments - Hip | fixed | 2017-01-20 12:51 | 2017-01-16 | 372344 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 347 5434417 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | N/A - Identified known wear item | Wear | distributor reported that it was hard to set the hip width. Found the adjustment tool to be worn causing the customer facing issue. | Distributor replaced bit on tool from Distributor stock. | Head of screw showed excessive wear due to repetitive use. It was the tool that was worn out and not hip abduction adjustments. | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler on 3/14/2017)
3490 | DUPLICATE - Hip width adjustment right is stuck | closed | kmettler | Adjustments - Hip | duplicate | 2017-01-20 12:44 | 2017-01-20 | 302874 | 131 | ULM - Germany | Dominik Mayer | dominik.mayer@rku.de | +49 731 1775530 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Undetermined | hip width adjustment right is stuck | Duplicate of #3463 | (Required for complaints)
3489 | DUPLICATE - R swing trajectory | closed | kmettler | Adjustments - Hip | duplicate | 2017-01-19 18:01 | 2017-01-19 | 540133 | 106 | RIC | Kristen Hohl | khohl@ricres.org | (312) 238-1000 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Unassigned |  | Undetermined | Issue: Just wanted to touch base with you after this morning's session with one of the two subjects that we were having trouble with last Friday. The first 10 minutes of walking looked normal (smooth swing phase and normal placement on the right). Then I started to noticed the same thing that I was noticing before with the non smooth trajectory and adduction whip on the right. I've included a video. -I set the hip abduction value at the # he was had previously been using. -Upon noticing this, I checked the position of the hip abduction values - and there was no appreciable change. Any further thoughts? Can someone investigate our toe sensors again to see if it started to misread again? | Duplicate to TRAC #3438 and TRAC #3447 | (Required for complaints)
3488 | Uneven steps | closed | kmettler | System | fixed | 2017-01-19 17:23 | 2017-01-19 | 70078 | 254 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | Issue: Yesterday our patient at 4:00 pm ran into some issues regarding the loaner EKSO once again. I was not present but the employees described it as walking uneven and in a ‚Äúclunky‚Äù fashion. Today we had a demonstration using the loaner EKSO and it decided to act up on us again. I was not there to witness the demonstration but the employees said the EKSO was acting up in the same way it was yesterday. The steps were not smooth and uneven, and the aide said the patient actually was off balance at times. This is pretty rare for this patient as he‚Äôs had plenty of experience and needs very little assistance when using the robot. I checked the foot calibration again and everything seems to be working just fine in that department. Yesterday we talked about moving up to the leg calibration and I was wondering how to do that. I‚Äôll be available at any time and if you prefer to talk over the phone I‚Äôll be more than happy to. | Foot sensor calibration passed during SWO-02427. | Possible foot sensor calibration issue or customer use issue. Manual foot sensor cal seemed to reduce the noise in the sensor (under thresholds in tech menu). Both sensors appear to be reading correctly.
3487 | Pant tears from the heel cup | closed | duanee | Ankle | fixed | 2017-01-17 19:47 | 2017-01-17 | 22552 | 229 | TIRR | Marcie Kern | Marcie.Kern@memorialhermann.org | (713) 799-5051 | United States | No injury | (same as contact listed) | ekso 1.2 | FALSE | 3 - Next Service | Minor | Investigation Needed | Design | Issue: TIRR‚Äôs advanced training and the team reports multiple scrub pants tears from the heel support when walking with patients. | Heel cups updated to most recent version per SWO-02438. | Known issue with previous revision of heel cup. | The heel cup used initially on this device (EKSO-0229-021507006) had the unintended consequence of causing slight tears, etc., in the pant cuff of some users in the field when caught. There is no safety hazard caused by this, although it is probably annoying to most who experience it. The heel cup (103835) on the customer unit was updated to the most recent revision. No further action is necessary. (D. Edmonds ‚Äì 16 Oct 2017)
3486 | DDL lever difficult to release | closed | duanee | Don-Doff Link | fixed | 2017-01-17 19:44 | 2017-01-17 | 22552 | 229 | TIRR | Marcie Kern | Marcie.Kern@memorialhermann.org | (713) 799-5051 | USA |  |  | ekso 1.2 | FALSE | 3 - Next Service | Minor | Investigation Needed | Materials | Issue: Also, the team reports increased challenge with silver lever release on R Hip DDL. I referenced the maintenance log of weekly moving the DDL Lever. The team reports weekly use and this continues to be an issue. One of the PTs states that improved release of DDL when bilateral ADDuction pressure is applied through legs, as typically taught, but one of the PTs is too petite to perform this way. I also suggested to apply greater force through the R Hip adduction and release the R hip first, then move to the L hip. They report attempting this way, but no significant improvement. | Updated the following service, via SWO-2438. | Known alignment issue between the bullet and the t-block in the DDL. | This issue with the DDL lever being more difficult than normal to release represents an inconvenience to the customer but does not present a safety hazard. The DDL components on the customer unit were replaced with updated components (updated in CO-11684). No further action is necessary. (D. Edmonds ‚Äì 16 Oct 2017)
3485 | Short R Step | closed | kmettler | System | fixed | 2017-01-16 20:40 | 2017-01-11 | 70078 | 254 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | N/A - Identified known wear item | Determined Unknown | Issue: Customer: Confluence Unit #254 (Loaner) Issue: PT called the hotline and said with their most recent session they were consistently getting a much shorter R step then Left. Are we able to look into the data to see if they are having foot sensor issues, I can walk them through a foot cal if this is the problem. Thanks! | Foot sensor calibration passed during SWO-02427. | Possible foot sensor calibration issue or customer use issue. Manual foot sensor cal seemed to reduce the noise in the sensor (under thresholds in tech menu). Both sensors appear to be reading correctly. | No investigation needed as identified foot sensors as known wear items that are recalibrate or replaced periodically at the request of the customer. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 364-C for similar risk analysis information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/9/2017)
3484 | cracked left hip housing | closed | kmettler | Actuators | fixed | 2017-01-13 09:21 | 2016-01-12 | 766332 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 347 5434417 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | distributor repoerted a crack in the left hip housing | Exchanged broken left hip housing for a new one. swo 2435 | Known design flaw | No new investigation needed as previously investigated and deemed not a safety rick or a reportable event. See FMEA [1.2 FMEA-390-C] for original investigation information and mitigations put in place. (Evaluation done by Brian Hanson on 3/6/2017)
3473 | Step count reset | closed | ckealoha | Software | wontfix | 2017-01-11 22:17 | 2017-01-11 | 18397 | 231 | University of Calgary - Canada | Kyle McIntosh | Kyle.McIntosh@albertahealthservices.ca | 403-220-7240 | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | No investigation needed - see Justification | Design | Issue: Just wanted to follow up with my question about the step count, walk time and up time in the ‚Äúabout‚Äù section on our device remote. The device today has a step total of 33,298 while we have over 100,000 documented steps from our study participants alone. I haven‚Äôt calculated the ‚Äúup‚Äù and ‚Äúwalk‚Äù times but they also appear proportionally out of whack. I can understand that the pulse data might be subject to a lag but shouldn‚Äôt the device itself be accurate? Thanks Lily ‚Äì and Happy New Year to you! | No further action is required. There is no safety risk associated with this behavior. | The core mechanism for this behavior has not been determined. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Brian Hanson on 5/23/2017)
3472 | Short R Step | closed | kmettler | Foot Sensor | duplicate | 2017-01-11 22:10 | 2017-01-11 | 33024 | 254 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Issue: PT called the hotline and said with their most recent session they were consistently getting a much shorter R step then Left. Are we able to look into the data to see if they are having foot sensor issues, I can walk them through a foot cal if this is the problem. Thanks! | Foot sensor calibration passed during SWO-02427. | Possible foot sensor calibration issue or customer use issue. Manual foot sensor cal seemed to reduce the noise in the sensor (under thresholds in tech menu). Both sensors appear to be reading correctly. | DUPLICATE
3470 | unit #258 - Pasori/France - shank is rotating | closed | kmettler | Lower Leg | fixed | 2017-01-10 15:31 | 2017-01-10 |  | 258 | CRF Pasori - France | Frederique Bordet | f.bordet@pasori.fr | +33 3.86.28.43.89 | France | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Wear | customer reported that a leg is twisting. they had to stop mid exercice to untwist the leg and then go again. Additionally she wrote that 'they changed the rings'. | Replaced adjustments screws and cleaned assembly. Work completed per SWO-2459. | Found that over time the adjustments screws stretched and caused a reduction in clamping force. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 41 & 42 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3469 | unit #141 - EMAC/Villa Adria - left knee cover broken | closed | kmettler | Other | fixed | 2017-01-09 12:34 | 2017-01-09 | 766332 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | left knee cover broken | Distributor replaced component per SWO-2378 | Cover was damaged during normal use
3468 | Joint Cap Failure | closed | kmettler | Actuators | fixed | 2017-01-09 12:27 | 2017-01-09 | 766332 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | No investigation needed - see Justification | Design | washer in hip broken | Distributor installed new parts per SWO-2378 | Joint cap failure occured but went unnoticed. The washer was then free to move between the hip housing and the actuator link causing the damage as found in the attached pictures. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 768 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/2/2017).
3467 | CI not engaging | closed | duanee | UI - Crutch Interface | fixed | 2017-01-06 18:54 | 2017-01-06 | 6332 | 296 | VA Long Beach | Lisa Hughes | LHughes@eksobionics.com | 484-753-5650 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | Investigation Needed | Materials | Issue: Hi team - When we went to plug in and try the crutch interface for the first time today it is not working. No green light and not engaging. Please advise. Thanks, Lisa SN - 022608001 PowON- 1603 Stand - 789 Walk - 216 Steps - 6332 | CI mailed to customer per SWO-02372. | Material Failure. | The component in question, the HMI crutch interface, failed in such a way that it cannot be used with the device. As a result, the failed component does not present a hazard to the safety of the operator, spotter, or patient. The component has been replaced in the field, and the replaced component is not available for further analysis. At this time, this appears to be a very isolated failure. No further action is necessary. (D. Edmonds ‚Äì 12 Oct 2017)
3466 | Hip width stuck | closed | kmettler | Adjustments - Hip | fixed | 2017-01-06 18:12 | 2017-01-04 | 250376 | 175 | Barrow Neurological Institute | Nathan West |  | (770) 867-3400 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | Issue: Hip Width will not move past 3 | Hid width tube replaced per SWO-02389. | The pin that holds the brass piece in the hip width tube walked and was grinding against the inside of the hip abduction assembly.
3463 | unit #131 - RKU Ulm - hip width adjustment is stuck | closed | bhanson | Adjustments - Hip | fixed | 2016-12-29 12:16 | 2016-12-29 | 302874 | 131 | ULM - Germany | Sebastian Lux | sebastian.lux@rku.de | +49 731 1775530 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Customer Use | customer reported that after setting hip width to maximum it can not be adjusted anymore | replaced parts, swo 2371 | Customer attempted to put a patient larger than the maximum hip width into the device. This caused the device's hip settings to bind and no longer adjust. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3462 | unit #131 - RKU Ulm - right foot drags while walking | closed | kmettler | Foot Assy | fixed | 2016-12-29 09:58 | 2016-12-29 | 302874 | 131 | ULM - Germany | Sebastian Lux | sebastian.lux@rku.de | +49 731 1775530 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported a dragging right foot preventing the step to finish correctly | Recalibrated foot sensors swo 2357 | Foot sensors require periodic calibration
3460 | Right hip is blocked | closed | kmettler | Actuators | fixed | 2016-12-21 15:54 | 2016-12-19 | 718447 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Wear | customer reported: Right hip is totally block and it' is impossible to unlock (picture attached). During the last training is seems a sensor problem, but when we try to go from standing to sitting, the leg was totally block. The patient that was using the ekso told me that he notice a weird feeling in his leg, but we check him and he is alright. | SWO-02345 - Exchange Potentiometer at right Hip, because wires where damaged under Motor cover. Couldn't exchange Hip Motor because it wasn't available at that time. SWO-2518 installs all parts which were not originally available. | Wear on ball screw caused customer facing issue. Known wear component | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 17 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/28/2017)
3458 | Knee stuck at 90 degrees | closed | lmcneil | Ball Screw | fixed | 2016-12-20 22:14 | 2016-12-20 | 812787 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Materials | Issue: Knee stuck at 90 degree angle, after talking to the PT we believe it may be a ball screw failure. | Joint board replaced per SWO-2351. | The joint board in the knee was no longer functional. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3454 | L leg seized up | closed | lmcneil | Adjustments - Leg | fixed | 2016-12-16 22:31 | 2016-12-16 | 779504 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Materials | Issue: PT‚Äôs called and the L Leg has seized up and when able to move it‚Äôs making a very loud clicking sound. We are bringing this device back to Hq for repair. | Joint board replacement per SWO-2384. | The brake switch would appear to randomly lock the brake while being depressed, however wiring and switches were not the issue. The joint board ultimately was replaced, as it was traced to be the offending part. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 574 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3453 | Hip Abduction Slipping | closed | kmettler | Adjustments - Hip | fixed | 2016-12-16 22:27 | 2016-12-16 | 37592 | 202 | Marianjoy Rehab Hospital | Shantel Firpi | SFirpi@eksobionics.com | 203-449-5009 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Design | Issue: I am doing a training at Marianjoy on their second Ekso SN 021603001 Step count :37592 For both sides of the device, the hip abduction has slipped a full number or more off of what it was set to. This has occurred a few times. Please let me know what other information you may need. Shantel | Walked customer through process to remove excess play | Particular joint will have play due to design tolerances
3452 | Step Abort | closed | kmettler | Foot Sensor | fixed | 2016-12-15 22:40 | 2016-11-17 | 812787 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: Can you look into this issue? During a max assist PT session today, we were having trouble with gait and ‚Äústep abort‚Äù was posted several times during 1:30 to 1:50 today. Clinically, there was a stall in standing on left stance leg with difficulty evoking step on right. I have clips if you need them. Tech, Colton, went through the manual re-set to adjust the foot plate sensors into normal ranges. Despite different force factors, the left heel gave an insufficient load reading. Tech took the sensor rubber plate off and the seal between the black and red surfaces was broken such that the sensor could be exposed. Thinking we need new sensor plates‚Ä¶? | Sent new parts via SWO-2278 | Known Wear Item
3448 | L Hip Slipping | closed | bhanson | Adjustments - Hip | fixed | 2016-12-15 00:03 | 2016-12-06 | 370402 | 24 | TIRR | Marcie Kern | Marcie.Kern@memorialhermann.org | (713) 799-5051 | USA | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Determined Unknown | Issue: The L hip joint on our Ekso 1.1 is having difficulty staying in place. The small screw that I tighten with the TBar as the back up to help hold the hip width adjustment doesn‚Äôt seem to be staying tight. During a session I hear creaking and popping (same as I have reported in the past) but have noticed the hip width adjustment begins to slip and at the end of a session it will have moved. For example, I have it set up at hip width of 7 and have tighten both bolts (the one with the TBar and the one with the other wrench (sorry, my brain is not thinking of the proper words at the moment)). At the end of the session the device was at 10 on the L and 7 on the R. I have been tightening each bolt 2 clicks with the torque wrenches and doing so in an alternating fashion. Is there something else I should be doing differently, or is it something wrong with the device? | Unit test walked per 103939 without issue. Work completed per SWO-2386. | Possible customer use issue, unable to replicate issue. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See D100117 - Risk Management Spreadsheet (FMEA Tab) # 169 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3447 | DUPLICATE - Broken Hip Housing | closed | kmettler | Actuators | duplicate | 2016-12-14 23:51 | 2016-12-14 | 540133 | 106 | RIC | Kristen Hohl | khohl@ricres.org | (312) 238-1000' | USA |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Undetermined | Issue: Just wanted to touch base with you after this morning's session with one of the two subjects that we were having trouble with last Friday. The first 10 minutes of walking looked normal (smooth swing phase and normal placement on the right). Then I started to noticed the same thing that I was noticing before with the non smooth trajectory and adduction whip on the right. I've included a video. -I set the hip abduction value at the # he was had previously been using. -Upon noticing this, I checked the position of the hip abduction values - and there was no appreciable change. | This is a follow up from the customer, this ticket will be closed as a duplicate entry, as the same issue as ticket #3438 | (Required for complaints)
3445 | L Hip Sticking | closed | kmettler | Adjustments - Hip | fixed | 2016-12-13 20:56 | 2016-12-13 | 109843 | 194 | Parkwood Hospital - Canada | Barry Lynam | Barry.Lynam@sjhc.london.on.ca | (519) 685-4000 Ext 42977 | Canada |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Materials | Issue: L Hip Sticking, as noted by a few PT's over the past few months. | All parts were cleaned and reinstalled per SWO-02402. | The key that locks the hip abduction to the hip width tube was binding due to material buildup on the key. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 4/21/2017)
3440 | Broken Epoxy Bond | closed | bhanson | Lower Leg | fixed | 2016-12-12 22:44 | 2016-12-12 | 109843 | 194 | Parkwood Hospital - Canada | Barry Lynam | Barry.Lynam@sjhc.london.on.ca | (519) 685-4000 | Canada |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Design | Issue: Issue: while walking a patient today they noticed the foot was rotating but the leg was staying in place, it sounds like they have broken the epoxy bond at the ankle. Patient wasn‚Äôt hurt and the PT saw it immediately and canceled the session. We are working to get this unit back to Hq for further inspection. | Shank replaced per CO-11538 during SWO-02402. | Epoxy bond on lower shank broke.
3438 | Issue w/ R leg | closed | kmettler | Adjustments - Leg | fixed | 2016-12-09 22:18 | 2016-12-09 | 423536 | 106 | RIC | Kristen Hohl | khohl@ricres.org | (312) 238-1000 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: I was hoping you could help me problem solve through a GT issue. Today we are having a bit of an problem with the right leg in that it seems to be a non-smooth swing trajectory with adduction (even to the point of stepping on the other foot). This has occurred in two users who are relatively new to the device/crutch use so at times it seems to improve when their lateral weight shift is blocked from over shifting but still occurs to some degree. I've attached a video of one of the subjects. Any ideas? '''FOLLOW-UP FROM CUSTOMER ON 12/14:''' Just wanted to touch base with you after this morning's session with one of the two subjects that we were having trouble with last Friday. The first 10 minutes of walking looked normal (smooth swing phase and normal placement on the right). Then I started to noticed the same thing that I was noticing before with the non smooth trajectory and adduction whip on the right. -I set the hip abduction value at the # he was had previously been using. -Upon noticing this, I checked the position of the hip abduction values - and there was no appreciable change. Any further thoughts? Can someone investigate our toe sensors again to see if it started to misread again? | Replacement of Hip Control Bars and a re-torquing of the tapered roller bearings in the hips. Roller bearings have finally "settled in", as this is sometimes an issue with builds, where the bearing fits snug, but does not bottom out on the hip housing. Torqued to spec. Cleaned and inspected all joints and drivetrain components, re-lubricated and tested. Completed per SWO-02396 | A combination of worn hip control bars, and slightly loose tapered roller bearings on Hip Housings.
3436 | Replacement for DDL skirts | closed | kmettler | Don-Doff Link | fixed | 2016-12-07 13:16 | 2016-12-01 | 11180 | 250 | Distributor - Italy | Matteo Carelli | Matteo Carelli <m.carelli@emac.it> | +39 347 5434417 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor | N/A - Identified known wear item | Wear | distributor reported that the device needs replacement for the DDL internal cover. | distributor will replace ddl skirts, swo 2317 | Known wear component | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Brian Hanson on 3/13/2017)
3434 | Hip abduction moves while walking | closed | kmettler | Adjustments - Hip | fixed | 2016-12-06 15:50 | 2016-12-05 | 41283 | 211 | Technomex - Poland | Michal Janus | Micha≈Ç Janus <mjanus@technomex.pl> | +48 32 23 16 660 w. 62 | POLAND |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Materials | Distributor reported that when patient walks with ekso GT the hip abductions moves. | distributor cleaned the threads on the abduction | screw on abduction was oily
3433 | Right Knee Motor Noises | closed | kmettler | Motor | fixed | 2016-12-06 00:05 | 2016-12-02 | 23556 | 231 | University of Calgary - Canada | Kyle McIntosh | Kyle.McIntosh@albertahealthservices.ca | 403-944-5358 | Canada |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: The right knee motor makes the same noise whether Ekso is powered on and legs free or if it is powered off and I am using the knee brake release. Neither 9V battery I am using is an energizer, they are both industrial Duracell but we have no issues with the left knee. When moved passively, there is no appreciable difference in resistance at the right knee vs the left. Additionally, when I walked in Ekso this afternoon I did not feel any difference side to side. Hope this information is useful. We would be able to accommodate a tech visit either this week or next. Next Wednesday we have two presentations/demonstrations of that involve the Ekso so it‚Äôs imperative that we have a functioning unit at that time. Currently our Ekso seems to function as it should with the exception of the noise. | Right knee motor was replaced per SWO-02315. | Ball screw on right knee became worn and was making an excessive amount of noise.
3432 | Ankle knob difficult to set | closed | kmettler | Adjustments - Ankle | fixed | 2016-12-05 15:47 | 2016-11-25 | 637042 | 198 | Distributor - Italy | Matteo Carelli | <m.carelli@emac.it> | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | N/A - Identified known wear item | Wear | distributor reported that it is hard to set the ankle dorsiflexion knob of the left foot. With the patient inside is a bit easier, but without is very hard to change the knob position. | Installed new components per SWO-2299 | Known wear components | No investigation needed as identified as a known wear items that are cleaned or replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler on 3/7/2017)
3431 | Footplate has separated from ekso | closed | kmettler | Foot Sole | fixed | 2016-12-05 13:27 | 2016-12-05 | 161014 | 155 | Linkoping - Sweden | Niklas Franson | Fransson Niklas <Niklas.Fransson@regionostergotland.se> | +46 10-103 15 75 | Sweden | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major | No investigation needed - see Justification | Design | The customer contacted us that the footplate somehow has been separated from the Ekso leg, that four screws for some reason has been cut off. | replaced foot assemblies, swo 2314 | Known design failure | No new investigation needed as previously investigated and deemed not a safety rick or a reportable event. See FMEA for original investigation information and mitigations put in place. (Evaluation done by Brian Hanson on 3/6/2017)
3430 | Hip width difficult to set | closed | bhanson | Adjustments - Hip | fixed | 2016-12-02 15:43 | 2016-12-01 | 329911 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | Distributor reported that it is hard to set one of the two side of the hip width setting. You have to apply force with the screwdriver to enlarge or reduce the hip width. | Visited the customer, but couldn't see any problems with the hip width. Work completed per SWO-2505. | Issue did not occur during service. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 3/2/2017)
3429 | Noise in right hip | closed | jhanff | Actuators | fixed | 2016-12-01 08:46 | 2016-11-30 | 363181 | 206 | Rolling Ball - Ireland | Kim Gregg | millsygregg@icloud.com | +44 7740 839736 | United Kingdom | no injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | N/A - Identified known wear item | Wear | Michael reported a new issue with unit 206 (see video): I was there twice already, solved some other issues and checked everything but could never create the sound by myself. The unit works well. It has no cracks in the hips. The sound appears from time to time. I torqued the screws in the sidebars holding the hip tube and also the screws from the carbon plate on the back of the unit. Any other ideas? | replaced hip tube and keys per swo 2410 | Hip tube showed signs of wear due to repetitive use. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Brian Hanson on 3/10/2017)
3428 | Strange walking | closed | kmettler | Foot Sensor | fixed | 2016-11-30 15:13 | 2016-11-29 | 1522476 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | distributor claims unit walks strange | parts replaced, swo 2307 | Known wear item
3425 | Issue with performance of device | closed | kmettler | System | fixed | 2016-11-29 18:13 | 2016-11-22 | 121087 | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com | 703-408-7713 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor |  | Customer Use | Issue: Just wanted to bring to your attention some differences I noticed with this loaner and my Ekso. This loaner is a lot quieter and smoother than mine. Also, some patients numbers are much lower in adapt and path assist etc than I get with mine. | All joint covers were removed and all internal components were inspected for wear or defects per SWO-02310. All ball screws were cleaned and re-greased. | No issues were found during service of the device that would result in performance issues. The ball screws/nuts are slightly louder than a new unit, but the noise is very minimal compared to other units with similar step counts.
3424 | Torso making a creaking sound | closed | kmettler | System | fixed | 2016-11-29 01:08 | 2016-11-28 | 493014 | 207 | Good Shepherd 3 | Alyssa Hauck | AHauck@gsrh.org | (610) 776-3247 |  |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: I noticed that our other unit which does not typically make noise has been creaking on the left hip even with the most advanced walkers. | The torso clamp bars were replaced per SWO-02327. | The bolts that secure the torso clamp bars to the tubes and torso housing became loose.
3423 | Torso making a creaking sound | closed | kmettler | System | fixed | 2016-11-29 01:08 | 2016-11-28 | 606134 | 220 | Good Shepherd 3 | Alyssa Hauck | AHauck@gsrh.org | (610) 776-3247 |  |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: I am emailing regarding two different issues we are currently having with our Ekso units at Good Shepherd. First, Charlene wanted to report that the hip of one of the Ekso units has been making an increasingly loud creaking sound which may require maintenance. | The torso clamp bars were replaced per SWO-02328. | The bolts that secure the torso clamp bars to the tubes and torso housing became loose.
3421 | No data in statistics menu after session | closed | kmettler | Software | fixed | 2016-11-28 11:20 | 2016-11-26 | 53806 | 268 | Distributor - Italy | Dr. Salvatore Calabro | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Training | PT reported: Bonino Pulejo costumer has contacted me as in their last 6 sessions the data for the Forward, Min and Path Assistance don‚Äôt appear in Statistics Menu at the end of their sessions. This issue is happing in Adaptive (both legs) either in Fixed Numbers (both legs). They are doing a Research so they do need these data. So it is a kind of urgent for them to fix this issue. | Informed the customer that 1. The assistance values will only appear when adaptive assistance is used. 2. If the assistance is changed from adaptive to any fixed level then the data values will disappear. Requested info if maybe the assistance type was changed during the adaptive sessions to a fixed level which would explain why no data was displayed. If that was not the case, I've asked for info in which sessions the problem occurred (date/time) so we can check in the transferred data what settings got used exactly. | Customer did not fully understand when the assistance values would be shown on the user interface.
3419 | unit #260 - BGU Hamburg - clasp on feet are opening | closed | bhanson | Foot Assy | fixed | 2016-11-23 12:53 | 2016-11-23 | 120000 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | +49 40 7306 - 3670 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Workmanship | customer reported that the clasps on the feet open while walking | Replaced foot binding on right foot. Work completed per SWO-2526. No parts required as the ankle assembly with the incorrect part was a replacement. The original ankle had correct component and was used for this service. The supporting documentation, part number 103939, is not available. | left foot binding mounted on right foot | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 369 for original investigation information. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/14/2017)
3416 | Foot Sensors Split | closed | kmettler | Foot Sensor | fixed | 2016-11-21 23:28 | 2016-11-21 | 3190 | 282 | Marianjoy Rehab Hospital | Danielle Yenter | dyenter@marianjoy.org | (630) 462-4000 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Materials | Issue: 2nd time the same issue has occurred, 2 weeks after a set of replacement sensors have been sent. A picture is attached - the footplate is coming unglued | Sent new sensors via SWO-2290 | Softer durometer cap sole material can more easily tear when turning on surfaces that create additional friction with the foot sensor
3415 | Shank Brace piece coming apart | closed | kmettler | Lower Leg | fixed | 2016-11-21 22:34 | 2016-11-21 | 3190 | 262 | Magee | Eric Beck | EBeck@mageerehab.org | 215-630-5926 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Workmanship | Issue: I noticed a piece to our ekso may be coming apart. Or may not be, but figured I‚Äôd see what you thought. The picture attached shows everything. The right leg along the knee in between the two leg straps is where the part seems to be separating a little. The left leg is normal but I added in the picture so you can compare the two. Does this piece need to be replaced or is it something that can be adjusted? Let me know if that makes sense and if you have any questions. Thank you for the help! | Repaired by customer with parts and tools sent via SWO-2288 | During inspection of the returned material it is probable that the main screw which had loosened was never torqued to the specified limit. Screws which are torqued to the specified limit show clear installation marking within the head of the screw whereas this assembly did not.
3414 | unit #197 - EMAC-1 Montecatone - unsual free movement in hip abduction - broken hip housing | closed | kmettler | Actuators | fixed | 2016-11-21 15:54 | 2016-11-10 | 612527 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39.347.54.34.417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | There is a lot of unusual free movement if you force a bit both legs in adduction. That's probably the reason because they have to set the patients ABD at 4 to avoid the contact between the feet. -> Udo found that both hip housings are broken, device will be shipped to workshop | replaced parts SWO-2292 | Stress within parts overcame yield strength.
3412 | unit #152 - BGU Duisburg - broken hips | closed | duanee | Actuators | duplicate | 2016-11-21 14:55 | 2016-11-21 | 566692 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | 4920376883613 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Design | Michael found at service visit that both hips are broken | replaced broken parts, swo 2291 | Stress concentration at crack origin. | duplicate
3411 | Broken Hip Housing | closed | kmettler | Adjustments - Hip | fixed | 2016-11-18 15:51 | 2016-11-18 | 535880 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 |  |  | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Design | customer reported that the feet touch each other while walking | replaced hip housings, swo 2291 | Stress concentration at crack origin.
3410 | Noise in knees | closed | kmettler | Brake / Release / 9V Battery | fixed | 2016-11-18 08:54 | 2016-11-16 | 125976 | 251 | Rolling Ball - Ireland | Colin O Shaughnessy | colin@helpfulsteps.ie | +353 87 132 6600 | Ireland |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported that both of the Ekso knees make a noise when you straighten them with the release button. | replaced 9V batteries with SWO-02280 | Known wear component
3409 | Worn soles | closed | kmettler | Foot Sole | fixed | 2016-11-18 08:49 | 2016-11-16 | 125976 | 251 | Rolling Ball - Ireland | Colin O Shaughnessy | colin@helpfulsteps.ie | +353 87 132 6600 | Ireland |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Elite gym had worn soles. | replaced foot sensors, swo 2280 | Known wear component
3408 | Broken shank brace | closed | kmettler | Lower Leg | fixed | 2016-11-18 08:30 | 2016-11-17 | 142372 | 251 | Rolling Ball - Ireland | Colin O Shaughnessy | colin@helpfulsteps.ie | +353 87 132 6600 | Ireland |  | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor |  | Customer Use | custoemr reported lots of movement in the shank brace | replaced shank brace, swo 2280 | Customer accidentally dropped the device and caused the issue.
3407 | Battery Charger LED no longer | closed | lmcneil | Battery - Charger | fixed | 2016-11-16 21:21 | 2016-11-16 | 157256 | 135 | GF Strong - Canada | Ian Denison | ian.denison@vch.ca | +1 (604) 734-1313 | Canada |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: When I put any battery on charger A the light turns from green to amber and then when the battery is fully charged the light turns green. When I put any battery on charger B the light stays green all the time but the battery does become fully charged in the same time as the battery on charger A. | Sent parts via SWO-2272 | Known wear item
3405 | unit #260 - BGU Hamburg - motor error | closed | kmettler | Motor | fixed | 2016-11-15 15:58 | 2016-11-15 | 118000 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | +49 40 7306 - 3670 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Determined Unknown | Our customer in Hamburg just called and reported that they had a motor error and needed to do an emergency sit. The patient is experienced in walking with the ekso and they never had this problem before. Unfortunately, she could not remember the error number. After the emergency sit, she restarted the device and the patient walked a bit more and the error did not occur again, but she is very concerned that it will happen again, so they only walked very short distances back and forth afterwards. The only additional info she could give , was, that the motors were not very warm, so the error could not be caused by ‚Äòover-heating‚Äô. | customer walked device with different patients and error did not occur again | Device likely produced a 133 or 121 which is known as a system limitation. These issues are one offs as it does not harm the system.
3397 | Ankle knob adjustment | closed | kmettler | Adjustments - Ankle | fixed | 2016-11-11 00:36 | 2016-11-07 | 524621 | 247 | RIM | Phuong Vu | PVu@dmc.org | 313-745-9745 | USA | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Ankle knob adjustment was stuck | Installed new components per SWO-2242. | Found that part ankle latch body, part number 103428, had fractured during use. CO-11494 corrects this design issue. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 ‚Äì Risk Management Spreadsheet (FMEA Tab) #840 for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 3/8/2017).
3396 | #121 RK | closed | kmettler | Motor | fixed | 2016-11-11 00:10 | 2016-11-10 | 36126 | 16 | Santa Clara | Tawni Ybarra | Tawni.Ybarra@hhs.sccgov.org | (408) 885-5000 | USA |  | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major |  | Materials | Issue: Yesterday Fred had a motor error (I don‚Äôt know the error code). We called and reported it to you guys so you may have the info somewhere. Today we tried to use him again and had another error ÔÅå. This time it was #121-RK. It was acting funny after we stood up the patient, it seemed like it wouldn‚Äôt sense the weight shifts. I tried changing from prostep to first step, so in first step we were able to take a right step, left step then the next right step kind of picked up the foot and put it straight down. We decided to abort the mission and sit the patient down. About ¬º way through sitting, it put itself in safe mode. The left knee was flexed and the right knee was fully extended and 121-RK was on the screen. I‚Äôm guessing that ‚ÄúRK‚Äù means ‚ÄúRight knee‚Äù. So, Fred is out of commission. If you have any advice on where we should go from here, let me or Amy Millan know! Thanks!! After over 100K steps in the last 10 months, I guess Fred is tired. | Replaced Motors per SWO-2284. | Motor upgrades to unit solved the issue reported. All soft pots and encoders checked, system recalibrated. Old Motor was flagged as failure through the ELMO.
3395 | Foot Sensors Split | closed | lmcneil | Foot Sensor | fixed | 2016-11-10 22:24 | 2016-11-10 | 3190 | 282 | Marianjoy Rehab Hospital | Kim Furman | KFurman@marianjoy.org | 630-909-7150 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: Foot Sensors is tearing A picture is attached - it's not the black treads that we previously had an issue with coming loose - it's not something that can be cut off for sure and I'm not sure about gluing it back down either. Our footplates are solid red - there is not the black treading on them - I think they were replaced previously. | Sent new sensors via SWO-2260 | Known wear item
3392 | unit #197 - EMAC-1 Montecatone - problem with the right joint cap | closed | kmettler | Actuators | fixed | 2016-11-08 16:11 | 2016-11-08 | 612527 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Design | distributor reported a problem with the right joint cap: the joint cap is a bit out than the normal position (attached picture), the PT said that it is moving if you touch it and lose grease. | distributor replaced parts (SWO-02252) Verified with new parts per SWO-2282 | Stress beyond yield strength of joint cap material was present during normal use. | No new investigation needed as previously investigated and deemed not a safety rick or a reportable event. See FMEA [list lines from Risk field] for original investigation information and mitigations put in place. (Evaluation done by Brian Hanson on 3/2/2017)
3388 | unit #185 -Technomex Ergo Hestia - broken right knee housing | closed | kmettler | Actuators | fixed | 2016-11-07 11:19 | 2016-11-07 | 855433 | 185 | Technomex - Poland |  |  |  | Poland |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | distributor reported a broken right knee housing on his customer's device (see pic attached) | replaced knee housings, Swo-2259 | Through analysis of the issue a stress concentration was found which induced the originating crack site.
3382 | Broken spacer | closed | kmettler | Other | fixed | 2016-11-02 14:54 | 2016-11-02 | 101408 | 109 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | customer reported: As you can see in attached pics there‚Äôs a damaged part to the loaner device. As we started the second session with the loaner device there was a piece falling from left hip joint (backside). This happened while leaning forward with the device to change the hip width. We found the damaged part (see pic 1455), beneath the black part of the joint. | replaced broken spacer, swo 2250 | The screw holding the joint cap in place failed. This failure was unnoticed by the customer. The customer continued to use the device and the joint shaft shifted allowing for the joint spacer to come in contact with the housing and actuator link. The actuator link then crushed the spacer into the housing.
3380 | Hip abductor slipping. | closed | ckealoha | Adjustments - Hip | fixed | 2016-10-31 16:51 | 2016-10-21 | 28933 | 281 | Franciscan Health Lafayette | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Workmanship | Right hip abductor adjustment screw will not torque. | The hip width adjustment key (103195) was replaced per SWO-02201. | The screw that threads into the hip width adjustment key (103195) was stripped and would not torque.
3377 | unit #213 - EMAC-Agazzi/Italy - error #1180 | closed | kmettler | UI - Controller | fixed | 2016-10-27 09:54 | 2016-10-27 | 134868 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Materials | distributor reported: the Ekso gave for three times error #1180. Problem was with the UI and not with the connector. | distributor sent new UI from distributor stock | UI not working correctly
3376 | Hip Width Adjustment Screw | closed | kmettler | Adjustments - Hip | fixed | 2016-10-26 21:24 | 2016-10-26 | 320000 | 173 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Materials | Customer: Warm Springs Unit #173 Step Count: 320,000 Issue: When setting up the device today the PT was adjusting the L hip width and the screw began to come out. Please see attached picture. We are sending them a loaner today and we will have their device back late this week. Thanks Josh | Entire lead screw assembly (102224) was replaced per SWO-2239. | The pin that locks the hip width adjustment screw into the lead screw sheared off.
3375 | Possible Encoder Issue | closed | kmettler | PCBA | fixed | 2016-10-25 23:21 | 2016-10-25 | 111825 | 107 | ICORD - Canada | Franco Chad | fchan@icord.org | (604) 675-8852 |  |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Wear | Issue: Today when we turned on the Ekso, it gave us many errors. I've attached the pictures here. If you can give us a call and let us know how to troubleshoot them, that will be great. FYI - we replaced the left foot sensor back in Aug and the right foot sensor about 3 weeks ago. We did the calibration, and everything worked fine until today. | Hip CCA was replaced, and errors were no longer present. Additionally, a new AMP board was also replaced, due to programming issues; not accepting new SW and calibrations would not save. All work performed per SWO-02256. | Hip CCA found to be defective.
3374 | DDL hard to open after walking | closed | kmettler | Don-Doff Link | fixed | 2016-10-25 09:00 | 2016-10-24 | 389934 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | distributor reported: DDL of this unit is very hard to open after walking. | EMAC cleaned and greased the assembly. Work is captured under SWO-2221. | wear
3373 | DUPLICATE - Software Bug after initial fix | closed | kmettler | Software | duplicate | 2016-10-24 21:53 | 2016-10-24 | 687514 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Undetermined | Issue: I just wanted to let you know that the bug (the double click when changing the assist level while walking) is still there. It‚Äôs still inconsistent, but often we do have to double click before we can adjust the level of assistance. I don‚Äôt know if you are still looking into this or if you have heard about this problem from other people. We were also wondering if it‚Äôs possible to not have that ‚Äòconfirmation screen‚Äô after changing the assist level. It slows the whole process of changing both legs down. Just a thought.. | This ticket is a duplicate of TRAC #3251. | (Required for complaints)
3372 | Restricted range of motion on R Hip Actuator | closed | kmettler | Actuators | fixed | 2016-10-24 14:23 | 2016-10-24 | 924996 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Determined Unknown | In stand-to-sit position the Ekso blocked at hip level right side. Now we can not move no longer the hip joint (see videos in WeTransfer). | replaced flexure, swo 2249 | broken flexure, possibly because of screws came loose, see attached pics
3371 | unit #163 - Mark Pollock - limping problem | closed | kmettler | Foot Sensor | fixed | 2016-10-24 14:07 | 2016-10-04 | 320949 | 163 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | customer wrote: Michael you were over with me in Dublin fixing sensors in August. That seems to have fixed the problem that cropped up at that point. Since then I have only walked a handful of times because my electrical stim was temporarily out of action and our team were on various breaks. But most of those sessions have had a slight limp from time to time. This problem came up last November and you can see Edmundo's response. I think Neil followed the instructions and solved the problem. However, I am not 100% sure that this is the same problem or the attached fix is the correct thing to do. Note, during the last session (see videos) the first few steps were really bad and over the first minute it seemed to sort itself out. But I remember that this got worse and worse in the past. ---- customer did foot test as requested, but it did not solve the issue: The right foot is pausing at toe off. We tried recalibrating the foot using the auto calibration, but this did not resolve the problem. Also, Mark has noticed a clicking sound coming from the left ankle. | foot sensor calibration per SWO-2416 | reported issue could not be found
3370 | Loose and missing screws on hip cover | closed | kmettler | Actuators | fixed | 2016-10-20 11:46 | 2016-10-20 | 351082 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Process Not Followed | customer reported: I noticed one of the little screws that secure the panel adjacent to the hip was loose. On closer inspection it appears that all the screws have been loose and consequently some appear to be missing. I have tightened all screws and I think this is cosmetic but still wanted to report. | replaced loose screws, swo 2234 | no or not enough loctide used. Not clear as customer has screwed in a few times so loctide might have dropped off.
3369 | Broken Shank, Below the Knee | closed | duanee | Lower Leg | fixed | 2016-10-19 23:19 | 2016-10-19 | 625208 | 9 | Craig | Julie Hartman | jwaldie@craighospital.org | (303) 789-8000 | United States | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major | Investigation Needed | Design | Issue: Customer: Craig Unit # 9- 1.1 device Step Count: 625,208 Issue: PT just called from Craig and reported the following issue. While walking with a patient today they were finishing their hour session and were heading back to the chair to sit... while walking straight the lower Left Leg broke making a very loud sound, the R Leg locked at the knee and the spotter was able to balance the patient on one foot and perform a long leg sit. The patient and spotter are both fine. We are sending out a loaner today and their Ekso will be back at Hq no later than Friday. | Replaced components per SWO-2286 | Part failed during normal use. Unable to find anything which would suggest that this device saw abnormal loading which would result in this failure. Failure has been classified as due to high cycle fatigue loading. | This issue is isolated to the Ekso 1.1 device. Ekso 1.2 uses a one-piece upper shank design. The risk assessment was updated again in June 2017 based on the field data. Based on the information we have at the moment, the probability of failure is low. The assessment will continue to be updated, as necessary, when more information becomes available. No further action is necessary at this point. (D. Edmonds ‚Äì 29 January 2018)
3367 | Hip Abd changing position | closed | ckealoha | Adjustments - Hip | fixed | 2016-10-17 22:13 | 2016-10-14 | 55343 | 279 | Sunnyview Rehab | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Question: One item, the hip abduction gauge on the left at times slips minimally during treatment requiring adjustment back to original setting. Common issue?? | Hip abductor assembly (102171) was replaced per SWO-02217. | This issue is currently under investigation for root cause. Exact cause of issue is currently unknown.
3366 | Hip lock screw is stiff | closed | duanee | Adjustments - Hip | wontfix | 2016-10-17 21:57 | 2016-10-14 |  | 292 | Franciscan Health Indianapolis | Amy Horter | amy.horter@franciscanalliance.org | (317) 865-5000 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Design | Issue: Hip lock screw is stiff I am at the Minneapolis VA for advanced training We are using Ekso SN 021602002 We are seeing a slip of the left hip abduction setting one full number. This is the same side the has a previous issue with Thank you, Shantel | Not applicable, resolved as 'wontfix'. | Known issue caused by material buildup in the abduction assembly causing the key to stick to the hip width tube. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 393 for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 5/24/2017)
3365 | R Hip abduction shifting | closed | kmettler | Adjustments - Hip | fixed | 2016-10-17 21:40 | 2016-10-14 | 165090 | 263 | Rehab Institute of Kansas City | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Issue: R Hip abduction shifting | Replaced 102224 assembly. Replaced per SWO-02188. | Design issue, as reported by engineering. More pronounced in some units, no real screening other than standard test method.
3362 | duplicate-Batteries no longer working | closed | jhanff | Battery | duplicate | 2016-10-14 17:46 | 2016-09-27 | 763125 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Undetermined | Issue: Apparently the batteries in Herentals are fast deteriorating as two batteries are merely enough to let one patient complete a one hour session. Can something be done to help that and/or prevent a fast loss of battery power ? I don‚Äôt think that this problem is caused by installing SA software. Morely, I think this is a consequence of the intensity we work with the batteries. Every day they run up to 6 times empty (less than 25 %). We have to recharge them 6 times a day. It occurs with both GT‚Äôs and with the 4 pairs of batteries. | Sent new batteries via SWO-2200 | (Required for complaints)
3361 | Clicking in left hip | closed | kmettler | Adjustments - Hip | fixed | 2016-10-14 14:44 | 2016-10-14 | 679984 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | They hear clicking sound in left hip. This is not happening all the time but only from time to time on random patients. Can I kindly ask you to look into data that unit is sending and letting me know if you see anything unusual? | Walked the unit with a patient, checked motor, rail, and hip abduction. Cleaned needle bearings on hip, greased the bearings and reinstalled them. Work completed per SWO-2222. | Device required lubrication within hip joint.
3360 | unit #245 - Ostracon/Anagennisi-Greece - problem with soles | closed | kmettler | Foot Sole | fixed | 2016-10-14 12:38 | 2016-10-13 | 18739 | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | black rubber is coming off on the left foot | New foot sensors installed per SWO-2199 | Known wear failure/component
3359 | Hip abduction setting changes during use | closed | kmettler | Adjustments - Hip | fixed | 2016-10-14 12:31 | 2016-10-13 | 18739 | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Info from Michael: Anagenisi still has an issue with the settings of the hip abduction. I cleaned the screw and added some vibra tite to it which worked for some weeks but now the settings change again while walking (both sides). | New hip abductions installed per SWO-2199 | Found excessive wear in M4 screw threads
3354 | Broken joint cap on right hip | closed | kmettler | Actuators | fixed | 2016-10-12 14:56 | 2016-10-12 | 37987 | 209 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Major |  | Design | distributor reported: I needed to replaced right hip cup (see pics attached) | Replace joint cap and screw with parts provided via SWO-2211 | Screw spec'ed for this assembly can break under normal loading
3353 | Foot sensors failed | closed | kmettler | Foot Sensor | fixed | 2016-10-12 11:26 | 2016-10-12 | 570416 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | distributor reported: After the distributor replaced the upgraded ankles, he did the foot cal and the test failed. So he needs to replace the foot sensors. | Distributor replaced Foot Sensor - SWO-02194 | Foot sensors are a known wear item.
3352 | Legs rotating Inwardly | closed | kmettler | Lower Leg | fixed | 2016-10-11 23:28 | 2016-10-11 | 260557 | 59 | Able Bionics - Aspen | Amanda Boxtel | amanda@bridgingbionics.org | (970) 379-9260 | USA | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major |  | Materials | Issue: Thank you for the newly calibrated T-wrenches. Sadly, the lower leg shanks (especially left) continue to turn inward after about 20 steps. I walked this morning and cut my session short because we kept having to stop, turn the foot out and tighten it down again. What do you advise at this point? Is it possible to get new lower leg shanks sent to us? Would that help resolve it? Please email or call. We have four clients scheduled to walk next Monday. Help. | Replaced components per SWO-2210. | Found that the part did was not conforming to specification of 5 degree +/- 1 degree offset for the head of the clamping screw.
3351 | R step not initiating | closed | kmettler | Foot Sensor | fixed | 2016-10-11 17:01 | 2016-10-11 | 150197 | 135 | GF Strong - Canada | Ian Denison | ian.denison@vch.ca | +1 (604) 734-1313 | Canada |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: One of the PT's from GF Strong called the red line this morning. It sounds like they may be having issues with their foot sensors. With their first patient this morning after they completed the right step, the left step would not initiate. He also tried w/out the patient in the device and this was still occurring. They don‚Äôt have any other patients today. | Sent foot sensors via SWO-2191 | Known wear item, will not be receiving them back for RCA
3350 | Problem with batteries | closed | kmettler | Battery | fixed | 2016-10-11 08:30 | 2016-10-10 | 39406 | 162 | Hobbs - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | customer reported: we had a fault with a set of our batteries today I've tried charging the set (1A & 1B) twice but when connecting to the Ekso it beeps and reports 'critically low batteries 3%'. I assume it's an error as I charge them everyday and rarely let them go below 40%. | called customer and advised to check charger cables | cable not plugged in correctly
3348 | Cracked flexture | closed | kmettler | Ball Screw | fixed | 2016-10-10 19:35 | 2016-10-04 | 462792 | 115 | NRH - Med Star | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Materials | Issue: strange noise coming from Ekso device - ended up being a cracked flexture | Flexure replaced per SWO-02108. Rail and carriage, spoons and motor were also replaced. Issue was not resolved until the flexure was replace. | Material failure. There was a crack in the flexure.
3346 | unit #213 - EMAC-Agazzi/Italy - noise in left leg | closed | jhanff | Lower Leg | fixed | 2016-10-07 14:25 | 2016-10-07 | 105005 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor | No investigation needed - see Justification | Customer Use | customer reported a noise when moving the left leg (see attached video) | Directed customer to remove label. | Customer installed a label on the knee of the device which caused the audible issue when using the device. | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/7/2017)
3345 | Problem with moving left leg forward | closed | kmettler | Foot Sensor | fixed | 2016-10-07 13:13 | 2016-10-07 | 105005 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | distributor reported: the Ekso in Agazzi has a problem with the forward movement of the leg left. They heard a noise and difficult to move forward the leg. This happen with different patient and also when they tried it. I follow them to perform the foot sensor calibration and it passed. I‚Äôll send to you a video with a second mail. | Walked customer through recalibration process | Recalibration of foot sensors resolved issue
3343 | Binding strap retaining rod | closed | lmcneil | Foot Assy | fixed | 2016-10-06 21:46 | 2016-10-06 |  | 161 | Glenrose - Canada | Vickie Buttar | vickie.buttar@albertahealthservices.ca | 780-735-6207 | Canada | no injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Customer Use | During setup of the Ekso the PT noticed the retaining rod had pulled through the hole, see picture. | SWO 2183 was generated and parts were sent to the customer to replace broken foot binding. | Customer applied too much force causing part to fail.
3341 | Broken joint cap | closed | kmettler | Actuators | fixed | 2016-10-05 16:02 | 2016-10-04 | 399301 | 46 | Guttmann - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Design | customer reported a broken joint cap on the left - see pic | Install new Joint caps and exchange Joint cap screws for new type. Work done via SWO-2202. | Joint caps are still the old type. Preventive Maintenance and Close open CO's has to be done.
3340 | unit #243 - Step by Step/Barcelona - defective battery contact left | closed | bhanson | Battery - Connector | fixed | 2016-10-05 15:53 | 2016-10-03 | 605927 | 243 | Step-by-Step - Spain | Joan Romero | jromero@fundacionstepbystep.org | +34 93 223 16 63 | Spain | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | N/A - Identified known wear item | Customer Use | customer reported a defective battery contact (shows always battery load 0%) and has not broken/bent pins | Exchanged left battery connector per SWO-2525. The supporting documentation, part number 103939, is not available. | Customer installed battery incorrectly and damaged the battery interface | No investigation needed as identified as a known wear items that are replaced periodically at the request of the customer. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 3/14/2017)
3339 | Foot plate separated from ankle | closed | lmcneil | Foot Assy | fixed | 2016-10-04 00:13 | 2016-10-03 | 281000 | 164 | VA Minneapolis | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Design | Issue: Issue: PT called the redline and reported the foot plate separated from the ankle. The PT immediately saw this occur and was able to sit the patient safely‚Ä¶. Since this is a VA they can‚Äôt send pictures, we are sending them a loaner arriving tomorrow and we‚Äôll have their device back later this week. I‚Äôll update this email if any pics do come through from the PT. | Foot plate (102280) and ankle housing (103426) per SWO-02180. Refer to TRAC # 2568 for more details on resolution. | Design issue. Refer to TRAC # 2568 for more details.
3338 | Ankle dorsiflextion changes | closed | kmettler | Foot Assy | fixed | 2016-09-30 13:45 | 2016-09-30 | 556463 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor |  | Materials | distributor reported: Ekso Montecatone has the ankle dorsiflexion that change alone during walking. | Distributor replaced both ankles (SWO-02177, parts taken from EMAC storage) | Shank housing was found to be broken
3337 | Step Aborts | closed | lmcneil | Foot Sensor | fixed | 2016-09-29 18:28 | 2016-09-28 | 508714 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: ‚ÄúToday we received several step aborts (~5) during our brief session using the GT. Each time they occurred when in stance on L and beginning swing phase on R. All situations where in which I felt the patient was in a good position and well balanced. I am wondering if there may be a malfunction of one of the foot sensors?‚Äù | Sent new parts via SWO-2165 | Known wear item
3336 | Foot Sensor no longer engaging | closed | kmettler | Foot Sensor | fixed | 2016-09-29 18:13 | 2016-09-28 | 254282 | 59 | Able Bionics - Aspen | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: R Foot senor is no longer engaging. Due to high use Able Bionics has high turnaround on their sensors | Sent parts via SWO-2163 | Known wear item
3331 | Joint cap failure | closed | ebabaran | Actuators | fixed | 2016-09-27 22:04 | 2016-09-27 | 575632 | 220 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Issue: Joint Cap fell out while patient was walking straight ‚Äì in the middle of their session. Please see attached pictures. | Joint caps replaced per SWO-02185 | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3330 | Brake Slip | closed | kmettler | Brake / Release / 9V Battery | fixed | 2016-09-27 18:48 | 2016-09-19 | 3909 | 281 | Franciscan Health Lafayette | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: The customer reported that the Brake is slipping. | Left brake hub replaced with new materials. | Left brake hub was not sitting completely flush with the brake. Brake failed the 12-point torque test.
3329 | Error #137 RK | closed | lmcneil | Brake / Release / 9V Battery | fixed | 2016-09-27 14:47 | 2016-09-27 | 15194 | 268 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | distributor reported that this Ekso during training last week had a couple of time the error # 137 RK. Distributor followed by phone our PT to do the joint cal, but during the last two days is happened again for three times. With different kind of patients. | Exchanged knee motor for new one and greased ball screw. Work completed per SWO-2524. The supporting documentation for this service, part number 103939, is not available. | Parts were not returned to HQ for analysis. Unable to accurately determine root cause. | No new investigation needed as previously investigated worst case scenario of knee motor failure and deemed not a safety risk or a reportable event as safeties are already in place. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 622 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3328 | Noise from ball-screw on left knee | closed | kmettler | Ball Screw | fixed | 2016-09-27 14:40 | 2016-09-27 | 72926 | 266 | OrthoMedico NV - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported a noise in left knee | Replaced knee motor, SWO 2219 | Ball screw performance degraded overtime resulting in unusual sounds
3325 | Crack in Hip Housing | closed | duanee | Adjustments - Hip | fixed | 2016-09-26 18:18 | 2016-09-15 | 291150 | 181 | VA Bronx | Steven Knezevic |  | 718-584-9000 x3130 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major | No investigation needed - see Justification | Design | Russ found that the R hip housing was compromised in the field. | Hip housings replaced per SWO-02154. | Hip housing returned from field and given to engineering for further evaluation. Known issue with crack on left hip housing. Observation of walking pin on right hip housing gouging both inner walls | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 390 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 5/17/2017)
3324 | Request for Torso Pad's | closed | kmettler | Soft Goods - Torso / Sacral | fixed | 2016-09-23 19:34 | 2016-09-16 | 270005 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: MJ's soft goods were beginning to wear and they requested new Torso Pad's (highly used item) | Sent new parts via SWO-2152 | Known wear item
3322 | Charger does not charge battery to full even with green light | closed | kmettler | Battery - Charger | fixed | 2016-09-22 20:34 | 2016-09-22 |  | 279 | Sunnyview Rehab | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | Email from Melissa: Team, For past 3 days at Sunnyview, we have had seen 'fully charged' batteries placed on Ekso reading 89 or 90%. Inconsistent with which battery or side of Ekso, but appears consistent with one charger. I have labeled this Charger A. See pic attached. Can we send a replacement charger? Additionally, the customer has inquired we haven't labeled the chargers, in the same way we label the batteries, so they can keep track of which charher is which. Jeff, this has come up in the past as well. Is there a plan to label chargers on shipment, or can we make this a feature request? Please let me know if there is anything else you would like me to look at. Thanks, Melissa | New charger shipped out via SWO-2157 | Not financially feasible to investigate replacment accessory
3321 | Noise in the joints of the knees | closed | kmettler | Brake / Release / 9V Battery | fixed | 2016-09-22 08:10 | 2016-09-22 | 399301 | 46 | Guttmann - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Training | Our customer Institut Guttman in Spain, has a problem with their device (Ekso sn/011211003 (AM 476)). They said it makes a strange noise in the joints of the knees. | The "Noise"was finally the empty battery in the Knee (9V Block) After exchanging both of them the noise was gone. No SWO created device was working as designed. The change out of the 9V battery is a user defined task. | Customer was not aware that they could change out the 9V battery on their own to resolve this issue.
3320 | Loud clicking in the left hip after fall | closed | kmettler | Other | fixed | 2016-09-22 06:51 | 2016-09-20 | 1040563 | 147 | Mark Beack - South Africa | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Customer Use | issue reported from the PT: I unfortunately had my first fall with a patient last week and it seems the EKSO has been very slightly damaged. The patient is a tall and heavy cervical SCI. We were in fixed assist mode and the patient leant aggressively to the right before the foot had been firmly placed on the floor. I managed to pull him back but just not enough to stop us falling very slowly to the right. All seemed fine with the EKSO however the left hip started clicking loudly in the next session. I think the Ekso pivoted on the left hip as it fell to the right which put excessive force into the joint. The Ekso is working well except for the clicking however it does not want to activate the turning mode to the right (seems the left hip is not sensing correctly). 2016/09/30: Update from Michael: I am afraid a screw moved into the rail. Please stop using this unit. We need to repair it as the screw might move out and destroy the rail. ---- Additional info provided regarding the fall: Patient details: C3/C4 incomplete (ASIA B) traumatic SCI Mode of operation of device: Prostep, bilateral fixed assist on 80% with crutches (hands bound to crutches with Active Hands) Patient injury (if any): None Short description of how the fall occurred: The patient is a tall and heavy cervical SCI. We were in fixed assist mode and the patient leant aggressively to the right before the foot had been firmly placed on the floor. I managed to pull him back but just not enough to stop us falling very slowly to the right. He lay on his right side before turning him to lay him flat on his back to remove the Ekso. Patient was laughing after the initial shock of the fall. We are both fine, it is more my ego that is bruised! We are carrying on with business as usual, the patient has had two successful walks since the fall. You are probably wondering why I was using fixed assist with such a high level lesion patient. I found with this patient that he became lazy with maximal assistance. When he uses more arm strength to maintain a good posture and forward momentum then the leg can get through in fixed mode. I find this very useful to gauge how much effort the patient is putting in. The obvious danger is when the leg ‚Äúhangs‚Äù when a step cannot be completed at the set percentage. I have gone back to variable assist for now but I have found there is nothing that motivates a patient more to work harder than having the leg stop mid-swing as they have not contributed enough to the step. | Both Hips where exchanged with complete new Hip assemblys per SWO-02264 | Device was dropped.
3314 | Right step not initiated | closed | kmettler | Other | fixed | 2016-09-19 15:10 | 2016-09-14 | 763125 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Mishandling | Supporting on the left side, the ekso hesitates to initiate a step on the right side‚Ä¶ We have tried it with different patients and when they are walking in the other devices there‚Äôs no problem‚Ä¶ | flat white cable in the knee had a sharp bend (see pic) and caused a defective connection replaced flat cable, swo 2158 | The flat flex cable was found to be damaged which induced this issue. It is not possible to determine what damaged this cable.
3310 | Free leg heating up | closed | kmettler | Lower Leg | fixed | 2016-09-14 22:21 | 2016-09-14 | 652372 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | Issue: Issue: while walking a patient, R side freed, the free leg is heating up and they can feel the heat (no errors thrown). They thought this was unique as the motors on the free leg are heating up but the other side under power is fine. Does anyone have insight on why this may be happening? I have an email into the customer to get more pt details and a possible video. Eric, is this something we can look at through their data? They send consistently and the pt ID is 0066. | Re-lubricated all internal components on hips and knees as a maintenance procedure. Replaced all motor and flexure bolts as a preventive measure. No issues found. Work performed once unit was in house, under SWO-2384 | Even the "free leg" setting does use the motors as a dampening and a soft lock , so it is possible that the patient was leaning heavily into an out-of-range scenario for normal step tracking. This would cause the motor to hold an end-position without the brakes, therefore heating is generated. No error codes, within normal operating specifications. No issue found.
3308 | joint cap screw fell out | closed | kmettler | Actuators | fixed | 2016-09-13 18:14 | 2016-09-13 | 128590 | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com | (703) 408-7713 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Major |  | Design | Issue: see attached pics. I was under the impression we had fixed this issue and rolled this out as part of the maintenance program when we touch all units. Is it now the case that joint caps can still fail even after we've applied the most recent fix? Please advise. | We sent parts via SWO-2146 so that the customer can implement the fix. | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3307 | joint cap fell out | closed | kmettler | Actuators | fixed | 2016-09-12 23:25 | 2016-09-12 | 272859 | 42 | Montreal - Canada | Martin Vermette | martin.vermette.physio@gmail.com | (438) 826-0595 |  |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Major |  | Design | Issue: You will find attached a Word document with photographs of the bolt that got dislodged from the bearing assembly of the left lower hip joint of our Ekso. According to our technician, part no. 2 was probably originally attached to part no.3 (the ring). If you opte for sending us replacement parts, he would feel at ease to put them in place. We could call you for technical assistance if needed. If so, he wanted to know if a certain torque force needs to be applied to the bolted assembly. Also, we were wondering if this problem could occur to the other hip joint eventially and if it would be safer for you to send us 2 replacement bolts - rings. Finally, if you decide to go that route, let me know. I will then confirm to you our best shipping address; Parcels have a tendency to get lost in our institution! Thank you for your assistance. | Parts sent via SWO-02141. | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3302 | unit #193 - Technomex/Poland - broken joint cap | closed | kmettler | Actuators | fixed | 2016-09-09 12:02 | 2016-09-08 | 167614 | 193 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Major |  | Design | customer reported broken hip joint cap problem to his distributor | distributor replaced it adnd walk device (500 steps) from distributor stock. | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3300 | unit #209 - Technomex - replacement of foot sensors | closed | kmettler | Foot Sensor | fixed | 2016-09-09 11:11 | 2016-09-08 |  | 209 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor |  | Wear | Distributor reported that their device SN 021504002 had problem with right foot sensor. I replaced sensor and made calibration. Device is ok. | distributor replaced sensors with material from distributor stock. | Foot sensors are a regular wear item needing either periodic calibration or replacement.
3299 | unit #208 - EU Loaner currently at EMAC/Agazzi - broken joint cap on right hip | closed | jhanff | Actuators | fixed | 2016-09-09 10:50 | 2016-09-09 | 235731 | 208 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Major |  | Design | distributor reported a broken joint cap on the right hip | Shipped replacement parts to distributor with SWO-02134 | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3298 | unit #006 - Thevenaz/Switzerland - DDL stuck on right leg | closed | ckealoha | Don-Doff Link | wontfix | 2016-09-08 09:36 | 2016-09-07 |  | 6 | Thevenaz - Switzerland | Benoit Thevenaz | bthev111111@hotmail.com | 41787126152 | Switzerland | no injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | Customer reported a stuck DDL on the right leg. It is not possible to open it anymore at all. | Resolving as resolved:wontfix. | This cannot be explored further until we have more information. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437V ‚Äì Risk Management Spreadsheet (FMEA Tab) #212-C for original investigation information and mitigations already in place. (Evaluation done by Brian Hanson on 5/23/2017).
3297 | unit #245 - Anagennisi/Ostracon-Greece - defective battery charger | closed | jhanff | Battery - Charger | fixed | 2016-09-07 10:04 | 2016-09-06 | 62677 | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | distributor reported a defective battery charger for his customer. It is not loading correctly anymore, green light is not changing to yellow. | Sent new battery chargers via SWO-2131 | chargers not being returned.
3296 | unit #208 - EU Loaner currently at EMAC/Agazzi - broken foot sensor | closed | jhanff | Foot Sensor | fixed | 2016-09-07 09:47 | 2016-09-06 | 235731 | 208 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Major |  | Design | foot sensors broke again after just two days of usage - see pic attached reply of the distributor to our request if customer might not use the turn mode correctly, because it is the fourth time in a short time period that foot sensors needed replacement at this customer's site: The customer is trained to the use of turning mode and watching them, the turning operations are made correctly. Every time that I visit a customer I attend a session with Ekso, to check them and the device. They use it only in a large and very long corridor (about 170 steps) used to walk with the device, so also the turning operation are few than other our centers. Last Friday I were there to replace the ‚Äúold type‚Äù of foot sensor on this Ekso because it doesn‚Äôt work well. But I replaced it with the ‚Äúnew type‚Äù that actually we have in our stock. I personally feel the difference when I try the device or when I make the Auto Foot Test. If you put on the ground the two type of foot sensor (not yet attached on the device)and try manually to turn, you can feel the difference. During a normal straight walking you can hear the classical noise of rubber that rubs with the floor. If you turn, you feel very well the difference. Less relevant but different, are also the type of red color and the smell of the two kind of sensor. | Replacement treadless foot sensors to be sent to customer via SWO-2130. | Known issue with foot sensors - black tread catches and starts to come off.
3295 | unit #232 - To Walk Again/Belgium - foot sensor broken | closed | kmettler | Foot Sensor | fixed | 2016-09-07 09:35 | 2016-09-06 | 282087 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | there‚Äôs a problem with the left foot sensors because the right step isn‚Äôt initiated by stance support on left side and the protection at frontside of the left foot is loosen | replaced foot sensors, swo 2136 | Foot sensors can fail over time with regular usage.
3294 | DUPLICATE: torn foot sensor sole | closed | kmettler | Pick an option | duplicate | 2016-09-06 19:02 | 2016-09-06 |  |  | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Undetermined | Foot sensor torn at left sole after only 3 days of use. Details to follow from customer. | duplicate ticket entered #3296 | duplicate ticket entered #3296
3293 | unit #258 - PASORI/France - defective battery charger | closed | kmettler | Battery - Charger | fixed | 2016-09-06 08:30 | 2016-09-06 |  | 258 | CRF Pasori - France | Fr√©d√©rique Bordet | f.bordet@pasori.fr | +33 3.86.28.43.89 | France |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | We have a problem with the charger, one of the connector is pushed down and therefore does not connect and charge the battery anymore. We need a new one. | mailed replacement charger to customer, swo 2129 | Charger will not be returned for root cause analysis as not financially viable. Will revisit if this becomes a recurring problem.
3292 | unit #213 - EMAC-Agazzi/Italy - broken shoulder strap | closed | kmettler | Soft Goods - Torso / Sacral | fixed | 2016-09-02 14:22 | 2016-09-02 |  | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | +39 010 3290411 | Italy | no injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | No investigation needed - see Justification | Wear | distributor reported a broken shoulder strap | distributor replaced parts from distributor stock. | Known wear component | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 498 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/7/2017)
3290 | Right knee cover broken | closed | duanee | Actuators | fixed | 2016-09-02 08:27 | 2016-09-01 | 85469 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | 040/73063670 | Germany | No injury | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor | Investigation Needed | Determined Unknown | Michael reported from service appointment that the right knee cover is broken | replaced parts, swo 2969 | It is not possible to do a deep root cause analysis on this because the part that was replaced cannot be analyzed further. | A crack in the right knee cover was detected during service for this device. It has been determined that this does not present a safety hazard, and the part has been replaced. The original knee cover is not available for further analysis. No further action is necessary. (D. Edmonds ‚Äì 02 Oct 2017)
3286 | Possible Pot Encoder Issue | closed | kmettler | Encoder / Potentiometer | fixed | 2016-09-01 22:58 | 2016-09-01 | 35446 | 227 | Children's Hospital of Los Angeles | Lauren Dunn | bhanson50@gmail.com | 6162935907 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | Issue: They had several sensor errors (25) come up today. I walked them through a leg test in the morning and around 2:30, both times it passed. With different patient‚Äôs and PT‚Äôs walking both times. One of the PT‚Äôs also said another error message came up but couldn‚Äôt remember if it was 50 or 51. Can we look into the data to make sure nothing else is going on? Additionally both errors happened when the patients were stepping forward w/ the R Leg. This morning they were in left affected mode & in the afternoon w/ right affected mode. | The contamination was cleaned using IPA, per SWO-2140. | There was a contamination on the pot that was causing the a dead spot.
3285 | Foot Sensors Dying | closed | kmettler | Foot Sensor | fixed | 2016-09-01 20:39 | 2016-09-01 | 189217 | 222 | Kennedy Krieger Institute | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: It seems like our left sensor may not be functioning Their heel on the L side is no longer working as well as the toe on the R side So we checked out the sensors, and here were our values. Left ‚Ä¢ Hip: -320 ‚Ä¢ Knee: -20 ‚Ä¢ Heel: 1 ‚Ä¢ Toe: 515 Right ‚Ä¢ Hip: -333 ‚Ä¢ Knee: -4 ‚Ä¢ Heel: 280 ‚Ä¢ Toe: 40 | Sent new sensors via SWO-2124 | Known wear item, however cannot perform true RCA as we will not be getting the sensors back
3284 | Capacity of batteries are declining | closed | kmettler | Battery | wontfix | 2016-09-01 11:01 | 2016-09-01 | 500000 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Determined Unknown | customer called, because he thinks that the capacity of the batteries is declining. As they plan to use the device more in the future, they want to make sure that the batteries won't make any problems and so asked for a check of them now. He could not give any exact data how long they can currently work with one set, but it is much less than in the past. | Performed Elmo Calibration and Leg test. The aforementioned calibrations and tests did not resolve the issue as presented to the customer. Over time the customer has reported that the issue is no longer present. | Issue is no longer present | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler on 4/6/2017)
3283 | Left Leg Extends Backwards | closed | kmettler | Foot Sensor | fixed | 2016-08-31 23:14 | 2016-08-31 | 245314 | 23 | NeuroFit360 | Guy Romain | guypt@neurofit360.com | (954) 252-8020 | USA |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: I am writing this email from NeuroFit360 in South Florida. We are having an issue with our eksobionic device that we need your help to fix. When we are walking with someone in bilateral mode and the machine is ready for left step, the left leg extends backwards before going forwards. This occurs maybe 2 out of every 5 steps, on rarer occasions it extends fully at the hip & knee which causes the person to bend forward. | Sent customer new sensors via SWO-2311 | Most likely a wear item, however we will not receive the parts back to perform a proper RCA
3282 | unit #208 - EU Loaner currently at EMAC/Agazzi - broken foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-08-31 15:12 | 2016-08-31 | 228140 | 208 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | the distributor reported defective foot sensors for the loaner at his customer's site. | distributor replaced foot sensors - SWO-02123 | Known wear item.
3281 | Don-doff Link opens while walking | closed | kmettler | Don-Doff Link | fixed | 2016-08-31 14:53 | 2016-08-31 | 96475 | 109 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Major |  | Wear | customer emailed that this unit is starting to get problems due to the using. One of the locking security systems of the left hip sometimes unlock. info from further check on the device: right ddl: difficult to move, does not spring back; pin sitting very low left ddl: working ok on first view, but pin sitting a bit too high | Exchange both upper shanks and complete lower feet. SWO-2027 | Components had begun to show wear which induced this failure.
3280 | Software Issues with assistance values | closed | kmettler | Software | wontfix | 2016-08-30 22:33 | 2016-08-29 | 638768 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-597-2347 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor |  | Design | Issue: After Burke‚Äôs device was upgraded to the new software, they have had issues with changing the assist levels. They stated that it has been a constant problem. They have to hit the check twice consistently, every now and then (If you just go into the assistance screen) they only have to hit it once. They said it has happened with everybody who has been in the device, since their update on 7/28. Adam has started to look into this issue and said that it was most prevalent when changing the assist levels while walking. I will reach out to the customer and find out how much of an inconvenience this is. | This is an inconvenience for the customer, they will deal with it until next service. | Adam looked over data from Burke‚Äôs device and can definitely see them pressing the check button twice to start editing assistance values. Some notes on what‚Äôs happening: ‚Ä¢ It‚Äôs not every instance, but ~50% of them ‚Ä¢ All the instances I saw were when they were modifying the assistance values while walking (although I didn‚Äôt see them attempt to modify the assistance much prior to walking in the 2 sessions I looked at) ‚Ä¢ Technically the first button press triggers an edit, but the device is immediately exiting the edit screen before the customer even sees it, so it appears like the first check does nothing. After reviewing the code, I believe this is an issue with the RC5 version of the code sent to the researchers for IRB use prior to release. I can‚Äôt find the same bug in the RC6 version of code released. And the issue wasn‚Äôt noticed during the Verification of the RC6 code. If the inconvenience is significant enough, then we can update their software to RC6 once we have the RC6 compatible FES cables in-house. Prior to doing this, please give Software Engineering a heads-up, so we can double-check this diagnosis prior to shipping the device or service member to update the software.
3278 | Adjustment knob | closed | kmettler | Ankle | fixed | 2016-08-29 13:44 | 2016-08-29 | 591526 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | customer contacted us as they had a problem with their unit after a short walk. The spring retention screw on the ankle adjustment broke. | Work completed per SWO-2128 Took out the adjutment mechanism from the Loaner and repaired with it the Unit 243, as the Loaner goes to Depot where we exchange the feet anywy due to close open CO's | Previously the design called for Loctite 243 with provided a bond which was not strong enough for the assembly. CO-11185 changed this requirement to Loctite 263 which is suitable for this process.
3277 | Issues with the controller | closed | adamzoss | Software | fixed | 2016-08-26 19:49 | 2016-08-24 | 633166 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: We aren‚Äôt sure if it is an issue from the controller or just the way they changed to robot, but if we want to change the assist level, we have to press the check button twice, sometimes once, so it‚Äôs inconsistent.. * This device has smart assist. | Sent new controller via SWO-2102 | Controllers are known wear items. Did not request back so unsure of root cause.
3276 | Error 1111 | closed | ebabaran | UI - Controller | fixed | 2016-08-26 19:45 | 2016-08-26 | 157358 | 196 | Children's Hospital of Atlanta | Amber Richards | arichards@chla.usc.edu | 323.361.2568 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: PT reported after a session this morning they got several software errors in a row and they were looping not allowing them to clear and return to walking. The one # they caught was a 1111 error among others but they didn‚Äôt remember the #‚Äôs associated. | New UI shipped with SWO-02107 | Known wear item, however getting back for additional RCA. Dan Brown - 8-31-2016 - Adding links to bad UI handset (controller) in EE/ME engineering database: https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2925 https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2765
3275 | Repeated error 137 LK | closed | duanee | Brake / Release / 9V Battery | fixed | 2016-08-26 13:39 | 2016-08-25 | 4940 | 267 | Jannah Armstrong - UK | Liz Armstrong | theheath@btinternet.com | +44 1886833551 | United Kingdom | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Wear | We had a repeated 137 LK error today: First occurred in ProStep on Bilat Max, after poor facilitation but no misstep ( error appeared on heel strike/loading - after sitting and rebooting it appeared ok Second occurred in First Step Bilat Max, patient was walking well/good facilitation, again happened on heel strike/loading ‚Äì same as above cleared error Third time occurred in First Step Bilat Max, good facilitation, happened later in the cycle during mid stance. After the third error we abandoned the session. I jumped into the Ekso and we tried various ways of replicating the error, but unsuccessful. On discussion with Michael he wondered as the error is triggered by the mild vibration of the floor (the therapy hall has a hydrotherapy pool underneath with a movable floor which allows for mild bounce). When retrying the Ekso with me in it and assistants jumping on he floor to increase the vibration, we still could not replicate the error. Please note, this is the same error they had a couple of months back during the second day of ITW ‚Äì it was fixed on site by Michael. After discussion with Barry and Juliana the plan for now is: I will return on Tuesday next week and we will send a loaner ‚Äì if the same error occurs with the loaner in the same therapy hall, then we can work on the hypothesis that the vibration of the floor is to blame. Should the error not occur with the loaner, then maybe we need to reexamine the device. | cleaned brakes, burnished brakes. Verified that issue was no longer present per SWO-2257. | reduction in holding power of brake pads (possibly oily) | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Brian Hanson on 6/8/2017)
3274 | Problem with hip abduction lock | closed | kmettler | Adjustments - Hip | fixed | 2016-08-26 13:15 | 2016-08-26 | 525763 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | distributor reported from customer site that even after replacing the spring plunger the abduction lock opens while walking | Michael sent new revision control bar to distributor for replacement via SWO-2104 | Previous designs of this part allowed for a small portion of parts to unlock. Design was corrected with CO-11310.
3273 | Problem with right battery connector | closed | kmettler | Battery - Connector | fixed | 2016-08-26 12:19 | 2016-08-26 | 109213 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | customer contacted distributor because of a defective battery connector, pic shows bent, twisted pins | Michael send parts to Udo - SWO-02106 Exchange both battery connectors and close all open CO's. | Pins are known to bend if force applied.
3271 | Errors #133 and #121 | closed | kmettler | Electrical Interconnect | fixed | 2016-08-25 23:24 | 2016-08-25 | 394102 | 165 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Determined Unknown | Issue: One of the PT's called after their morning session. They were unable to use the device. With their first patient, the attempted a sit to stand several times and the device stopped 3/4 of the way up. The PT then attempted to stand the device by itself and it immediately through an error #133 and then an error #121. We had them attempt a joint calibration, it passed successfully, however once she tried to stand the device on it's own, she was brought back to "initiation" and was unable to free the legs or use the device. | Using a set of batteries that is above 30% charge prevents the error from occurring. Work conducted on device via SWO-2138. | 121/133 errors ocurred as designed. Unable to reproduce reboot issues seen at customer.
3270 | Multiple Step Aborts | closed | kmettler | Foot Sensor | fixed | 2016-08-25 17:20 | 2016-08-22 |  | 220 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: GS reported multiple step abort errors, walked them through a foot cal, failed the calibration. | We sent new foot sensors, customer replaced and successfully calibrated the device. Parts sent via SWO-2095 | Foot sensors are a wear component
3268 | Fall at Marianjoy - No Injury | closed | kmettler | Customer Use | fixed | 2016-08-24 23:39 | 2016-08-24 | 262688 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue: I had a patient fall in the Ekso (our original Ekso) today ‚Äì patient ID was 11 ‚Äì I was able to control the fall and thankfully the patient was not injured. He fell forward and to the right ‚Äì he was on 0.2 step height and medium swing completed on fixed assistance. I‚Äôm not sure what happened honestly, but seems similar to the previous fall that we had here. This will need to be analyzed as the previous fall was to determine the cause. After the fall ‚Äì the Ekso was vibrating while the patient was on the ground and the tech that was assisting also said it was vibrating as he was falling‚Äì not sure if that is normal or not. | Clinical had a follow up call with customer | Clinical user error.
3267 | Crack in Hip Housing | closed | kmettler | Adjustments - Hip | fixed | 2016-08-24 23:26 | 2016-08-23 | 729621 | 154 | Confluence Health | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | Team- Confluence Health in Wenatchee has been in touch w/ us recently regarding concern over right leg hitting left leg during swing. In video, we have seen a drop of the left hip. I had asked them for more details today. Jeremy contacted me reporting he can see some movement of the left hip bolt in/out of Ekso on both sides when the torso is stabilized and the left leg is moved in/out for hip abduction/adduction. Video attached- though low res, not sure it can be seen. Jeremy reports there is slight movement of bolt in this video. Pic attached confirms bolt he is concerned about. They have patients scheduled now and for rest of afternoon. I advised them not to use Ekso until EB follows up, and to plan to alternate treatment at this time. -second video attached is video they sent last week of patient w/ complete paraplegia walking; can see left hip dropping during left stance with right leg mildly crossing midline to hit left, and left leg swinging out slightly during walking. Josh/Lily- Jeremy and Tricia called me directly. Jeremy can likely take the follow up. Please advise if you want me to follow up, or if you will. | Replaced both hip assemblies per SWO-02097. | Stress concentration at crack origin. parts given to engineering for further evaluation. Parts to be evaluated in engineering ticket #3003 ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/3003#comment:1
3265 | Cracked hip housing | closed | kmettler | Adjustments - Hip | fixed | 2016-08-24 13:01 | 2016-08-23 | 1416104 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | distributor contacted us because the device generates strange sound during walking - see video attached | Exchange both hip housings and close all open CO's. Work completed via SWO-2145 | Hip housing can crack with regular usage. Evaluation for improved design with ‚Äãhttps://berkeleybionics.trac.cloudforge.com/med_eng/ticket/3003#comment:1
3264 | Problem with DDL | closed | kmettler | Don-Doff Link | fixed | 2016-08-23 14:08 | 2016-08-23 | 151034 | 171 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Process Not Followed | distributor reported problem with ddl. He tried to clean it but it did not help. It looks like spring inside is broken. | Replace DDL SWO-2144 | Found that the adhesive which holds the drill pushing had come in contact with the spring, and thereby causing this issues. Processes already in place to prevent this issue.
3262 | Broken right hip housing | closed | kmettler | Actuators | fixed | 2016-08-19 14:59 | 2016-08-19 | 1220159 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Design | distributer reported a broken right hip while checking his device internally, pic is attached | replaced def. parts, swo 2063 | Stress concentration at crack origin. Hip housing cracks evaluated for design improvements with ENG Ticket: ‚Äãhttps://berkeleybionics.trac.cloudforge.com/med_eng/ticket/3003#comment:1
3261 | Problem with shifting weight to the left | closed | kmettler | Foot Sensor | fixed | 2016-08-19 14:49 | 2016-08-18 | 81971 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | +49 40 7306 - 3670 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | customer reported that even able persons have problems with shifting the weight to the left. Manual foot test did not show any errors. | Walked customer through foot tests over the phone. No SWO created as no piece parts needed to complete this work. | Calibration values on foot sensors can drift with usage. Recalibration of sensors can readjust these values.
3260 | unit #260 - BG Klinikum Hamburg - problem with charger | closed | kmettler | Battery | fixed | 2016-08-19 14:46 | 2016-08-18 | 81971 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | +49 40 7306 - 3670 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Determined Unknown | customer reported a 'defective contact' with one of the batteries update: customer updated his description with: defective contact is on the charger and not on the battery. The charger is showing the green light even if the battery is not fully charged yet. | replaced charger, swo 2089 | Will perform RCA on defective charger, not financially viable.
3259 | Problem with legs moving back | closed | kmettler | Foot Sensor | fixed | 2016-08-18 07:47 | 2016-08-17 | 170108 | 6 | Thevenaz - Switzerland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | customer sent video (see attached) and following description: Sometimes I have A problem With The legs . She move out on The back, and sometimes its dangerous | customer did foot test | Foot sensors require calibration (foot test) periodically.
3257 | Leg won't hold in place | closed | ckealoha | Don-Doff Link | fixed | 2016-08-17 00:12 | 2016-08-16 | 22986 | 274 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Customer Use | Issue: The leg will not hold in place, this has happened across the board with multiple patients. This is most likely an issue with the T-Block. | It looks like a piece of the plastic cover was getting jammed in there not allowing it to lock in place. | A piece of the plastic cover was jammed into the leg when it was locked.
3254 | Step Aborts | closed | lmcneil | Foot Sensor | fixed | 2016-08-15 20:48 | 2016-08-15 | 98702 | 107 | ICORD - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: They were having several step abort issues, they were walked through a auto foot cal and had multiple 171 errors after repeating the test several times. I had them test the sensors and they found that the L Heel is dead. | Sent new foot sensors via SWO-2066 | We will not be able to determine true RAC as we will not be receiving the parts back.
3245 | unit #261 - Step by Step/Tenerife - turning right leg | closed | uduell | Lower Leg | fixed | 2016-08-10 12:59 | 2016-08-09 | 25381 | 261 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Training | customer called Udo and reported turning right leg | Udo explained customer to tighten the screws a few more times | Customer was not using correct tightening method
3244 | Hip Width Stiff | closed | kmettler | Adjustments - Hip | fixed | 2016-08-09 22:17 | 2016-08-09 | 314312 | 173 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: The hip width on the R side is super stiff. | Replaced key(103193) and moved hip through range of adjustment to confirm smooth actuation. Work completed via SWO-2059 | Locking key(103195) showed signs of wear at point of engagement to the torso tube.
3243 | Left Hip getting stuck | closed | kmettler | Adjustments - Hip | fixed | 2016-08-09 21:17 | 2016-08-09 | 7071 | 281 | Franciscan Health Lafayette | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | Issue: The left hip is getting stuck in full extension during setup, it‚Äôs occurred 3 times and is being brought back to Ekso Hq for more analysis. | Replaced linear rail in Left Hip Assembly with parts via SWO-2092. | Linear rail was found to be defective.
3241 | unit #163 - Mark Pollock - device 'not quite in sync' with patient anymore | closed | kmettler | Foot Sensor | fixed | 2016-08-09 14:30 | 2016-08-08 | 296151 | 163 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | customer reported: Over the last few sessions I have noticed that my GT is feeling like it is not quite in sync with me. There is nothing specific that I could say is wrong but it is feeling strange. It feels like the device doesn't know quite what is going on - like it is slightly confused! I don't know how to fix this but I feel that it might need recalibrated or reset? update from 2016-08-11: Just to follow up on the GT issue. We ran the Auto Step Test yesterday and the test ran successfully. However, as soon as we began walking after this test, we received several #168 error messages. The Ekso was stalling at right foot toe-off, just before the swing phase on the right side. We tried walking in both Prostep and Prostep+ but the same error occurred with both. I have attached an image of the error message. | replaced foot sensors using parts from swo 2076 | foot sensors faulty
3240 | unit #125 - Technomex/Poland - left foot epoxy issue | closed | kmettler | Foot Assy | fixed | 2016-08-09 14:23 | 2016-08-08 | 1220159 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Wear | Device is currently used as loaner at customer site. Customer reported that the left foot has rotated inwards during patient swapt (see pic attached). | Udo replaced parts per swo 2063 | Found that the annodization layer on the parts had begun to wear down on the lower shanks, resulting in a smaller than normal tube. This allowed for a decrease in clamp strength
3239 | unit #125 - Technomex/Poland - broken left knee | closed | kmettler | Lower Leg | fixed | 2016-08-09 14:00 | 2016-08-08 | 1220159 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Materials | Device is currently used as loaner at customer site. Customer reported a broken left knee (see pic attached) | replaced def. parts, swo 2063 | Tool chatter marks violated surface finish requirements and caused stress concentration at crack origin. Current production parts do show signs of this defect.
3236 | unit #144 - BG Klinikum T√ºbingen - potentiometer error on left hip | closed | kmettler | Encoder / Potentiometer | fixed | 2016-08-02 13:24 | 2016-08-02 | 403055 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | customer called, because patient got the left foot caught on the crutches, which made the left knee got stuck. They tried emergency sit down, but this did not work anymore, so they switched the device off and set the patient down 'manually' on a chair. But even then the left knee stayed straight. After switching the device on again without patient they got asked to made a leg test, which they made with help of Udo and which finished sucessfully. But when they tried again with a patient they got a 'potentiometer error' for the left hip (unfortunately no error number was noted down) repeatedly. | replaced hip pot, swo 2078 | hip pot defect
3235 | unit #166 - Sheffield Hospital/UK - problem with foot rotation | closed | lmcneil | Lower Leg | fixed | 2016-08-01 15:55 | 2016-08-01 | 134753 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | customer reported again problem with foot rotation | replaced foot assemblies, swo 2054 | shank tube diameter to small
3234 | unit #166 - Sheffield Hospital/UK - problem with hip width | closed | lmcneil | Adjustments - Hip | fixed | 2016-08-01 15:51 | 2016-08-01 | 134753 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer has problem with adjusting the hip width. Already received the info to clean the tubes with alcohol from Michael, but this only solves the problem for one or two days and then the problem is the same again | added grease to adjustment screw per work performed by SWO-2054 | hip with is hard to set on the left side.
3233 | unit #166 - Sheffield Hospital/UK - foot binding not closing | closed | kmettler | Fasteners | fixed | 2016-08-01 15:47 | 2016-08-01 | 134753 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Customer informed us that the foot clip/fasteners kept coming loose while walking patient | replaced foot assemblies according swo 2054 | Cheap part that is easy to replace we will investigate further root cause.
3232 | Broken Shoulder Screw | closed | kmettler | Ankle | fixed | 2016-07-29 22:25 | 2016-07-29 | 234335 | 149 | SDSU | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: The customer called and reported broken should screw on Right Ankle | Sub-assembly replaced per SWO-2039 | Known issue with vendor supplying SS shoulder screws when alloy was requested.
3231 | Crack in Torso Shroud | closed | duanee | Torso | fixed | 2016-07-29 22:21 | 2016-07-29 | 226322 | 181 | VA Bronx | Steven Knezevic |  | 718-584-9000 x3130 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor | Investigation Needed | Customer Use | Issue: The shroud was cracked, found while servicing the device for another issues | Parts replaced per SWO-02026. | Unit dropped into equipment in therapy room. | It has been determined that the initial cause of the crack in the torso shroud is likely to have been damage caused by the unit being dropped. There are no apparent design or process issues that contributed to this. No further action is necessary. (Duane Edmonds ‚Äì 10 Sep 2017)
3229 | Error #121 | closed | kmettler | Motor | wontfix | 2016-07-28 15:18 | 2016-07-27 | 125976 | 251 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | customer reported: On Saturday we got an error #121 motor error and the left motor seemed to stop working. We have used it since and it has worked fine. The person on Saturday was an incomplete spinal injury whereas the 2 people who have used it since are complete spinal injuries. | no issue found | System limitation
3228 | unit #213 - EMAC-Agazzi/Italy - problem with right foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-07-27 12:02 | 2016-07-27 | 48164 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | distributor reported a broken foot sensor. Sole is peeling off completely. Pic is attached | replaced foot sensors, swo 2044 | Known wear item.
3227 | Broken joint cap on right hip | closed | kmettler | Actuators | fixed | 2016-07-27 11:29 | 2016-07-27 | 643914 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Major |  | Design | distributor informed about problem on the hip for the customer. Pics show a broken joint cap on the right hip. | distributor replaced caps, swo 2103 | part weak due to design, no mishandling by customer
3221 | unit #212 - EMAC/India - battery charger not working anymore | closed | jhanff | Battery - Charger | fixed | 2016-07-22 15:44 | 2016-07-22 |  | 212 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | distributor reported that one of the customer's battery chargers does not charge anymore | Parts mailed to distributor via SWO-2037 | Not cost effective to investigate root cause on battery charger failures
3220 | unit #197 - EMAC-1 Montecatone - problem with velcro on foot bindings and broken torso clip | closed | jhanff | Soft Goods | fixed | 2016-07-22 15:32 | 2016-07-07 | 424657 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | distributor noted during servicing the device at the customer's site that the velcro inside the foot bindings peels off and that the chest strap is broken | Parts mailed to distributor via SWO-2036 | These are normal wear components on the ekso
3219 | unit #197 - EMAC-1 Montecatone - problem with hip abduction lock | closed | jhanff | Adjustments - Hip | fixed | 2016-07-22 15:26 | 2016-07-07 | 424657 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | distributor reported a problem with the hip abduction lock, tried to fix it with help of Michael, but problem persisted. customer therefore needs new spring plungers. | Mailed replacement parts to distributor via SWO-2036. | Ball plunger set screw can get deformed causing it to not function properly.
3217 | Walker un-even | closed | najero | Cane / Walker / Crutch | fixed | 2016-07-19 17:40 | 2016-07-19 | 22537 | 246 | Baylor | Christina Fazio | cfazio@bir-rehab.com | (214) 820-9393 | United States | No injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor | Investigation Needed | Determined Unknown | Issue: Doing an advanced training at Baylor and their walker is really wonky and uneven. Please send them a new walker. | Sent new walker via SWO-2028 | We will not be receiving the walker back so cannot perform true RCA | The trained PTs are aware of the proper (and safe) operation of walkers and will not use a walker that presents a safety hazard to a patient. The walker was replaced. The original walker was not returned to HQ for analysis. It is not possible to further investigate this issue. The complaint database was reviewed for similar issues. No further action is necessary. (D. Edmonds ‚Äì 02 Oct 2017)
3216 | Hips width moving from set number | closed | kmettler | Adjustments - Hip | fixed | 2016-07-18 23:31 | 2016-07-18 | 223772 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: The hip here is moving from its set number. Video attached of the issue | Replaced 103195, per SWO-02026 restoring functionality as intended. | 103195 Hip Key Lock was worn, appeared as normal wear through the use of patient adjustments.
3215 | Cosmetic crack on the Knee cover | closed | jwitherspoon | Actuators | fixed | 2016-07-14 21:57 | 2016-07-14 | 240979 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Customer Use | Issue: Can you send Darrell this picture: There is a crack in the front of the EKSO that one of the tech‚Äôs noticed today. Kim doesn‚Äôt think it‚Äôs a big deal and shouldn‚Äôt affect the operation, but I thought he should see this before they come so if anything needs to be addressed they can bring that part as well. | Darrell hand carried parts via SWO-2013 | Exposed plastic cover that can break during regular usage.
3209 | Shoulder Strap Broke | closed | kmettler | Soft Goods - Shoulder Straps | fixed | 2016-07-12 21:29 | 2016-07-12 | 135111 | 64 | Dan Rose | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: One of the shoulder straps on my device broke. It is the strap on the top of the back pack strap that secures it to the top of the device near the remote. If you could send two of them because the second is looking a little frayed as well. | Sent new shoulder strap via SWO-2001 | see CAPA #2296
3206 | Error 21LH | closed | uduell | Encoder / Potentiometer | fixed | 2016-07-11 12:02 | 2016-07-07 | 7957 | 109 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Wear | Error 21LH on EU loaner unit | Executing Leg Test/Recal with SWO-2027 | Recalibration needed periodically
3205 | unit #213 - EMAC/Istituto Agazzi-Arezzo - problem with batteries | closed | jhanff | Battery | fixed | 2016-07-11 09:04 | 2016-07-08 | 48164 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | I received a confirmation from Agazzi about a problem with the other batteries couple. (2A-2B) They informed me last week that it lose charge very quickly jumping from higher percent to lower. I suggested them to make a completely discharge and a completely charge a couple of time and let me know. This morning they said that nothing is changed, so probably they need to replace also this couple of batteries. Why this is happened? Could be a problem with the charger? Update from customer: With the same patient, they said that the new couple of batteries discharge in normal way. With this they means that the discharge is gradually, like they are used to see. | Replacement batteries being sent directly to customer via SWO-2038 | Batteries lose capacity with regular usage
3202 | Request for new CI | closed | lmcneil | UI - Crutch Interface | fixed | 2016-07-05 23:14 | 2016-07-05 | 108889 | 14 | Spaulding | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | We are having an issue with our crutch interface. It is only working on occasion because the connection where the wire attaches to the computer/backpack is no longer locking. So we can only use ProStep when the connection is good (took multiple trials to get today). | Sent parts via SWO-1985 | Known wear item
3200 | Error 21 & 27 RH | closed | kmettler | Adjustments - Hip | fixed | 2016-07-05 22:29 | 2016-07-05 | 242200 | 175 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Process Not Followed | The customer called and reported Error #21 & #27 RH occurring continuously. They were not able to use the device. | Encoder was replaced due to damage during removal. Work done via SWO-2093. | Encoder hub was not set on motor shaft properly.
3199 | Cracked Hip Housing | closed | kmettler | Actuators | fixed | 2016-07-05 14:51 | 2016-07-05 | 26373 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Ekso making a noise on the left side difficult to determine exactly where from. Not effecting walking but very distracting. additional info from customer after receiving the loaner: -------------------- -it sounds significantly quieter than our Ekso -it feels smoother to handle both in walking and in the preparation stage prior to walk e.g.. when testing the leg ROM the mechanism feels smoother -it generally feels a lot different to our Ekso Sorry can‚Äôt be more specific but I could really tell the difference between the two devices. -------------------- | replace hip housings, SWO 2064 | hip housings cracked
3198 | Broken hip housing | closed | kmettler | Actuators | fixed | 2016-07-04 16:58 | 2016-07-04 | 573770 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Determined Unknown | After the customer reported the broken joint cap, Michael asked for pictures from the hip housing. They show that the hip housing is broken, too. | Parts sent per SWO-01993. Hip housings were replaced with new parts. | Will not evaluated these parts as this is a repeat occurrence of a now known issue.
3197 | Screw on right heel cap broken | closed | kmettler | Fasteners | fixed | 2016-07-04 16:50 | 2016-07-04 | 573770 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | the screw of the right foot that I was talking with Udo was broken | Old screw removed and replaced per SWO-1928. | Material failure.
3196 | Joint cap broken | closed | kmettler | Actuators | fixed | 2016-07-04 16:42 | 2016-07-04 | 573770 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Design | this morning we hear a noise on the left hip and it's also broken. It seems to be the same problem that happened last time, but on the other leg. | Joint cap replaced per SWO-1993. Down rev version of joint cap can break during regular usage. New version implemented via CO-11476. | Design.
3195 | Black rubber on sole peels away | closed | jhanff | Foot Sole | fixed | 2016-07-04 16:35 | 2016-07-01 |  | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | The customer Anagenisi rehab Center reported that they have problem with the detachment of a material on the outsole. We know that this is a common problem and that you try to find alternative solutions. Please, let us know what to advise the customer and what to do. Michael advised to glue the sole material. Reply from customer: we discussed with our customer the option of glueing the sole material as a temporary solution. However, they aren‚Äôt satisfied with this option and they require a more permanent solution as they are concerned on patient safety. Could you please propose an alternative? Is there an option of replacing the outsoles? | Sent parts via SWO-01983 | Tread can peel during regular usage.
3194 | Error 21-RH | closed | kmettler | Electrical Interconnect | fixed | 2016-07-04 16:25 | 2016-07-04 | 534019 | 54 | Prosperius - Italy | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | We have a problem with our Ekso. When I do the hip right flexion the Ekso writes that there is an error. The number is #21-RH. I try to reboot but there is always the same error. | Customer could solve issue doing Leg Calibration. No SWO create as device was fixed over the phone. | Certain use will require the system to do a leg calibration in order to have the device function properly.
3193 | Error #133 | closed | kmettler | Motor | worksforme | 2016-07-04 16:21 | 2016-07-04 | 117387 | 50 | GEMI - Poland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | customer contacted us with: During a walk a device is switching to an emergency sitting mode. Got as additional info: error 133 2016/07/08: happened twice by now, customer could not give info additional info about error message (like L or R) because he always only checked for the error number | - Michael sent parts to Technomex via SWO-01988 -Technomex was not able to reproduce the problem. They made around 500 steps in device trying to provoke error but without success. -SWO has been canceled. | 121/133 errors can occur through normal operation. Distributor was unable to replicate error.
3190 | L Hip Abduction is slipping | closed | kmettler | Adjustments - Hip | fixed | 2016-06-27 21:39 | 2016-06-27 | 281493 | 164 | VA Minneapolis | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Workmanship | Issue: Our EKSO L hip abd is slipping. I think this may have been happening during training as well, as I had noticed an asymmetry a couple of times when resetting the device for our pts. I confirmed with others to make sure that they check to make sure things are symmetrical. Over our training time, I have noticed continued asymmetry when resetting for other patients. I have mentioned it to others and they had started to feel like something else was going on. Today, I made sure both sides were symmetrical at .5, check was symmetrical, after first 100‚Äô still symmetrical. During the last have of the walk, we noted that pt was starting to catch L L/E on R. After he sat down, the abd marker was at -2. What should we do? - The PT also added that it got gradually worse over time and was happening across the board. | Plunger (102266) was reseated into the hip abduction housing per SWO-02180 using step 4.8 in MPI 102170. | Spring plunger (102266) was not seated properly in the left hip abduction housing.
3189 | Tread peeling on R Foot | closed | kmettler | Foot Sole | fixed | 2016-06-27 21:16 | 2016-06-27 | 231362 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: Josh, some of the sandy grip material on the bottom of the right foot of the Ekso has peeled away. Don't want to play around with gluing it. Should I send it back or do you want to send us two new rubber foot plates. It looks they Velcro on so we could just swap them out and pitch the old ones. Thoughts??? | New foot sensors installed on device via SWO-1768 | Normal wear.
3187 | Hip motor getting hot | closed | kmettler | Motor | fixed | 2016-06-24 10:17 | 2016-06-22 | 157004 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | We noticed that the motor units on hip level (both sides) get really hot, even after a few minutes walking. A ventilation system is starting up and also the backpack side gets really hot. This occurs with different patients and it results in an important decreased autonomy of the batteries. For example we have to stop a session after 25 minutes of walking due to low battery level where we normal can continue the session up to 50 minutes (with the same patient). There is no error occurring on the interface. Details : Ekso GT SN 021508003 Total step count : 146160 Starting of ventilation system on motor units level : walk time : 9‚Äô34‚Äù, up time : 10‚Äô01‚Äù, steps : 406 Patient settings SCI AIS A Th11 44 Years old Height = 184 cm; weight = 80 kg Ekso experience since January 2016, frequency walking sessions : 2 times/week | ELMO calibration, leg test per swo-2268. | Unknown why elmo values change over time.
3186 | Left foot turning, epoxy issue | closed | kmettler | Lower Leg | fixed | 2016-06-24 09:56 | 2016-06-21 | 12334 | 234 | Centro Sinapse - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Process Not Clear/Complete | yesterday we saw that the left foot of the Ekso turn to the right (internal rotation). I attach two photos where you can see the final part of the leg. That part it's free and the control during the steps is very complicate because the foot does internal rotation | Both lower Feet (Complete) and the two upper shanks where replaced for the new versions. SWO-01906 | Glued shanks on device were built to down rev version at time of original assembly.
3185 | Issue with the left battery connector | closed | kmettler | Battery - Connector | fixed | 2016-06-23 15:01 | 2016-06-22 | 291353 | 13 | Mark Beack - South Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | Ekso 1.1 has an issue with the left battery connector. The connection was previously extremely tight and now some pins have bent. | Customer bent battery pins back into place | Customer incorrectly installed battery
3184 | Error 1180 | closed | kmettler | UI - Controller | fixed | 2016-06-22 09:41 | 2016-06-22 | 403063 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Customer Use | customer got LCD-error 1180 three times in a row, all 3 patients must sit down | #NAME? | Found the outer casing of the UI to be damaged which would infer repetitive drops. Also found blunt force damage on the inside of the UI which would infer that the component was dropped multiple times.
3183 | Right leg is delayed while walking | closed | kmettler | Foot Sensor | fixed | 2016-06-22 07:23 | 2016-06-22 | 170108 | 6 | Thevenaz - Switzerland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | I have a problem with the first time when I Walk. The Right legs Wait when I walk. | Walked customer through foot tests over the phone. No SWO created as no piece parts needed to complete this work. | Calibration values on foot sensors can drift with usage. Recalibration of sensors can readjust these values.
3181 | Foot Sensor Issues | closed | kmettler | Foot Sensor | fixed | 2016-06-21 15:28 | 2016-06-21 | 197783 | 55 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | The Ekso 1.1 is not working properly for all the session. Sometimes we are able to start and finish the treatment without problems, Sometimes instead we need to stop and to start more than one time. Problems are the message "reboot" and also that is difficult for the patient do a proper shifting in lateral and in forward position (when it is in Prostep mode). | customer did foot tests for recalibration | periodic foot recalibration needed as part of the design.
3180 | Broken joint cap - Right | closed | kmettler | Actuators | fixed | 2016-06-21 15:25 | 2016-06-18 |  | 212 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | The big Screw of the Hip Joint of the Left leg is practically outside of the joint. This problem occurred on Saturday evening during the last treatment. We immediately stopped using Ekso and it was put into sitting position. | Parts to be delivered for fix via SWO-1963 | Down rev version of joint cap known to fail during use. Part not designed sufficiently for uneven loading.
3179 | Broken joint cap right | closed | kmettler | Actuators | fixed | 2016-06-20 13:04 | 2016-06-20 | 85371 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Major |  | Design | distributor reported broken joint cap (see pic) patient is fine | Customer replaced parts according swo 1980 | Old version of joint cap can break with regular usage.
3173 | unit #213 - EMAC/Istituto Agazzi-Arezzo - problem with foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-06-16 14:01 | 2016-06-15 |  | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor |  | Design | Distributor reported: Because I follow them by phone on the auto calibration, but the problem persist. So they need to replace it. Additionally already did by phone, following PT, the foot auto test. It not passed. When I will be there I'll check also the sensor reading. | distributor replaced parts sent via SWO-1908 | Known wear item
3172 | unit #199 - EMAC/Italy - problem with left foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-06-15 12:52 | 2016-06-15 | 120595 | 199 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | the left foot sensor doesn't read well and it not launch the right leg. | distributor replaced parts from swo 2050 | Known wear item.
3171 | Adjustment rod on left ankle broken | closed | kmettler | Adjustments - Ankle | fixed | 2016-06-15 12:45 | 2016-06-15 | 120595 | 199 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Customer Use | Distributor reported problem with left ankle with broken adjustment rod. The video shows the problem. | distributor replaced parts from SWO 2050 | Parts returned to engineering for RCA. Further analysis to be performed in Engineering Ticket #3002 ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/3002#comment:1
3170 | Stiff Leg | closed | lmcneil | Adjustments - Leg | fixed | 2016-06-14 23:33 | 2016-05-19 | 13259 | 262 | Magee | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: We have been working with the EKSO GT at Magee Rehab Hospital. The Right upper leg seems to be getting stuck when trying to make adjustments to shorten or lengthen. It seems to be easiest when unlocking the thigh rotation component. Is this normal? Are there any other tips we can be using to make these adjustments easier? | Resolved per SWO-01946. The thigh clamp key (102089) was replaced and the thigh bore and lower donn doff link were cleaned with IPA. | The thigh clamp key (102089) was binding up on the donn doff link (102242).
3169 | Broken joint cap | closed | kmettler | Actuators | fixed | 2016-06-14 11:49 | 2016-06-14 | 710598 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Materials | broken joint cap on left hip | Exchange outer Joint Caps for new ones.Replace inner Joint caps for new ones. Exchange screws for new (black) ones. Grease bearings in all 4 joints.Replace the missing ball/roll of the RH bearing with a new one. SWO 2117 | material fatigue on inner Joint cap of RH.
3166 | right foot binding doesn't fasten properly | closed | kmettler | Foot Assy | fixed | 2016-06-13 15:54 | 2016-06-13 | 46991 | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Determined Unknown | The clip that locks the right ankle doesn't fasten properly. The therapists referred to the clip that secures the right ankle, possibly meaning the adjustments. | Binding needs to be replaced according SWO-01962 | Will not inspect bindings for further defects as this is not cost effective.
3165 | Hip abduction setting changes during use | closed | kmettler | Adjustments - Hip | fixed | 2016-06-13 15:51 | 2016-06-13 | 46991 | 245 | Ostracon Ltd - Greece | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | The hip abduction setting is moved/changed during use (especially at the left hip). The hip settings change with every patient and according to the therapists, Eva also noticed it during their advanced training a couple of weeks ago and suggested to them to report it if it would continue. The extent of the change is of at least 1 unit. | cleaned adjusting screw from oil - NO SWO SINCE DISTRIBUTOR FIXED | adjusting screw oily
3164 | Screws on left foot are broken | closed | lmcneil | Foot Sole | fixed | 2016-06-13 14:59 | 2016-06-13 | 513638 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | The screws of one foot are broken. Please find attached the last picture. It happened when we were going to turn left while walking. | Exchange both lower feet for new ones, including the new V3 solid Shank. Work completed with worked performed via SWO-1928. | See CAPA #2856 for complete RCA
3163 | unit #234 - Centro Sinapse/Spain - error with batteries | closed | kmettler | Battery | fixed | 2016-06-10 16:40 | 2016-06-07 | 12334 | 234 | Centro Sinapse - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | customer reported 'error with right battery' and asked for help to find out if the problem is with the battery, the charger or the ekso. 2016/06/22 - update: we had a session with the Ekso and all worked well, but at the end it came the same error ("Falta Bater√≠a derecha. Sustituya y reinicia #4"). But now, when we were preparing the next session, we wanted to put the batteries, and now it's impossible to put the left batterie. The up part of the batterie must conect with the Ekso, but doesn't connect now. | Batteries replaced with SWO-1996. | Battery 1B is not recognized by system. Will not be able to return battery for further root cause analysis.
3162 | Sensor error | closed | kmettler | Foot Sensor | fixed | 2016-06-10 16:33 | 2016-06-09 | 111.986 | 139 | Asepeyo - Madrid | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported a 'sensor error' | Problem solved with the last Service on SWO-01901 | Regular wear seen on foot sensors will require periodic calibration in order to allow to function properly.
3157 | Error 1241 RK | closed | kmettler | Electrical Interconnect | fixed | 2016-06-10 16:10 | 2016-06-09 | 557088 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Materials | The unit #233 had a software error #1241 RK a few minutes ago. After rebooting it still occurs. | replaced motor, decoder cable and cca. SWO 1932 | short circuit in decoder cable.
3155 | Repeated #133 and #121 Errors | closed | kmettler | Electrical Interconnect | fixed | 2016-06-08 23:19 | 2016-06-08 | 233873 | 59 | Able Bionics - Aspen | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Materials | Issue: The PT's reported that they were having issues with each patient in the device on 6/13 and 6/6. On 6/8 the errors happened so often they were not able to continue the use of the device. | SWO-01955. All four motors (102292 & 102294) were replaced. | Left hip motor had a large amount of resistance and was replaced. Both knee motors were also determined to cause motor errors and were replaced.
3147 | Fall at Marianjoy | closed | kmettler | System | fixed | 2016-06-03 16:56 | 2016-06-01 | 227367 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Customer Use | Issue: Customer: Marianjoy Unit #202 Issue: They had a patient walking in the device this morning, incomplete SCI T4-T6. They were walking in adapt and then put her into a fixed setting, they believe it was 50 on the left and 70 on the right. During the session the pt just finished a left step and then the left knee began to buckle, this was followed by the R knee buckling causing the pt to loose balance. The PT and aide walking with the pt weren‚Äôt able to secure them and they fell forward almost directly onto their knees, right before the knees hit the ground the Ekso straightened causing her face to hit the ground first. They said the Ekso then went into safe mode, no errors were heard prior to this event. They rolled the pt over onto their back and called for emergency help. She‚Äôs been taken to their hospital and is being examined right now. They did say the pt split their lip and chin and was put into a cervical collar. They don‚Äôt believe they will need stiches but this area was bloody. Patient details: 5‚Äô2, 120 #‚Äôs. 5th time using Ekso. Doug, we are waiting on finding out the status of the patient, is there any further information we need to ask the customer about this fall? Katie once we get the data please let me know and hopefully we can find out what happened in more detail. We are bringing this device back to Hq and it should be here early next week if not sooner. | Refer to SWO-01938. Unit was put through a thorough inspection to ensure that the unit was functioning properly. No errors or damage to any components, other than the right foot plate, were found. | Customer use, refer to narrative from Darrell on 6/1.
3146 | DUPLICATE: Problem with internal rotation of lower leg and foot | closed | kmettler | Lower Leg | duplicate | 2016-06-03 10:51 | 2016-04-08 | 415505 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | The problem of the internal rotation of the lower leg and foot part in opposite of the upper leg still occurs. Sometimes this is solved by using other bolts (but we don't have any new one). It is always the left side which is rotating. It occurs with different patients, but I do think it occurs more frequently with heavier patients. | Ticket duplicate to #3057 | Ticket duplicate to #3057
3145 | DDL opens sometimes | closed | kmettler | Don-Doff Link | fixed | 2016-06-03 10:00 | 2016-01-21 | 4489 | 162 | Hobbs - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | right DDL opens while placing the patient into the ekso | Pressed drill bushing out, cleaned it and pressed it in again according MPI. Previous fix superceded by the installation of the latest version of the DDl assembly. Latest work completed via SWO-2060 | Fix implemented, not able to determine root cause. drill bushing not fix and moved outwards (1mm)
3142 | Worn R foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-06-03 06:08 | 2016-06-01 | 232570 | 59 | Able Bionics - Aspen | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | I am at the Snowmass Club today with all three PT's. We determined that it is the RIGHT foot sole that is torn with the sensors compromised. It is currently duct taped together and functional. This is the same sole that was replaced a little over three weeks ago. Could you please overnight a new right foot sole with sensors to replace again? We still have the left foot replacement, which is new and unused. | Sent new sensors via SWO-1926 | Most likely - This is a known wear item, will not be receiving the sensors back so cannot perform true route cause analysis
3140 | Error 137 | closed | kmettler | Brake / Release / 9V Battery | fixed | 2016-06-02 13:44 | 2016-06-02 | 19777 | 259 | Zentralklinik Bad Berka - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported error 137 (brake slipping) | cleaned the brake with alcohol, burnished brake, no parts needed. Work conducted on servie done via SWO-1992 | Break can clip due to accumulation of dust. Brake cleaned and re-burnished.
3139 | unit #259 - Bad Berka - Left foot rotating | closed | kmettler | Lower Leg | fixed | 2016-06-02 13:40 | 2016-06-02 | 19777 | 259 | Zentralklinik Bad Berka - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | the customer just called me an claimed a rotating foot. On my last visit I checked the clamps and the screws didn‚Äôt touch the inside of the collar so no need to add more washers. The customer said the tube is turning so it is not the epoxy issue. | replaced foot assemblies to SWO 1992 | Dimensional tolerancing and stack up between mating parts can allow for rotation given force customer exerts on device.
3138 | problem with batteries | closed | kmettler | Battery | fixed | 2016-06-02 09:13 | 2016-05-16 | 129361 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | customer: EMAC/Istituto Agazzi-Arezzo Batteries 1A/1B doesn't maintain the charge. In a session of 9 minutes with one free leg they passed from 60% to 25%. | The Customer will receive a new pair of batteries and will send us the defective one (1A/B) to our HQ in Freiburg, Germany.Batteries sent to customer via SWO-1909. Incorrect TRAC listed in closed SWO. Update added to comments section. | I walked the Unit by myself and could find a difference between pair 1A/B and 2A/B in behalf of the power loss. Meanwhile pair 2AB show a normal power drop after a certain amount of steps, pair 1AB dropped from 96% to 81% after only 5min walking time. We are unable to ship back bad batteries for further evaluation due to shipping restrictions.
3136 | both ddl's cracked | closed | kmettler | Actuators | fixed | 2016-06-01 14:16 | 2016-05-17 | 557088 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Design | both ddl‚Äôs show beginning cracks. | replaced ddl's, swo 1854 | Design requirements did not properly anticipate customer use in terms of aggressive split stances during turning. RCA evaluated with CAPA #2489
3132 | Broken Copper Rivet | closed | kmettler | Soft Goods | fixed | 2016-05-31 20:37 | 2016-05-19 | 82349 | 107 | ICORD - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | Issue: One of the copper rivets came off from our Ekso at ICORD (Vancouver). Can we reuse these rivets? Do you have instructions on how to properly put them back on? | Sent spare parts via SWO-1919 | Will not receive the part back, so root cause will be unknown
3131 | Error #1180 | closed | kmettler | UI - Controller | fixed | 2016-05-31 18:48 | 2016-05-26 | 588795 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: We were getting a reboot #1180 error message 3-4 times today when moving the EKSO around before the patient was in it; when they were walking it didn‚Äôt have any issues. I haven‚Äôt put the low friction sensors on yet, so if this error message is because of the sensors, let me know. | Sent new parts via SWO-1922 | Could not determine root cause as the UI will not be returned to HQ.
3130 | Joint cap broken | closed | kmettler | Actuators | fixed | 2016-05-31 15:37 | 2016-05-31 | 487576 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Determined Unknown | We have a problem with our ekso unit. Just this morning we have detected that the right leg was doing a weird movement in ABD when doing steps. By the end, we detected that one screw from the hip seems to be broken. | Exchange Joint Caps and screw for new ones.Check all housings for possible cracks, but everything is ok. The broken items will be send to our HQ in Richmond for further investigation. Addressed with work performed via SWO-1993 | Root Cause will be evaluated on Engineering Ticket #2997: https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2997
3129 | hesitation in triggering the right swing phase | closed | kmettler | Foot Sensor | fixed | 2016-05-30 09:48 | 2016-03-15 | 352380 | 232 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | here‚Äôs a hesitation in triggering the right swing phase when supporting in stance phase on the left side. The patient has reached sufficiently the targets which are desired but then there‚Äôs a hesitation in the initation of the right swing phase. This problem doesn‚Äôt occur always but it changes between patients and also within one session. | foot sensor calibration fixed the issue. No SWO created to account for the completion of this work. | Foot sensor values can drift with regular usage
3126 | unit #045 - Justin Smith/SouthAfrica - batteries | closed | jhanff | Battery | fixed | 2016-05-25 08:14 | 2016-05-17 | 832686 | 45 | Justin Smith - So. Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | In addition our one set of batteries packed up this morning and have stopped charging so we will be using the other set of batteries a lot, so we will need another 2 sets of batteries. | Sent new batteries via SWO-1884 | Batteries are known wear items. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
3120 | unit 109 motor makes loud noise | closed | kmettler | Actuators | fixed | 2016-05-24 15:28 | 2016-05-24 | 386165 | 109 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Determined Unknown | motor makes loud noise, Technomex already replaced the motor, I assume the AMP board is defect | motor makes noise while moving, Technomex replaced motor but issues still exists. Technomex replaced amp board. No issues reported after work completed. | Replaced possible root causes.
3117 | DUPLICATE-Binding Broken | closed | kmettler | Foot Assy | duplicate | 2016-05-23 17:22 | 2016-05-23 | 157788 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Undetermined | Issue: | DUPLICATE TO #3121 Sent new parts via SWO-1899 | DUPLICATE TO #3121
3111 | problem with the Hip ABD | closed | kmettler | Other | fixed | 2016-05-19 13:17 | 2016-05-16 | 129361 | 213 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | They have a problem with the Hip ABD. It change values during the training. For example, at the beginning is setted to 1 and at the end of the training it changed to 0. | Clean the screws from the Hip Abduction adjustment from the protective Oil. Verified this issue has been addressed when device last serviced via SWO-2106. No additional parts needed. | Excess grease found in interface between adjustment plate screw and control bar threads in hip abduction assembly. Unsure of cause for excess grease.
3110 | Unit 185 left hip housing broken | closed | kmettler | Adjustments - Hip | fixed | 2016-05-19 12:47 | 2016-05-19 | 588370 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | customer reported left hip is broken | replaced hip housings with SWO-1890 | no mishandling or mis-use visible. Cracks in hip housing can form through regular usage.
3109 | Skin breakdown at the trochanter | closed | kmettler | Shipping / Handling / Packaging | fixed | 2016-05-18 17:27 | 2016-05-18 | 89898 | 135 | GF Strong - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Major |  | Materials | Issue: I am doing an advanced training at GF strong in Vancouver. One of the PTs here also works at ICORD in Vancouver, as well as a researcher from there who is observing the training at GF strong. They told me they recently had some skin breakdown at the trochanter of a patient at ICORD. They reported it to be on one side, more of a friction rub than a pressure ulcer, and has since healed. He has not been back in the Ekso since, mostly due to his schedule. He has used Ekso for several weeks without issues before this happened. I advised them to re- check his hip width setting as the patients weight may have fluctuated. My question is, is there anything I need to do since they have told me this? It is a bit unusual as the information is about a clinic which I am not visiting, and can't not inspect their Ekso. I am in Vancouver until tomorrow evening. | This was a minor skin irritation issue. | The thigh pad can cause irritation
3106 | Both hip housings cracked | closed | kmettler | Actuators | fixed | 2016-05-17 17:58 | 2016-05-17 | 557088 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Design | The black bolt on level of the hip joint (right side) is popping out. It occured during walking a patient. We have immediately stopped the session. Patient‚Äôs profile (walking session with problem occured) SCI AIS A Th 6/Th 7 Experienced with Ekso walking since June 2015 (2 times a week) Height : 171 cm Weight : 75 kg Poor spasticity (MAS = 1 or 2) No hip extension or knee extension restrictions. ---- Michael found the right outside joint cap deformed but not broken. He checked the housings too and found a big crack on the right hip and a beginning crack on the left hip. In addition both ddl‚Äôs show beginning cracks. | replaced hip housings, swo 1887 | Stress concentration at crack origin. Parts have been redesigned
3102 | Strap of binding not holding | closed | najero | Foot Assy | fixed | 2016-05-17 17:10 | 2016-05-17 | 54503 | 195 | Miami Project | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: The strap of their foot binding comes out if you touch it lightly | Sent new parts via SWO-1880 | Cannot determine root cause since the part was not returned to HQ for analysis. SIMILAR TICKETS: Inspected the locking mechanism with Al and Edmundo. There doesn't appear to be a problem with the locking mechanism when evaluating the straps that locks the foot binding in place. Although there appears to be a trick to it by pressing the buckle to the strap and then locking it.
3100 | Step Aborts | closed | kmettler | Foot Sensor | fixed | 2016-05-17 00:08 | 2016-05-13 | 137272 | 222 | Kennedy Krieger Institute | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | Issue: 222 has had some step aborts so I checked it out. The right foot is pretty noisy. A recal may help as the noise is pretty close to 500, but we should look at getting that replaced at some point. | Walked the customer through a re-cal over the phone, it passed successfully. No SWO required to capture the completion of this work. | Gain values needed to have foot sensors work properly can change as foot sensors are used over time. If still functional this can be addressed and fixed by performing a foot recal
3097 | Foot Sensor Issue | closed | kmettler | Foot Sensor | fixed | 2016-05-13 21:22 | 2016-05-13 | 473864 | 28 | Mt. Sinai | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Determined Unknown | Issue: We walked them through a foot cal, it failed and we had to send them new sensors to replace their damaged ones. | Sent new sensors via SWO-1877 | We will not know the true route cause of this issue as we did not receive the sensors back.
3095 | Foot Sensor Issue | closed | kmettler | Foot Sensor | fixed | 2016-05-13 18:55 | 2016-04-26 | 52040 | 110 | UHS | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue: My name is Gabby Canales, a Physical Therapist at University Health System. Today, I was working with a patient in the Ekso. I had an issue that came up today and wanted to know your thoughts on it. We had been walking for about 12 minutes when we decided to sit for a rest break. Afterward, when we began to attempt walking again, the machine gave us some problems, however no issues were detected with the device. The pt was in ProStep. On the R First Step, it was as though the device wasn‚Äôt reading the targets that were entered. I had to take the pt outside of the targets to get a reading. We went to stand still after which I checked the pt‚Äôs settings which were correct. We attempted again, however, the second time, the pt‚Äôs left leg went into extension during the R First Step. We were unable to get the targets read for the pt‚Äôs weight shifting and couldn‚Äôt continue with walking. We ended treatment after this. I‚Äôm not sure why this happened and thought I could try to get feedback from you. Also, there is a piece on the left foot that comes out from time to time. I‚Äôm not sure if it is a sensor, but it is red-orange color. This patient‚Äôs settings were in Bilateral Max Assist. | New sensors sent via SWO-1861 | Known wear issues due to the design. This is a low risk problem. The defect will be corrected in the next version of the device. Additional analysis to be done with Engineering Ticket 2998 then updated here. ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2998
3090 | Problem with the rigidity ankle selector | closed | kmettler | Ankle | fixed | 2016-05-13 13:28 | 2016-05-13 | 1081743 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | the Villa Beretta's Ekso has a problem with the rigidity ankle selector. They set it at 4 and during walking it move to 2. More over if they move manually the ankle, they hearing a noise of a part that move inside. | Exchange complete foot ankle assembly. Parts provided by SWO-1820. | Additional analysis to be done with engineering ticket #2996: https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2996
3087 | Error 1065 R Knee | closed | kmettler | Electrical Interconnect | fixed | 2016-05-11 09:08 | 2016-05-11 | 126320 | 113 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Issue: when switching on the device this morning and setting the legs to ‚Äúlegs free‚Äù mode, they immediately got "error 1057 RK ‚Äì Software error‚Äù. This is reoccurring when they re-start the device. 26379, 36517, 121, 3, 64, RK, Elmo motor failure bit set 26378, 36491, 1057, 3, 66, RK, CAN PDO received with error 26377, 36491, 1065, 1, 66, 1073762125, Elmo MO command rcvd with error | Replacement of knee cable and other components being sent via SWO-1879. Identified replacement of external cable as fix needed. See attached photos. | Removed parts to be returned to HQ for further evaluation Dan - 8-22-2016 - No parts received at HQ Dan - 8-22-2016 - Based on the UI errors that show in the description, engineering thinks that the source is more consistent with an ELMO problem that we have been tracking since early this calendar year and ELMO themselves have admitted that they had an issue causing the 1057 errors. The following is a link to the TRAC ticket capturing the 1057 ELMO errors: https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2994
3086 | UI stopped working | closed | kmettler | UI - Controller | fixed | 2016-05-10 21:56 | 2016-05-10 | 537191 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue: their UI has stopped working multiple times today. Attached is a picture of what they are seeing. | Sent new UI Via SWO-1863 | Similar behavior captured in tickets in engineering database. Behavior likely related to off the shelf Crystalfontz LCD module. But it seems likely that the module is being damaged prematurely in the field. https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2765 https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2925
3084 | Possible cabling/knee joint board errors | closed | kmettler | Actuators | fixed | 2016-05-10 21:23 | 2016-05-09 | 32255 | 263 | Rehab Institute of Kansas City | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | RIKC called with the following errors on Tuesday 5/10 They had the following errors at around 3-4 (their time): 1060, 1061, 1241, 1054. They also reported an errors #160 on 5/9 - All related to the User Interface ---- | Unit was test walked very roughly in order to throw the error again, but there were no errors. Work was performed under SWO-1866. | The knee breakout CCA, joint board, and external cable were all replaced one at a time in order to determine the root cause of the error. After replacing the amp board and having the same error, the original amp board was reinstalled. After that, the error went away and I was unable to replicate it during the test walk.
3079 | Request for battery chargers | closed | kmettler | Battery - Charger | fixed | 2016-05-02 22:15 | 2016-04-28 | 81317 | 59 | Able Bionics - Aspen | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue: They called and said that their battery chargers were no longer holding the charge on the batteries. | Sent new charger via SWO-1849 | Known design issues. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
3077 | Possible wirring issues | closed | kmettler | Electrical Interconnect | fixed | 2016-04-29 18:24 | 2016-04-29 | 114859 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Workmanship | Issue: The PT had a patient session morning of 4/29. They went to turn on the device and a beeping noise came up and the UI screen was blank. This still persisted even when unplugging the UI and plugging it back in, turning off the device and turning it back on | UI cable was seated fully onto the amp board per SWO-01855. | UI connector was not fully seated onto the amp board and had walked out after training.
3073 | Possible brake release issue | closed | kmettler | Brake / Release / 9V Battery | fixed | 2016-04-28 23:50 | 2016-04-28 | 529161 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue: I work at the Burke Medical Research Institute in White Plains NY with EKSO. A couple of weeks ago I replaced the batteries in the knees (flexion/extension when the robot is turned off). This worked totally fine, but since last week the right leg battery part seems broken. The leg doesn't move and I don't hear the 'click' sound if I press the button. I tried to replace the battery again, but that didn't fix the problem. Hope you can help me with this. The break release did not work even when opening the casing and moving around the wire. | Battery cable (102550) was replaced per per SWO-1856. | The crimp on the ground wire on the 9V battery cable (102550) was pulled out of the connector.
3072 | Batteries Draining quickly | closed | kmettler | Battery | fixed | 2016-04-28 21:05 | 2016-04-07 | 103747 | 14 | Spaulding | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Issue: The batteries seem to be draining very quickly actually - we started at 98% and 47% during a 2 hour period of mostly just having the device powered on. We took 66 steps total in that time. | Sent new batteries via SWO-1850 and via SWO-1872 | The batteries were very old so this is most likely a wear item however we won't be able to complete true root cause as we won't be getting the parts back.
3071 | Batteries Draining quickly | closed | kmettler | Battery | fixed | 2016-04-28 21:05 | 2016-04-07 | 103747 | 14 | Spaulding | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Issue: The batteries seem to be draining very quickly actually - we started at 98% and 47% during a 2 hour period of mostly just having the device powered on. We took 66 steps total in that time. | Sent new batteries via SWO-1850 | The batteries were very old so this is most likely a wear item however we won't be able to complete true root cause as we won't be getting the parts back.
3070 | Rotating lower legs | closed | kmettler | Lower Leg | fixed | 2016-04-27 14:45 | 2016-04-27 | 8902 | 259 | Zentralklinik Bad Berka - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor | No investigation needed - see Justification | Customer Training | customer contacted us because of rotating lower legs. Already provided info about washer which did not help at all - customer had added 6 washer in the end without any success. | Cleaned tubes, explained correct usage of t-wrenches. No SWO required to apply fixes noted. | Showing the customer the proper techniques contributed to proper functionality of the device. | No new investigation needed as previously investigated and deemed not a safety risk or a reportable event. See 102437 - Risk Management Spreadsheet (FMEA Tab) # 41 & 42 for original investigation information and mitigations already in place. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager on 3/16/2017)
3068 | Issue with CCA in knees | closed | lmcneil | Other | fixed | 2016-04-26 11:49 | 2016-04-26 | 345720 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Workmanship | rattling sound while unit is switched on. Replace CCA in knees | replaced cca 102286 in left knee. Parts from SWO-1845 Dan - 8-2-2016 - Tried out thigh board in known good system and was able to replicate brake chattering at least once. On subsequent freeing of the legs brake did not chatter though it did once after 5 minutes in the free'd state. RCA not precisely known though I suspect some design marginality in the boost converter that energizes the brake coil in which the supply might be hitting a current limit temporarily. Opened an engineering ticket for further investigation to this issue, ticket 2995 in EE/ME engineering database. Link here: https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2995 | Dan - 11-2-2017 - Root cause for brake chattering has since been determined to be the resettable fuse on the 12V supply on the knee connector opening and closing due to current limit which in turn results in the on and off behavior of the brake power supply which results in brake chatter. This in part may be due to a design issue which has since been addressed in a 102284 Rev N with a fuse with increased current rating but more likely is due to soldering workmanship of the fuse part and aging over time of the connectivity. Workmanship would explain the infrequent nature of this problem. Dan - 8-2-2016 - Tried out thigh board in known good system and was able to replicate brake chattering at least once. On subsequent freeing of the legs brake did not chatter though it did once after 5 minutes in the free'd state. RCA not precisely known though I suspect some design marginality in the boost converter that energizes the brake coil in which the supply might be hitting a current limit temporarily. Opened an engineering ticket for further investigation to this issue, ticket 2995 in EE/ME engineering database. Link here: https://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2995
3067 | Broken hip housing | closed | kmettler | Don-Doff Link | fixed | 2016-04-26 11:33 | 2016-04-26 | 322950 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Design | during the normal walking, this Ekso broke the left Hip, like the attached picture. | Parts sent per SWO-1940. Hip housings were replaced with new parts. | Design requirements did not properly anticipate customer use in terms of aggressive split stances during turning. RCA evaluated with CAPA #2489
3066 | Step Abort Issue | closed | kmettler | Foot Sensor | fixed | 2016-04-25 23:13 | 2016-04-25 | 130949 | 225 | Miami Project | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue: I work at The Miami Project to Cure Paralysis and we are working with our last subject for our current Ekso study. She is almost walking without me having to assist with her weight-shifting. Today during our training session the Ekso cued multiple step aborts, and always during a right-stance/left-swing step. It didn't seem like she was weight-shifting back toward the left to cause the abort to happen. We tried to video it but of course whenever we stopped the video it would happen. Is there a re-calibration that we need to perform or anything? Our subject comes back Wednesday. If we could work to fix this issue before that, that would be great! | Sent replacement foot sensors via SWO-1847 | Known wear issues due to the design. This is a low risk problem. The defect will be corrected in the next version of the device. Additional analysis to be done with Engineering Ticket 2998 then updated here. ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2998
3065 | Noise from Hip | closed | kmettler | Adjustments - Hip | fixed | 2016-04-22 20:40 | 2016-04-19 | 490797 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Wear | Issue: Not far into training our first participant (CJC), we heard a snap noise from around the left hip area. As the training progressed the sound happened with increased frequency until it was happening with nearly every step. We tried to record it, but the recording was not clear. Some snaps sounded louder that others; some you could feel through the handles or see at the distal thigh. It appeared to mostly occur around left terminal stance. | The old bolts were removed, the threads were cleaned, and new bolts were torqued in place. New parts provided via SWO-1878 | The bolts in the clamp bars (103068/103069) were loosened slightly, causing the bars to shift during step. Rivets were seen being bulled out of part 103067
3063 | unit #199 - EMAC/Italy - problem with right knee and foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-04-22 10:42 | 2016-04-22 | 14083 | 199 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | this Ekso showed a no identify error number on the RK. When I checked it I found that the left heel sensor floats inside the foot sensor. I did the foot calibration and it passed, but during the normal walking sometimes lose right step and two time showed an error on the RK. Could you maybe check on the Ekso DataDump the error number, because with the User and Passw that you gave us the access is denied, so I can't check. However, for sure is necessary replace the foot sensor. Afterthat let you know if the knee error will be showed again. | New sensors sent to distributor for installation via SWO-1848. | Down rev version of foot sensor allows the heel portion of the sensor to float into the toe sensing region
3061 | Side closure no longer holding | closed | lmcneil | Soft Goods - Torso / Sacral | fixed | 2016-04-21 21:35 | 2016-04-21 | 86635 | 135 | GF Strong - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: I'm finding the velcro strap for attaching the torso pad has been coming loose from time to time, is there a way to fix this or do we need new velcro straps for the Ekso? | Sent new side closures via SWO-1840 | Known wear item. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
3060 | DDL lever is tight | closed | kmettler | Don-Doff Link | fixed | 2016-04-21 16:06 | 2016-04-08 | 47569 | 180 | VA Oklahoma City | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Workmanship | Issue: The right swing away lever (DDL) is tight as well. No other issues. Soft goods are fine | Donn-doff links were updated to the current version with SWO-01823 | Found excess loctite within the lower DDL housing causing the customer facing issue.
3058 | unit 233 left foot turning | closed | ebabaran | Adjustments - Leg | fixed | 2016-04-21 11:29 | 2016-04-21 | 286952 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | The customer reported:"The problem of turning inside of the left foot occurs again more frequently. It appears with different patients and it‚Äôs always the left side. Walking 100 metres we have to adjust it up to 3 to 5 times. It happens both during straight walking as turning (left/right).The problem occurs with the two Ekso GT suits." | Required parts ship to EU via SWO-1854, relaced foot assemblies and upper shank | gap on clmap off upper shank not sufficent
3057 | unit 232 left foot turning | closed | kmettler | Adjustments - Leg | fixed | 2016-04-21 11:28 | 2016-04-21 |  | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | The customer reported: "The problem of turning inside of the left foot occurs again more frequently. It appears with different patients and it‚Äôs always the left side. Walking 100 metres we have to adjust it up to 3 to 5 times. It happens both during straight walking as turning (left/right).The problem occurs with the two Ekso GT suits." | Parts provided via SWO-1852, replaced upper and lower shanks | Tolerance stack up between OD of shank tube and ID of upper shank allowed for the possibility of this slippage. Parts provided to implement CO-11538 shown and larger OD on the shank tube used.
3053 | Foot Sensor and Velcro Issue | closed | kmettler | Foot Sensor | fixed | 2016-04-20 23:03 | 2015-04-20 | 215855 | 59 | Amanda Boxtel | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: One of our clients who has a lot of spasticity (Leah) made the left leg internally rotate a little yesterday. The Velcro on the right foot sole is not holding. I've duct taped it together. üòî hmmmm? Here are pics... | Sent replacement parts to customer via SWO-01839 | Customer use on carpet accelerates the breakdown of an already known wear item.
3052 | Left Foot Heel Dead | closed | kmettler | Foot Sensor | fixed | 2016-04-20 21:40 | 2016-04-07 | 79729 | 205 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Customer called and was having issues with their R step not triggering. Walked the PT through the back analog sensors and found that the left foot heel was not reading. | Sent new sensors via SWO-1837 | Known wear issues due to the design. This is a low risk problem. The defect will be corrected in the next version of the device. Additional analysis to be done with Engineering Ticket 2998 then updated here. ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2998
3048 | Unit #243 - Step by Step/Barcelona - rotating foot | closed | kmettler | Lower Leg | fixed | 2016-04-19 12:54 | 2016-04-18 | 324596 | 243 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | It continues to rotate the foot and he has done what you told him to clean and lubricate and replace the screws and it continues to rotate. | 25.04.2016 I exchanged the torx screws for new ones and add additional washers. I walked the Unit by myself for around 1200 steps without issues. Verified fix in place when device serviced via SWO-1928 | Stack tolerance issue in materials requires addtional washers to create necessary holding froce to prevent shanks from turning.
3046 | unit #197 - emac/Istituto di Montecatone Ospedale di Riabilitazione - error # 1180 LCD | closed | kmettler | UI - Controller | fixed | 2016-04-15 15:12 | 2016-04-15 | 296729 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | the customer called me saying that during the Ekso walking appeared the error # 1180 LCD. The Ekso stop to walk and it doesn‚Äôt answer at any kind of input from the remote control. The PT have sat the patient with the legs extended and put him out from the device. After that when they tried to switch on the device, the remote control display was completely white, without words on it. Also with the device off, was impossible to move the knees. I suggested them to remove the batteries and try and know they unlocked the knee. But when they retried to switch on the Ekso, now the remote control display the screen correctly, but when they press any button, nothing happen. Could be necessary replace the remote control? | Sent new UI via SWO-1835 | Part not returned for evaluation.
3045 | Hip makes noise | closed | kmettler | Other | fixed | 2016-04-14 15:08 | 2016-04-13 | 3748 | 260 | BG Klinikum Hamburg - Germany | Waltraud Kemper | w.kemper@bgk-hamburg.de | +49 40 7306 - 3670 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Process Not Clear/Complete | customer called with info about noise from the hip | added grease to ball bearings. No SWO created as no stock parts required to implement fix. | ball bearings on hip joints dry
3044 | unit #206 - Neurokinex/Injured Jocky Asociation Ekso (Rolling Ball) Ireland - rotating lower shank | closed | kmettler | Lower Leg | fixed | 2016-04-14 14:44 | 2016-04-14 | 241415 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | One of the PTs just contacted me stating that they had issues with a rotating lower shank (after 20-40 steps). | sent new t-wrenches via SWO-1936 | Torque wrenches in use were out of calibration
3039 | unit 232 broken screw on footplate | closed | kmettler | Foot Sole | fixed | 2016-04-12 11:15 | 2016-04-08 | 415505 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | unit has a broken screw on footplate | Replaced ankle housing. Addressed with parts from SWO-1826. Only #3038 listed on SWO for foot sensor replacement. Mention of this ticket added to comments section of SWO. | Found broken screw on footplate while replacing the foot sensor.
3038 | unit 232 step aborts, step lengh not ok | closed | kmettler | Foot Sensor | fixed | 2016-04-12 11:11 | 2016-04-07 | 415505 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | customer reported "right step length is really small. We have the experience that the 'step aboarded error' occurs more." | replaced foot sensors with stock from EU inventory via SWO-1826 | Known wear item
3037 | Noise on right knee | closed | kmettler | Actuators | fixed | 2016-04-12 11:00 | 2016-04-07 | 453136 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Customer reported "crackling sound on level of the right knee" right knee makes noise, ball screw makes noise and needs to be replaced, planned for 18.04.2016 | replacement of knee motor with parts taken from EU stock. Work performed via SWO-1854 | ball screw makes noise while moving the leg
3035 | unit #125 - Technomex/Poland - problem with foot sensor | closed | kmettler | Foot Sensor | fixed | 2016-04-11 12:35 | 2016-04-07 | 1456456 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | One of our units SN 021403006, steps: 1456456 had problem with foot sensor. I needed to replace both of them otherwise making foot calibration was impossible. | distributor replaced both foot sensors from there stock. Parts were provided to distributor via SWO-1851. | Foot sensors are a known wear item. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
3034 | Broken ankle | closed | kmettler | Ankle | fixed | 2016-04-11 12:24 | 2016-04-07 | 712099 | 109 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Loaner unit has problem with ankle ‚Äì therapist heard metallic sound Ankle needs to be exchanged. | customer (technomex) replaced latch assembly, swo 2115. | latch assembly broken. The rev A version of this part is known to break with regular usage.
3032 | Device cracked around the shroud | closed | kmettler | System | fixed | 2016-04-08 17:46 | 2016-04-07 | 114567 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue: I‚Äôm emailing in regards to our unit suffering a little crack. I was informed that one of our participant‚Äôs wheelchair knocked over the chair that the device was sitting in and it then tipped over the chair and unit that knocked into our Lokomat system. The cracked occurred on the right side of the cover by the power button. We do not have anyone scheduled today, but we do have patients scheduled for tomorrow. | Old shroud and parts removed from the torso and replaced with a new shroud. Completed with SWO-1822. | Customer dropped the unit and cracked the shroud.
3031 | Errors #1054 and #1241 RK | closed | kmettler | Electrical Interconnect | fixed | 2016-04-08 17:04 | 2016-04-08 | 3129 | 186 | Tecno Logica - Mexico | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Good Morning SN: 021501003 Customer # 0116 EKSO UNIT 0186 CRIT in Mexico (Serviced by TLM) reported errors 1054 RK repeatedly and error 1241 RK. This appears to be a cable issue, as communication was disrupted. Suggested cable replacement and take knee boards just in case. | SWO-01836 | Could not replicate, unknown root cause. Device was fully functional before and after service.
3030 | Cracked knee housing | closed | kmettler | Other | fixed | 2016-04-08 15:22 | 2016-04-07 | 1063644 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Materials | cracked knee housing | replaced knee housings including motors with parts provided by SWO-1820 | Tool chatter marks violated surface finish requirements and caused stress concentration at crack origin. Current production parts do show signs of this defect.
3029 | unit #232 - To Walk Again/Belgium - small step length right | closed | kmettler | Lower Leg | fixed | 2016-04-08 15:14 | 2016-04-08 | 415505 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Sometimes the right step length is really small. Perhaps I can fix this by doing a leg test? We have the experience that the 'step aboarded error' occurs more. | foot sensor calibration.No SWO requireed to complete this work as no stock components needed. | Foot calibration periodically required as routine maintenance. This is a low priority, low risk problem.
3028 | unit #233 - To Walk Again/Belgium - crackling sound in right knee extension | closed | kmettler | Actuators | fixed | 2016-04-08 15:08 | 2016-04-07 |  | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | There's a crackling sound at the end position in knee extension on the right leg side. I don't have a video of this, because I was not on site today. The robot does not give any error and we can still walk with it. | SWO 1932, replaced right knee motor | Ball screw performance degraded overtime resulting in unusual sounds
3026 | Noise from left hip motor | closed | uduell | Motor | fixed | 2016-04-07 10:23 | 2016-03-29 | 492941 | 54 | Prosperius - Italy | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | mechanical noise from inside the left hip motor | SWO-01692: Exchange CCA Covers of both Hips (B100115-101/201 Rev.:L) for new ones. Change of these and proper installation potentially removes pinch for wires which cause issues described by customer. | Unable to ensure root cause of issue with fixes applied but was able to eliminate customer issue. No presence of pinched wires communicated.
3024 | Stiff Leg Adjustment | closed | ebabaran | Adjustments - Leg | fixed | 2016-04-07 00:07 | 2016-04-06 | 513511 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Burke has been having issues with the stiff leg adjustment | Thigh bore was reamed per CO-11217. Completed per SWO-1821 | Thigh bore had not been reamed per CO-11217.
3020 | Foot Sensor Issues | closed | kmettler | Foot Sensor | fixed | 2016-04-05 21:10 | 2016-03-15 |  | 204 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: In addition to that our one device (serial # 021503008) does this abnormal thing after heel strike. The leg gets stuck in a few degrees of knee flexion and then pauses and extends for stance phase of gait(I will send you a video). The patient appears to be in good wt shifting position and it has happened with Kevin who is a super user. It also has thrown multiple step aborts. (I did do a sensor re-cal 2x in the past and it has helped the step abort issue) Other therapists also reported that the leg occ gets stuck in hip extension, just prior to swing forcing patients to go into stand still(I could not get that on video). | Foot sensors were replaced and the machine was recalibrated per SWO-1780 | Bad left foot sensor in the toe. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
3015 | unit 232 adjustment rod on left foot not fix | closed | mblaschke | Adjustments - Leg | fixed | 2016-03-29 11:51 | 2016-03-25 | 384771 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | adjustment rod on left foot not fix | replaced spring plunger assembly. Verify completion of fix via replacement of entire foot ankle assembly via SWO-1852 | loose spring plunger. Plunger can potentially wear over time and loose springiness.
3014 | unit 53 walkes "different" | closed | kmettler | Foot Sensor | fixed | 2016-03-29 10:27 | 2016-03-25 | 584258 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | patients mention unit walks "different" to the others | Foot sensor calibration performed. No SWO required as fix was done over the phone. | Foot sensor values can drift with regular usage.
3013 | unit 233 left foot sensor teared off | closed | kmettler | Foot Sensor | fixed | 2016-03-29 09:52 | 2016-03-25 | 413029 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | unit 233 left foot sensor teared off | replaced foot sensors in field via SWO-1793 | Regular wear for foot sensors
3010 | #168 Error Messages | closed | kmettler | Foot Sensor | fixed | 2016-03-28 16:33 | 2016-03-16 | 527862 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: We have been using the EKSO a lot lately. We are frequently getting an error code of #168; which causes us to abort gait. And we restart, sometimes, it errors right away again, other times it won't do again. Hoping you have some insight into this. I will be out of office the rest of this week, but I may check my email and I can contact the other users if there is something we should or should not be doing. | New foot sensors sent via SWO-1792 | Known Wear Items - Won't be able to perform true root cause analysis because we won't be receiving the sensors back.
3002 | Toe guard broken | closed | kmettler | Foot Assy | fixed | 2016-03-23 20:45 | 2015-06-24 | 76785 | 205 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: This device has a broken toe guard. | We a new toe guard and glue to replace the broken toe guard via SWO-1424 | Known wear item. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2998 | L Hip Joint Locked | closed | kmettler | Adjustments - Hip | fixed | 2016-03-22 20:16 | 2016-03-11 | 294696 | 24 | TIRR | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major |  | Wear | Issue: During my session the left hip joint began making a noise. I think you can hear it in the video. I stopped the session and after transferring the patient out I noticed the left hip joint was locked. Not locked in a abduction/adduction but locked and not moving in to flexion/ extension. The device was in free mode already so I turned it off and back on but not change occurred. I tried to make it sit to stand but it tells me to move the joints and I can move only right hip and knee and left knee. Since left hip doesn't move it won't engage for sit to stand. Any suggestions on how to get the left hip to work? | Fixed per SWO-1806. Motor and flexure were replaced on both sides. | Ball screw got locked up on the motor shaft.
2997 | Multiple Joint Test Req./ Software Issue | closed | kmettler | Software | fixed | 2016-03-22 20:07 | 2016-02-26 | 41753 | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com | (703) 408-7713 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Customer reported multiple joint test required errors/ Error #24. It has happened numerous times over a span of several weeks (2/26-3/18) | SWO-01815 Reset the VectorNAV frame rotation. | Out of date vendor software.
2996 | unit 128 hip width adjustment not possible | closed | kmettler | Adjustments - Hip | fixed | 2016-03-22 19:34 | 2016-03-22 | 150133 | 128 | Dennis Veit - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | customer walked a big patient and had to set the hip width to 20. Since the patient walked the unit the hip width could not be changed. | replaced latch pins. Unit fixed in field with parts via SWO-1555 | latch pin deformed and seizes up through regular usage.
2992 | unit 185 joint cap broken | closed | kmettler | Actuators | fixed | 2016-03-21 14:28 | 2016-03-21 | 397980 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | unit 185 joint cap broken | Implement new version joint cap via SWO-1783 Technomex used the new version of joint caps! | Down rev version of joint cap can fail with regular usage
2991 | unit 185 ankle screw broken | closed | kmettler | Foot Assy | fixed | 2016-03-21 14:22 | 2016-03-21 | 397980 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | ankle screw broken on right ankle | Right ankle replaced on SWO-01783 | Could be design or materials. Will see if there is a repeat issue.
2990 | feet touching each other independent of the hip abduction setup | closed | kmettler | Adjustments - Hip | wontfix | 2016-03-21 14:19 | 2016-03-21 | 320003 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | unit 176 feet touching each other independent of the hip abduction setup | Won't fix this is normal function of the device. Customer was notified this was normal function of device and no further action was to be taken. | Won't fix this is normal function of the device.
2989 | Pro Step + wouldn't initiate step | closed | kmettler | Foot Sensor | fixed | 2016-03-18 17:57 | 2016-03-13 | 502567 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | PT's walking in Pro Step + couldn't get the R leg to initiate even when they hit the appropriate targets. | Sensors replaced on device via SWO-1751 | Known wear issues due to the design. This is a low risk problem. The defect will be corrected in the next version of the device. Additional analysis to be done with Engineering Ticket 2998 then updated here. ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2998
2985 | unit #193 -Technomex - Problem with foot sensors | closed | kmettler | Foot Sensor | fixed | 2016-03-15 16:46 | 2015-03-04 | 101434 | 193 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Actually strange one. During walking device stoped reading from sensors which made walking impossible. At beginning of session all readings were OK. I tried to made auto foot calibration which failed. After a while readings were changing only in range from 2 to 25. | Info from Distributor: I exchanged foot sensor made calibration and made 200 steps in device. At this moment device works OK. Sent parts via SWO-1774 | Known wear issues due to the design. This is a low risk problem. The defect will be corrected in the next version of the device. Additional analysis to be done with Engineering Ticket 2998 then updated here. ---- https://berkeleybionics.trac.cloudforge.com/med_eng/ticket/2998
2984 | unit #053 - To Walk Again/Belgium - L Foot Sensor | closed | kmettler | Foot Sensor | fixed | 2016-03-15 15:57 | 2016-03-15 | 579370 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: We have the feeling that the left foot sensors doesn‚Äôt work accurately. It is a big difference for the patients to walk in this device or in the GT‚Äôs. | calibrated foot sensors already on device. SWO not required as new parts were not issued to device. | Foot sensor values can drift through regular usage. Foot calibration periodically required as routine maintenance. This is a low priority, low risk problem.
2983 | Ankle stiffness adjustment | closed | kmettler | Other | fixed | 2016-03-15 15:54 | 2016-03-15 | 352380 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | The adjustment for the ankle stiffness on the left side doesn‚Äôt work. If we want to set them on ‚Äú3‚Äù for example, it does not stay in this position. There‚Äôs a hesitation in triggering the right swing phase when supporting in stance phase on the left side. The patient has reached sufficiently the targets which are desired but then there‚Äôs a hesitation in the initation of the right swing phase. This problem doesn‚Äôt occur always but it changes between patients and also within one session. | Adjustment rod on left ankle lose, replaced spring plunger assembly, swo 2114 | Adjustment rod on left ankle lose. Existing plunger found to not engage sufficiently. Likely on spring of plunger being able to hold force.
2980 | L leg swing backwards | closed | kmettler | Foot Sensor | fixed | 2016-03-14 21:52 | 2016-02-25 | 123233 | 64 | Dan Rose |  |  |  | USA |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: Issue: My device recently started to swing back with the left foot before stepping forward. It doesn't happen every step and sometimes it kicks back further than what I was able to capture in the attached clip. Sorry for the poor lighting and angle of the clip, it was difficult to find a way for my spotter to spot and capture what was happening on film. I was wondering if there was something we could do to fix the problem. | sent new foot sensors via SWO-1746 | Known wear item. Foot calibration or sensor replacement periodically required as routine maintenance. This is a low priority, low risk problem.
2978 | Hip abduction not fix | closed | kmettler | Adjustments - Hip | fixed | 2016-03-11 14:26 | 2016-03-11 | 23362 | 235 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Hip abduction setting changes while walking in device. | added vibratide to set screws to ensure set keeps proper engagement with adjustment mechanism. No SWO created as no additional part required. | Set screw can move as no thread locker is specified for the assembly.
2977 | Leg internally rotating | closed | ebabaran | Adjustments - Leg | fixed | 2016-03-10 18:05 | 2016-02-18 | 197725 | 59 | Amanda Boxtel | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Customer Use | Issue: Over the past couple weeks they have been having issues with the leg internally rotating. It has interfered with several different patients they‚Äôve had. All of them were within our spec‚Äôs. We decided to bring their device back since we had walked them through a number of things and it was still happening. The Loaner device arrived with them last Wednesday. Amanda walked in the device and everything was fine. Another patient walked on Friday and she said that the leg was internally rotating again. | Foot sensors changed out on device for non-treaded version. Fixes applied by SWO-1758. | Customer is walking on carpeted area
2976 | Unit #258 - PASORI/France - left leg tends to rotate/tubes are hard to slide | closed | kmettler | Other | fixed | 2016-03-10 16:20 | 2016-03-09 |  | 258 | Pestori - France | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 7 - Clinical Visit | Minor |  | Materials | The distal part of the left leg tends to rotate on the left when in use ( the foot ends up in medial rotation of several degrees) even with the correct setting ( 2 clics) Also when we are setting the tubes they are hard to slide. | Sandy added third washer under screws. As device issue was addressed during clinical visit by the adding of additional washers no SWO was written to address the completion of this work. | Tolerance stack ups can allow for this to happen even with everything in spec.
2971 | Joint cap failure | closed | kmettler | Actuators | fixed | 2016-03-09 05:48 | 2016-03-08 | 517018 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | Issue: The customer did not notice that the cap was no longer flush with the Ekso Customer Email - Yesterday the washer that houses the allen bolt for the hip broke. I believe that its purpose is to hold the cylinder in place that‚Äôs inside the hip joint allowing it to pivot. I wanted to see if we should send our unit in to be updated now if you guys are ready or just have a replacement part sent. Please give me a call when you have a chance, my personal cell 509-741-7500 or they can transfer you to my ascom at 509-665-6156 option 1. | Joint caps changed in field via SWO-1762 | Down rev version of joint cap can break during regular usage. New version implemented via CO-11476
2968 | unit #160 - Moritz Klinik German clinic (EU) - broken toe guard | closed | ebabaran | Foot Assy | fixed | 2016-03-08 15:57 | 2016-03-08 | 3133 | 160 | Moritz Klinik - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | broken toe guard | replaced foot sensors, SWO-01760 | broken while walking a patient. Known wear items. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2967 | RH/RK pot encoder errors | closed | kmettler | Encoder / Potentiometer | fixed | 2016-03-07 19:37 | 2016-02-18 | 10749 | 227 | Children's Hospital of Los Angeles | Lauren Dunn | ldunn@chla.usc.edu | 949-412-0314 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Materials | Issue: There‚Äôs some RH and RK pot/encoder mismatch errors that are outside of our allowable range‚Ä¶ Nto sure what‚Äôs causing it. Looks like they ran a leg diagnostic and failed at least once (2/16). The PT stated the errors came up several times after L Turn. They were slowly pivoting a patient. The errors always came up in ProStep + at the exact same spot on 2 occasions. It was a max of 60 feet from the start. The motors were really hot after each error with a funny smell. The first time the error came up at the end of a sit to stand with a R Hemi Patient, in crutch mode. The patient was impulsive at times. The 2nd time was with a thin & gentle patient. The 3rd time the error came up was with a PT. | Replaced encoder per SWO-01775. | Found that by replacing encoder the issue was no longer present. Left part in engineering RCA box. DB - 4/8/2016 - Removing the encoder is destructive so no further analysis can be done on the part. That being said, this is an infant mortality an electronic part that we buy off the shelf so there are no design implications on our side.
2965 | Issues with locking the right hip | closed | kmettler | Adjustments - Hip | fixed | 2016-03-02 11:32 | 2016-03-02 | 191536 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | We are having issues with locking the right hip and we heard a loud snapping sound when walking a patient this morning. | pressed drill bushing out and pressed it in again using Loctite 609. Device fixed via SWO-1750. | drill bushing on ddl moved out a bit. Failure of loctite sufficiently holding drill bushing in place.
2964 | unit #45 - Justin Smith / South Africa - error 168 | closed | kmettler | Foot Sensor | fixed | 2016-03-02 10:09 | 2016-03-01 | 736866 | 45 | Justin Smith - So. Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Our unit is giving us hassles again. Giving lots of 168 errors. To the point where the sensors on right foot doesn't pick up during pro step. I have attempted to to calibrate. All is fine apart for the right rear sensor. It doesn't want to calibrate | Sent new sensors via SWO-1756 | Known wear item. Foot calibration or sensor replacement periodically required as routine maintenance. This is a low priority, low risk problem.
2958 | Broken joint cap | closed | kmettler | Actuators | fixed | 2016-02-25 16:17 | 2016-01-25 | 315792 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Major |  | Design | unit #185 -Technomex Ergo Hestia - broken joint cap | Sent parts to distributor via SWO-1741. | Down rev design of joint caps known to break through regular usage.
2956 | Foot Sensor issue on device | closed | kmettler | Foot Sensor | fixed | 2016-02-24 23:00 | 2016-02-23 | 262545 | 24 | TIRR | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: I was working with an individual yesterday and we had 2 different situation where Ekso was acting up. The first occurred when we were in L affected so her R leg was free without motors. She would advance her R limb at her own pace and power. However there were at least 3 instances when the R leg would freeze. It was as if she was beginning to advance the limb with her hip muscles and the leg would go into a flamingo stance with the hip in neutral extension and the knee in flex. The foot would actually rise up and hit me in the thigh. If we didn't go immediately into stand still the R leg would begin to quiver our bounce. I wasn't able to discern a trigger or cause and thought it odd since the motors should have been off. The second issue was when we trialed R affected so the motors on the L were off. It was as if the motors on the R were not kicking and recognizing she had achieved her weight shift onto the L. We had once instance where she met her weight shift left (I heard the lateral beep) and the patient began to step the right on her own power and had completed an entire step before Ekso suddenly triggered and then began to step her, almost pulling her leg forward and out into a splits position. I hit stand still and we had a hard landing to the ground with the R leg. Afterward it would not let us start again. I shifted her to the L and tried to trigger the first R step and it would make a noise (like a beep but a lower pitch to it). We changed modes, putting it back in B mode and then back to R affected and then it worked. This happened on 2 separate occasions. My only thinking on what may have caused it was if maybe the patient had not fully shifted lateral and forward as I only had the lateral beep activated. Thanks for any feedback on what may have caused the above. I will be working with her again today so we will see if these issues occur again. | Sent new foot sensors via SWO-1739 | Known Wear Item. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2946 | Step By Step Error 1057-RK Unit 242 | closed | kmettler | Software | fixed | 2016-02-19 12:36 | 2016-02-18 | 16921 | 242 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Determined Unknown | Customer: Step by Step, Barcelona , Spain Error #1057-RK on Unit 242 (0215010005) From the customer:''' Google Translated''' Cuando llevamos unos 20 minutos utiliz√°ndolo da un error de software # 1057- RK y se desbloquea la pelvis por lo que no podemos continuar y tenemos que hacer una bajada de emergencia. Ahora mismo tenemos las dos unidades de Ekso inutilizables porque esta no funciona y la otra unidad la cadera se sigue abriendo. Por favor ponte en contacto conmigo urgentemente. When we use it gives about 20 minutes a software bug # 1057- RK and pelvis so we can not continue is unlocked and we have to make an emergency descent . Right now we have two units Ekso unusable because it does not work and the other unit continues to break her hip . | Amp board replaced on device via SWO-1736 | Board has been removed from device and will be returned to HQ/vendor for further evaluation. PCBA returned to HQ and given to engineering for further valuation. The 1057 / 1065 error has occurred on multiple devices. The issue is being tracked and resolved with this ticket: ‚Äãhttps://berkeleybionics.trac.cvsdude.com/med_eng/ticket/2994
2943 | unit 13 Mark Beak South Africa, foot sensor broken | closed | kmettler | Foot Sensor | fixed | 2016-02-18 09:49 | 2016-02-11 | 272279 | 13 | Mark Beack - South Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Design | foot sensor broken | replaced foot sensors via SWO-1717 | Foot sensor can break down with regualt usage
2941 | unit 185 (Technomex) thigh straps worn out | closed | kmettler | Soft Goods - Upper Leg / Pelvis | fixed | 2016-02-18 09:26 | 2016-02-11 | 710326 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | thigh straps worn out | sent new thigh straps via SWO-1710 | Known Wear Item. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2938 | Hook on torso handle broken | closed | kmettler | Torso | fixed | 2016-02-18 09:16 | 2016-02-17 | 185971 | 65 | Dale Messenger - UK | Dale Messenger | dalemessenger@btinternet.com | +44 7833 456767 | United Kingdom |  | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Design | Hook on Torso Handle is broken | Replaced the torso handle with newer 1-piece version via SWO-1712. | Previous two part design of handle can break as hook was glued in place.
2937 | Unit 232 To Walk Again, joint cap broken on left hip | closed | kmettler | Actuators | fixed | 2016-02-18 09:07 | 2016-01-29 | 226385 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | (Describe issue, including geographic location, if appropriate) | Replaced all inner joint caps per SWO-1697 | Downrev version of joint cap can break
2934 | BGU Duisburg, right foot sensor faulty, toe guard loose | closed | kmettler | Foot Sensor | fixed | 2016-02-17 15:04 | 2016-02-17 | 500000 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | BGU Duisburg, right foot sensor faulty, toe guard loose | replaced foot sensors, found broken shoulder screw on right ankle. Replaced ankle housings on both sides (couldn't get the screws out to mount the bearing plates). Work done via SWO-1729. | Foot sensors can wear/break down through regular usage. Toe guards can break off with regular usage.
2933 | unit #232, epoxy broken on left shank | closed | kmettler | Lower Leg | fixed | 2016-02-17 11:07 | 2016-02-15 | 3138 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Process Not Available | epoxy broken on left shank, foot is turning | replaced lower shank. Had to use a used one from unit 233. Work performed via SWO-1728 | Previous MPI allowed for variation in process.
2927 | Errors 22,24 and 25 LH | closed | kmettler | Encoder / Potentiometer | fixed | 2016-02-15 14:35 | 2016-02-15 | 1170339 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Materials | Unit 129 shows errors 22,24 and 25 LH. The error occurred again after a leg test. | replaced both knee housings via SWO-1748 | Both knee housing had cracks. Crack origins were examined and housings were found to have tool chatter marks which caused stress concentration at crack origins. Production housings were examined and found to have no tool chatter marks in same area.
2923 | battery charger pin bent | closed | kmettler | Battery - Charger | fixed | 2016-02-08 19:27 | 2016-01-28 | 436057 | 136 | Kessler |  |  |  |  | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Mishandling | Issue: One of the pins in one of our battery chargers is bent. | Sent new battery charger via SWO-1705 | Critical pin of battery charger was bent.
2922 | Batteries | closed | kmettler | UI - Crutch Interface | fixed | 2016-02-08 19:22 | 2016-01-28 | 436057 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | I have a couple of minor issues here: 1.) The cable at the attachment to our crutch interface is pulled away/exposed. | Sent new UI via SWO-1705 | Wear item. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2909 | Issue with R Knee Brakes | closed | kmettler | Lower Leg | fixed | 2016-01-29 17:12 | 2016-01-27 | 331977 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: Hi Josh, Never a dull moment here - at the end of today's training session with EKSO GT in R affected mode we encountered an episode where the R knee brakes seemed to be coming on throughout the swing phase. Video is attached. I will be calling you shortly to discuss. | Troubleshooting determined that the Knee PCBs (Breakout board and CCA) were causing issues consistent with the complaint. Boards will be replaced. SWO-01694. | Suspect low voltage to brake, PCB component not measured, but suspected. Brake functions as intended directly wired.
2908 | Hip adjustment fell out | closed | kmettler | Adjustments - Hip | fixed | 2016-01-28 17:55 | 2016-01-21 | 482877 | 9 | Craig | Julie Hartman | jwaldie@craighospital.org | (303) 789-8000 | USA | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Customer Use | Issue: The piece that holds the hip adjustment in place fell out. | MPF, 1/28/16: I found that the right side torso clamp nut had moved out of it's slot and was loose inside the dog house. I re-installed the clamp nut per SWO-1698. | Block drifted during usage due complete disengagement of adjustment bolt. In speaking with customer device would be put on side to help loosening up the hip width adjustment. Doing this while fully disengaging the hip width adjustment bolt allowed for this to happen. Customer was shown different techniques to help with hip width adjustment and told to never fully disengage the hip adjustment bolt.
2903 | Noisy R hip Motor | closed | ebabaran | Motor | fixed | 2016-01-26 17:23 | 2016-01-26 | 178070 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue: The past few sessions using the Ekso we have had a very noisy right motor, and it is getting significantly hotter than the left motor. This is for patients with both complete and incomplete spinal cord injuries so there is no reason why there should be more stress on the right motor. It is happening after about 20-30 minutes of walking. Do you know what we should do to fix the problem, or can someone come to have a look? It may involve a 30 minute walk to reproduce the problem. | Did an elmo calibration. No SWO created as no parts needed. | the hotter motor was caused by a heavy patient leaning far to the right. Another patient and myself caused normal temperatur allocation on the motors.
2902 | Left shank is slipping with one patient | closed | kmettler | Adjustments - Leg | fixed | 2016-01-26 16:39 | 2016-01-26 | 87177 | 225 | Miami Project |  |  |  |  | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | When talking with the PT we were able to facetime with the customer and we could tell the left shank slit wasn't closing as tightly as the right side. This allowed situations where the L Shank would begin to slip every session w/ their patient... They are trying to qualify this pt for a study but can't walk long enough to do so due to the slipping shank. | MPF, 1/27/16: Upon arrival it was possible to rotate the left shank by applying a large amount of torque to the footplate. The same amount of torque would not rotate the right shank. Upon inspection it was found that the bore of the left upper shank was out of spec (too large). I replaced the left upper shank with new. Completed with SWO-01687 | Parts were orignally reworked per NCR-11549 due to step found in main bore of component. Upon rework and receipt via PO 5672-01 we were able to confirm that effected issue addressed in NCR-11549 was measured but not the actual bore dimension.
2894 | battery is not charging | closed | kmettler | Battery | fixed | 2016-01-21 21:46 | 2016-01-05 | 137907 | 6 | Thevenaz - Switzerland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue: Monday the robot stop And the parts Up of the robot move down And the computer write " left battery empty". And after the computer of the robot write 76% (battery). But he was very dangerous because he stop just Like that... I think its good to have A people come to See the robot. | New battery sent via SWO-1649 | Additional analysis to be done with Ticket #2384 then updated here.
2888 | Foot Sensor Tear | closed | ebabaran | Foot Sensor | fixed | 2016-01-20 21:33 | 2016-01-15 | 179119 | 59 | Amanda Boxtel | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Issue: I was rushed in between clients taking these photos and now realize they are a bit blurry. Sorry. Both toes on the red soles are beginning to peel away. Red sole replacements would be optimal. | Sent foot sensors via SWO-1669 | Wear item. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2884 | User interface message is unclear (Spanish translation) | closed | duanee | Documentation / Labeling | fixed | 2016-01-20 16:50 | 2016-01-13 | 4216 | 216 | Tecno Logica - Mexico | Lina Ariza | ingenieria@tlmandina.com.co | +57 1520-7920 | Mexico | no injury | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor | Investigation Needed | Customer Training | Issue: We seem to have a bug in the crutch stand LCD message. All three patients today have had the same message after holding for stand (pics attached). When I tried to recreate at lunch, it displayed normally. Out configuration is Bil, Adapt, Pro+, crutch interface used. Despite the incorrect message, the device began sit and paused at 90, then stood normally. I took a video of the controller during the stand. | No additional resolution is necessary for this customer other than the verbal clarification about the device behavior. The translated message on the controller was changed with the 2.1.0 release of the language file in CO-11660. | (see Investigation Summary) | There was some amount of confusion about the nature of this issue initially. The root cause of this issue was initially thought to be a bug in the device software, however, upon further analysis, it has been determined that the core issue has to do specifically with the abbreviated Spanish translation of one of the user interface messages on the device controller. Translations are not handled in the device software directly, but are maintained in a separate language file. The display message "Begin Sit Up" in the English version of the display was translated to "Iniciar sent dcho" (abbreviated version of "Iniciar sentarse derecho") in the Spanish translation as seen in the attached photo of the controller. In this case, the translated "Begin Sit Up" message was misunderstood to mean "Begin Sit", or similar, causing confusion about the action the device would take. The operators correctly understood from training what the action would be and noted that the controller message was unclear or incorrect. The confusion may, in some ways, be due to differences in the version of Spanish used for the translation on the controller and the local Spanish spoken in the area of this facility. Initially, the reported behavior could not successfully be reproduced because the initial exploration was done with the controller in English. The translated message in question was updated in the 2.1.0 release of the language file (CO-11660). No further action is necessary. (D. Edmonds ‚Äì 12 Oct 2017)
2883 | Error #27 and #21 | closed | kmettler | System | worksforme | 2016-01-20 07:00 | 2016-01-12 | 133000 | 202 | Marianjoy Rehab Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Determined Unknown | Issue: MarianJoy GT - steps about 133,000 At approx. 4:10 pm CST, walking in First Step, Bil, Adapt, We got error #27 and then immediately #21 on Left Knee (pretty sure it was knee, not hip), resulting in Knees Locked error. Cleared both, rebooted, able to stand and resume as normal. Flagging this now in case there is recurrent concern. Katie, let me know if we need to do anything. | n/a | did not recur.
2882 | Toe guard fell out | closed | kmettler | Foot Assy | fixed | 2016-01-20 04:06 | 2016-01-19 |  | 240 | Advanced PT | Valerie Gibson | gibsondpt@gmail.com | (703) 408-7713 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: Advanced PT called and reported that the toe guard fell out from in between the foot plate and sensor. Pictures attached | Sent customer RP Gel in order to secure the toe guard | Known wear item/ Customer use. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2881 | Battery not holding a charge | closed | kmettler | Battery | fixed | 2016-01-19 20:01 | 2016-01-15 | 480210 | 9 | Craig | Julie Hartman | jwaldie@craighospital.org | (303) 789-8000 | USA | no injury | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | John a PT a Craig hospital reported that a battery is no longer charging | Sent new battery via SWO-1674 | Additional analysis to be done with Ticket #2384 then updated here.
2880 | Errors 121, 133, 140 on unit 147 (Technomex) | closed | kmettler | Electrical Interconnect | fixed | 2016-01-19 15:50 | 2016-01-18 | 1312713 | 147 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: shows errors: 121, 133, 140 which make session impossible. This errors are related to engine. Please advise what can be done. I do not know if problem is related to electronic, wire or motor itself. | Unit fixed in field by replacing hip housing during service conducted via SWO-1678. Parts to address hip housing crack which is believed to be related to the root cause for the issues were provided via SWO-1342. | Additional analysis done with CAPA #2489
2879 | loud noise from right hip on unit 125 (Technomex) | closed | kmettler | Other | fixed | 2016-01-19 15:47 | 2016-01-18 | 444820 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Loud noise from right hip | Replaced motor on device per SWO-1683. | Unable to get faulty motor back for further evaluation.
2873 | Device will not turn back on | closed | kmettler | System | fixed | 2016-01-08 18:22 | 2016-01-08 | 88676 | 196 | Children's Hospital of Atlanta | Amber Richards | arichards@chla.usc.edu | 323.361.2568 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Issue: Issue: The customer called this morning and reported the following‚Ä¶ While walking with a patient, one leg in Free and one Leg in Fixed the device lost power. The knees were in a bent position but the PT was able to safely sit the pt on a nearby chair. The Ekso won‚Äôt power back on after multiple attempts. We plan on sending a loaner unit to the customer today. | MPF, 1/20/16: Measured voltage across pins 5 and 6 on PWR SW header. Jumpered pins 5 and 6 but unit still did not power up. Tested a known good CPU board but did not power up. Replaced the amp board per SWO-1675. Unit powered up without issue. I bagged the old amp board and placed it in the ENG RCA Review bin. Dan Brown confirms that Elmo breaker component Q16 is failed short, a know cause of failure to power up. DWB - Board reworked at UMAI DWB - Board fixed and ran 102284 refurb verification test | Dan Brown confirms that Elmo breaker component Q16 is failed short, a know cause of failure to power up. Fix is to motor power circuit breaker design that will be on a future iteration of 102915 AMP board (see ticket #2464 for details)
2870 | battery draining at WakeMed | closed | kmettler | Battery | fixed | 2016-01-06 23:03 | 2016-01-06 | 71194 | 33 | WakeMed |  |  |  |  | no injury | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue: One of the batteries at WakeMed was draining and not holding a charge. | Sent new battery via SWO-1645 | Additional analysis to be done with Ticket #2384 then updated here.
2869 | Battery draining quickly | closed | kmettler | Battery | fixed | 2016-01-06 23:01 | 2016-01-06 | 203150 | 24 | TIRR |  |  |  |  | no injury | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue: One of the batteries at TIRR was draining and not holding a charge. | New battery sent via SWO-1644 | Additional analysis to be done with Ticket #2384 then updated here.
2864 | Walker with one leg that is slightly shorter | closed | najero | Cane / Walker / Crutch | fixed | 2016-01-06 18:52 | 2016-01-06 | 15569 | 228 | Sister Kenny | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Determined Unknown | Issue: I am a physical therapist that has been working with the Ekso Bionics system at Courage Kenny Rehabilitation Institute in Minneapolis. Our Ekso walker has one leg that is slightly shorter than the others. It has been this way since our training week, but we didn‚Äôt fuss too much with it (and honestly thought the hallway we were working in may have had a slight slant/grade). Now we are working in a different hallway with many patients and nearly all of them complain about the unstable feel of the walker. Do you think that we could get a replacement walker? Please let me know how we could go about doing this. | Sent walker via SWO-1642 | The walker did not come back to headquarters for analysis as it was not cost effective. A new walker was sent to the customer with no issues. Walker is not available for root cause analysis. I can only speculate that the walker may have been damaged during shipping since the complaint was initiated during initial training. The past 4 lots of walkers we received have been 100% inspected. Incoming Inspection verifies the legs are even and stable. Thus far, no complaints of uneven legs have been reported.
2860 | Stiff left DDL on Ekso Unit 0048 - Bad Aibling - Germany | closed | kmettler | Don-Doff Link | fixed | 2015-12-31 11:16 | 2015-12-28 | 824395 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Customer reported a stiff left DDL. The unit cannot be used. The left DDL is brand new and worked well as we replaced it. | No SWO created. Grease provided to customer and they were walked through the process for resolving the issue both for this instance and for any future occurrence. | Through usage bullet in DDL assembly can begin to bind up. Adding grease to the back side of the bullet when depressed and cycling the assembly helps to resolve this issue.
2859 | Batteries only charging to 29% | closed | kmettler | Battery | fixed | 2015-12-23 19:34 | 2015-12-23 | 258192 | 18 | RIC | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Issue: Recently we began using our original Ekso system more frequently and we are having problems with the batteries. One set of them is only charging to 29% and I think the other set might not charge fully either but goes much higher. This is despite the batteries being on the chargers over night.They also drain incredibly quickly- but this was happening before as well. Is there anything we should do differently or is it a possible to get a new set of batteries for the ones that will not take a full charge? | Sent new batteries and charger via SWO-1630 | Additional analysis to be done with Ticket #2384 then updated here.
2858 | Left foot sensor broken | closed | kmettler | Foot Sensor | fixed | 2015-12-23 10:24 | 2015-12-23 | 170748 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Eva reported the unit is walking "strange" even while being walked by a physio, it seems the left foot sensor doesn't work right. Calibration ended always with error 171. | Replaced foot sensors per SWO-1629 | Left foot sensors are non-mirrored versions of the right sensors which allows the solder joint to be downwards facing on the left side.
2846 | Foot sensor left heel not reading | closed | kmettler | Foot Sensor | fixed | 2015-12-17 14:30 | 2015-12-15 | 248031 | 170 | FLM - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue - Maria Customer reported multiple #168 error. Both auto and manual foot cal were unsuccessful. | Exchange Foot Sensors for new ones. Work accounted for via SWO-1803 | Glue went off. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2845 | Left foot Epoxy failure | closed | kmettler | Foot Assy | fixed | 2015-12-17 10:52 | 2015-12-16 | 141123 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Clear/Complete | Issue - Maria Customer reported Epoxy failure on left foot of Ekso. | Shank mount assemblies changed in field per SWO-1611 | Variability in previous process allowed for this failure.
2838 | Toe Guard Broke | closed | kmettler | Foot Assy | fixed | 2015-12-14 19:05 | 2015-12-14 | 157251 | 134 | Chuck Slyngstad | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Issue: The toe guard from Chuck's device broke. Replacement will be sent. | Sent new toe guards via SWO-1600 | Known Issue. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2837 | Ankle Adjustment Slider Broken | closed | kmettler | Adjustments - Ankle | fixed | 2015-12-14 18:39 | 2015-12-11 | 318624 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: Customer was walking a patient and heard a noise‚Ä¶ after the noise they noticed the foot could freely go through the range of motion and was no longer being controlled by the knob on the back. It doesn‚Äôt go further than intended range but it won‚Äôt hold any value now set by the PT. I‚Äôve told them to stop using the device. | Installed new ankle adjustment sliders on both left and right feet per SWO-1599. | Damaged parts returned for further investigation
2834 | Error #1240 | closed | kmettler | Software | wontfix | 2015-12-09 21:13 | 2015-12-09 | 303624 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Design | Technomex reported that Ekso threw error #1240 a few times in different occasions (cannot confirm when). This happened with different patients/ walking modes. The PT had to restart the device. | Informed customer - No additional action taken works as designed. | Device operates as designed. This is a known bug that does not effect usability of device.
2832 | L and R Foot strap open too easily | closed | kmettler | Foot Assy | fixed | 2015-12-09 20:53 | 2015-12-09 | 1044939 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue - Maria Technomex reported that L and R Foot strap open too easily | Foot binding sent to customer sight via SWO-1595. | Low cost wear component that can be easily replaced. No further analysis required.
2830 | Broken knee cover | closed | kmettler | Actuators | fixed | 2015-12-09 20:43 | 2015-12-09 | 1044939 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Mishandling | Issue - Maria Customer reported broken knee cover | Knee cover sent to customer via SWO-1595. | Cosmetic plastic cover can be easily broken if any excessive force seen by component.
2825 | Broken left and right adjustment knobs of ankle dorsiflexion | closed | kmettler | Adjustments - Ankle | fixed | 2015-12-05 01:14 | 2015-12-04 | 146771 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | Issue - Maria Matteo at EMAC reported that the adjustment knobs of the ankle dorsiflexion do not work in the same way. The right ankle seems to be looser than the left one. Today the screw and spring of the left knob came out and the PT tried to put screw back but it broke and got stuck inside. It seems that both sides will need to be replaced. | Work addressed via SWO-1594 | Down rev version of this assembly used only medium strength blue loctite to secure this screw. Process has since been changed to used high strength red locitie
2824 | Rotation lower shanks | closed | kmettler | Lower Leg | fixed | 2015-12-05 00:53 | 2015-12-04 | 146771 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue - Maria Matteo at EMAC reported that during walk the shanks rotate inwards (on the right side more than the left one) | Parts sent to customer via SWO-1594. This was added in the comments section of SWO in omnify as SWO was already closed when attempting to make this link. | Adjustment screws can bottom out on back wall of shank collar.
2820 | Faulty batteries | closed | lmcneil | Battery | fixed | 2015-12-04 23:26 | 2015-12-04 | 79417 | 30 | Merryweather - So. Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue - Maria Customer reported 3 batteries out of 4 are not working. The batteries were not charged for about 2 months. Only one battery seems to be working and if plugged to the right or left side it always shows as charged to 98%. The customer tested this battery combined with the other 3 and the Ekso throws a Missing Battery error on the side in which any of the 3 batteries are connected. They tried charging all batteries using both chargers and with both of them the light is immediately green. The device is now down as they don‚Äôt have a working set of batteries. | Talked to customer and problem did not occur again after first occurrence. This was a one off, user error. | Was not able to reproduce after first occurrence.
2819 | Epoxy failure Right foot | closed | kmettler | Ankle | fixed | 2015-12-04 16:36 | 2015-12-04 | 81540 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Process Not Available | Issue - Maria Epoxy failure on right foot. The customer noticed rotation of the foot whilst walking a patient so they stopped the session. I have asked them to stop using the device immediately. Udo is going on site early next week to fix this unit. | Unit fixed in field with new stuffed shank mount assemblies via SWo-1586 | Previous version of gluing allowed for too many inconsistencies.
2817 | Hip Motor Getting Warm | closed | kmettler | Motor | fixed | 2015-12-03 17:17 | 2015-12-03 | 97239 | 232 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | Issue - Maria During session with patient, PT noticed left hip motor getting warm. | performed elmo calibration. No SWO created as no parts required to resolve issue. | Drift of values governing the elmo calibration and motors cause sub-optimal performance of system.
2813 | Error #140 | closed | kmettler | Foot Sensor | worksforme | 2015-11-24 16:22 | 2015-11-23 | 224072 | 156 | Kerpape - France | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue - Maria The customer reported step issue. See video attached. They have performed the foot cal 3 times but each time instead of getting #171 or #170 they got error #140. Could you please advise. They have a TV going there on Monday and the Ekso needs to be working. The foot sensors are visibly new. | Informed customer. No further action required, error occurred due to customer use with no reoccurance. | Error Code 140 occurs when a test is initialized when a test is currently being conducted
2812 | Limping issue walking in Pro Step + | closed | kmettler | Foot Sensor | fixed | 2015-11-24 11:58 | 2015-11-23 | 160782 | 163 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue - Maria Mark Pollock reported limping problem whilst in Pro Step Plus. It is like the heel sensor is not sensing when the left heel is up. | Went on site via SWO-01504 and walked customer through the recalibration of foot sensors on device. | Wear on foot sensors allowed values to drift out of usable range
2811 | Error #133 | closed | kmettler | Electrical Interconnect | fixed | 2015-11-24 11:37 | 2015-11-23 | 142924 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue - Maria Error #133 occurred with a one patient at the end of the session - after 25 minutes walking in Pro Step+ adaptive. The patient has some spasticity. With regards the patient specific data, Colin will gather the details. He will also send me a video later on today. He is trying to restart the device but the legs won‚Äôt freeze. | Signal for entire hip assembly changed in field via SWO-1573 | Unable to perform root cause analysis as this type of repair troubleshooting eliminates opportunity to get this information
2808 | Stepping Issue | closed | kmettler | Foot Sensor | fixed | 2015-11-20 17:27 | 2015-11-19 | 65000 | 167 | Children's Hospital of New Brunswick | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Josh - I have a question regarding our EKSO. We have been using the device with a patient over the last 4 weeks. In the beginning we would get a display reporting ‚Äúright foot stepping‚Äù not allowing the patient to take a step with their left leg only occasionally. This last session we receiving approximately 7-10 of these messages. They were in Pro Step mode every time. We think it might be tone/spasm related limiting the ability to extend his right knee. It is just interesting as his weight is fully weight bearing through the right extremity at this time. There also seems to be a delay to when EKSO kicks in to assist the patient in achieving right knee extension allowing him to then take a step with his left leg. We tried increasing his swing time thinking it might help if it was spasm related and it didn‚Äôt help, almost seemed like it made it worse. Any feedback or advice would be great. | We installed the latest foot sensors to the unit per SWO-1578. | Individual cause of issue unknown. To be evaluated with larger assessment of failed sensors. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2806 | Error #21-RH and #27-RH | closed | kmettler | System | fixed | 2015-11-19 12:45 | 2015-11-18 | 53989 | 233 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue - Maria On one of our new devices we had a sensor error RH-#21 followed by a sensor error RH-#27. This happened while we were starting to walk (the first step of the session). It appeared twice (on Tuesday and today). Yesterday we just had to switch off and on the device. Today I did a Leg Test followed by a torso test and the problem didn‚Äôt reoccur anymore. Is there an issue? Height and weight of the patient Pt 1 : 186 cm; 63 kg Pt 2 : 164 cm; 79 kg Level of spasticity Pt 1 : average 1 (on MAS) Pt 2 : average 0 (on MAS) | replaced hip cable, encoder, encoder cable an fan cable via SWO-1571 | Unable to perform root cause analysis as this type of repair troubleshooting eliminates opportunity to get this information
2805 | Error #21-LK | closed | kmettler | System | fixed | 2015-11-19 12:22 | 2015-11-19 | 515880 | 54 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue : Maria Customer reported error #21-LK. Please see videos attached. They seem to have issues with doing the leg test. I advised them to try again with the Ekso on a tether pushing until hard stops. They have tried several times, but it failed. Currently they cannot use the device. | Replaced left knee, both foot sensors, encoder and cable (left hip) via SWO-1574 (Left knee signal changed was changed out to address this issue). | Unable to perform root cause analysis as this type of repair troubleshooting eliminates opportunity to get this information
2803 | Battery Charging only to 87 % | closed | ebabaran | Battery | fixed | 2015-11-17 23:36 | 2015-11-17 | 51273 | 205 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue: Barrow - One of their GT batteries is only charging to 87 % (B1) | Send new battery via SWO-1564 | Additional analysis to be done with Ticket #2384 then updated here.
2802 | R Hip Getting Hot | closed | kmettler | Adjustments - Hip | fixed | 2015-11-17 19:12 | 2015-11-16 | 449000 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | Issue: Josh - Burke called and told me while walking a pt today they were getting an burning smell from the R Hip motor. One person said they saw smoke coming from the motor but the person I spoke with couldn‚Äôt verify this happened‚Ä¶ She only was able to smell some a burning from the R motor. This is only happening with one of their pt‚Äôs. 5‚Äô0 180 Lbs Pro+ Adapt Patient has difficult time getting to their left due to pain in L foot. PT‚Äôs have to physically push to hit targets. Today they have a stand time of 9 min- walk time of 1:36 and took 20 steps. Within this time they got the burning smell and possible smoke from motor. Note- While walking this person in 1st step they don‚Äôt have this issue, it‚Äôs only been while in Pro+. This issue only occurs with this pt in Pro+ and not with any other of their pts. Katie- Can we pull data from this unit? I‚Äôd like to be able to get back to them or Clinical with some advice on how to avoid this going forward. Also- If there was smoke coming from the motor is there a bigger concern or would we need the unit back? | This issue was cause by a user error that they were able to rectify. | The use likely caused the right hip to get hotter and particulate may have been in the motor that burned. Based on the device continuing to function, it is not damage to the actual motor.
2801 | Hip Flexion Issue | closed | kmettler | Adjustments - Hip | worksforme | 2015-11-17 18:55 | 2015-11-17 | 253041 | 225 | Miami Project | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | Issue: Melissa The PT at Miami Project sent this concern and information. They have a GT, delivered approx. around end of summer. -Friday November 13th at around 2:30pm EST I don't remember if the pt id was entered but he would be 0005 We had hip flexion set in the device to 7.3 degrees I believe. Got the subject in the device, stood and started to walk. This subject is normally one of our most advanced ambulators, with canes and barely needs any assistance. However this day he was unable to get forward. After double checking all other areas of the device itself and the programming, we decided to increase hip flexion by about a degree. Normally the adjustment is very slight and the subjects can barely feel a change but you could visibly see the hip flexion increase and the subject could feel a drastic difference. The rest of the walking session went normally. I'm thinking that even though we had the device set for 7.3 degrees, there was some sort of error where the hip flexion didn't translate in standing or walking until we increased it more? Have you heard of this happening before? Knee flexion is set to 4.5 Normally hip flexion is 7.3, but we increased to 8.3 when we were having the issue RWA is at 5.0 We were in Prostep with the crutch control engaged | Based on dialogue issues seems to be user related and due to improper use. | Improper customer use lead to results seen. No further action will be taken.
2800 | Ankle has more plantarflexion than usual | closed | kmettler | Foot Sole | fixed | 2015-11-17 14:33 | 2015-11-16 | 207755 | 133 | Uppsala - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Issue - Maria : Customer reported: ‚ÄúWe have a problem with our Ekso, yesterday when we used it we heard a crack (twice) from the left leg. And it seems after that as the left foot is more plantarflexed than the right one and when the patient is walking the front part of the foot almost touch the ground during the whole cycle. I have tried it on myself and experience the same thing. When we look at the Ekso without anyone using it, it seems that the whole foot is in more plantarflexion than the right one.‚Äù | replaced complete ankle assemblies including shank mounts per SWO-1563 | Both shank mounts broken due to design. See CO-11291 for additional information.
2796 | left knee cover broken | closed | kmettler | Other | fixed | 2015-11-16 11:56 | 2015-11-16 | 253041 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Left knee cover broken | replaced knee cover, swo 1551. Confirmed via SWO-1845 that fix is properly in place for this device | not clear, unit didn't fall, possibly happened while placing the patient into the ekso
2793 | Multiple #168 errors | closed | kmettler | Foot Sensor | fixed | 2015-11-13 15:21 | 2015-11-12 | 829722 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Determined Unknown | Issue - Maria Matteo reported that Ekso at Villa Beretta is throwing #168 with multiple patients in several occasions. They believe the sensors are not responding correctly. | Foot sensors sent to field via SWO-1565 and swapped out by distributor technicians. | Unable to get removed sensors back from distributor
2792 | System errors #1241 - #21 - #25 | closed | kmettler | System | fixed | 2015-11-13 15:08 | 2015-11-13 | 85294 | 198 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | Issue - Maria Matteo at EMAC reported that the Ekso at Negrar threw error #1241 as soon as it was powered on. The clinicians have changed the batteries and moved the legs of the Ekso and the error disappeared. When they used it with a patient, Ekso threw errors #21 and #25, and they were asked to perform a Leg Test. After the Leg test the Ekso was working again as normal. | Distributor was talked through performing a joint cal over the phone. No additional parts required and no SWO created. | Device functioned as designed error 21/25 can come up through regular usage.
2791 | Rotating right lower shank | closed | kmettler | Lower Leg | fixed | 2015-11-13 14:48 | 2015-11-13 | 157486 | 170 | FLM - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue - Maria Customer reported: We are having an issue with the right lower leg. I happens when we are using Ekso with our patients. The problem is that the lower leg twists when we are walking and we are not sure if it¬¥s safe to keep walking. | Customer added aditional washer | Adjustment screws can bottom out on the back wall of adjustment collar making it so that the holding torque is not fully achieved. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2790 | Errors #133 - #121- LK | closed | kmettler | Electrical Interconnect | fixed | 2015-11-12 16:39 | 2015-11-11 | 26135 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Training | Issue - Maria The customer reported that whilst they were walking a patient the Ekso threw errors #133 #121-LK. | For issue explained in this ticket device operated as designed. Visited customer to perform other work via SWO-2790 and verified this is not an issue | No problem, device operated as designed.
2789 | Left ankle screw broken | closed | kmettler | Ankle | fixed | 2015-11-12 13:06 | 2015-11-11 | 26135 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue - Maria Customer reported left ankle is moving slightly. Upon reviewing the videos sent to us by the customer, it seems that one of the screws might be broken. | Unit serviced in field via SWO-1554 | Screw previously used did not meet the spec originally designed for.
2788 | Upper DDL Break | closed | kmettler | Don-Doff Link | wontfix | 2015-11-12 00:05 | 2015-07-14 | 1100000 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | TRUE | 10 - Field Service & Loaner | Major | Investigation Needed | Design | ADDITIONAL PICTURES ON CAPA#2489 Issue: Josh - We have an major issue with our ekso # 1. I was walking with a user in ekso # 1 and we heard a pop and the left leg completely fractured at the hip joint detaching it from the device. There was a tech that was out last week and assessed the entire unit. However, the fracture of the device was slightly distal to the hip joint. I am going to send you a few pictures. I am going to call you to inform you of this because I feel this is a true emergency. I‚Äôve spoken with Good Shepherd and no one was hurt as a result of this failure. They had a PT spotting from the back, an aid helping with the walker and someone observing the session. After the failure they were able to sit the patient into a chair and get them out of the Ekso. Our current plan of action is to send Russ out to Good Shepherd so he can remove the broken part of the device and pack the rest of the unit into a crate. We‚Äôll be shipping Good Shepherd loaner unit #220. | Unit has been retired and will not repair. Unit will be torn down for further learning opportunities and will be engineering owned. | Design requirements did not properly anticipate customer use in terms of aggressive split stances during turning. RCA evaluated with CAPA #2489 | INV-0002
2784 | Screws rusting | closed | kmettler | Fasteners | fixed | 2015-11-11 14:22 | 2015-11-11 | 95204 | 128 | Dennis Veit - Germany | [6 empty] | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Specification Not Correct | Issue - Maria Dennis reported that some of the screws are starting to rust. He doesn't have spare ones. | Customer replaced screws. Rust proof screws implemented per CO-11608 | Screws were not specified to have rust proof coating | No investigation needed as complaint was identified as not affecting patient safety, the device function, labeling, packaging, or any of its specifications. (Evaluation done by Kelly Mettler, Regulatory Affairs Manager, on 4/24/2017)
2783 | Rusting batteries | closed | kmettler | Battery | fixed | 2015-11-11 14:15 | 2015-11-11 | 95204 | 128 | Dennis Veit - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Issue - Maria Dennis reported that all 4 batteries are starting to rust | The batteries were not replaced - This was a cosmetic issue. Michael cleaned the batteries off while he was there for other issues and determined new batteries did not need to be sent. | Known wear item - cosmetic issue
2780 | Broken hip cap | closed | kmettler | Actuators | fixed | 2015-11-09 15:36 | 2015-11-09 | 78534 | 193 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Major |  | Design | Issue - Maria Technomex reported: Private customer from Kudowa Zdr√≥j reported us today problem with screw in right hip. Please see attached pictures. I hope just replacement of broken part will make the trick. SN: 021502002 steps: 78534 Whole situation happened during walking with patient. Fortunately nothing bad happened to anyone. | Joint caps replaced to V2 style via SWO-1545. | Joint caps under designed for loading seen and can break from normal usage.
2779 | Right step issue | closed | kmettler | Foot Sensor | fixed | 2015-11-09 15:28 | 2015-11-09 | 248548 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue - Maria Customer reported that the right step doesn‚Äôt trigger properly with some patients, however when the PT tested the device there were no issues with it. I have asked the PT to perform a foot calibration to see if anything changes. | replaced foot sensors via SWO-01551 | Design of foot sensor can wear and become inoperable over time. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2778 | Hip width adjustment Left screw loose | closed | kmettler | Adjustments - Hip | fixed | 2015-11-09 15:23 | 2015-11-09 | 248548 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue - Maria The left screw to adjust the hip width is loose. | Replaced complete hip abductions on both sides per SWO-1551. | Parts evaluated on engineering ticket #2966
2776 | Error #133 and #121-RK | closed | kmettler | Motor | fixed | 2015-11-06 15:30 | 2015-11-06 | 425061 | 36 | Hobbs - UK | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Materials | Issue- Maria Customer called me reporting they were using Ekso with a patient when it threw errors #133 and #121-RK. The right knee motor is making a grinding sounds, however the knee moves smoothly. I have given the PT instructions on how to perform a joint test and I‚Äôve asked them to let me know the outcome. | MPF, 12/18/15: The source of the errors was traced back to the amp board. A new amp board was installed. Work done via SWO-1541. | Bad/damaged components on the Amp Board were the root cause of this issue. Unable to determine if certain operations performed on the device made this occur on the board.
2775 | Short left step | closed | kmettler | Foot Sensor | fixed | 2015-11-06 15:19 | 2015-11-06 | 90850 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | Issue - Maria Customer reported that when walking in First Step, the left step is shorter than the right one. I have asked Matteo at EMAC to guide them through a foot cal. | Ekso passed auto foot calibration. | Known wear item on the device.
2774 | R Hip screw loose | closed | kmettler | Actuators | fixed | 2015-11-06 09:36 | 2015-11-06 | 197642 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue - Maria Customer reported that the screw in the right hip is loose and some oil came out. Please advise if they should stop using the device. | Unit has been repaired, will get V2 joint caps as part of retrofit. Work perfromed via SWO-1542 | Original part had stress concentrations.
2773 | 121 Repeated | closed | kmettler | Electrical Interconnect | fixed | 2015-11-04 23:22 | 2015-11-04 | 368252 | 9 | Craig | Julie Hartman | jwaldie@craighospital.org | (303) 789-8000 | USA |  | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Design | Issue: Unit got 121 errors at 3 different times during a walk session. Code 12288 ‚Äì undervoltage This may be a resistance issue from that specific patient. We should continue to monitor. | MPF, 11/24/15: We removed the amp board on SWO-01553 and Ken attempted a repair, but upon re-installation errors persisted. We replaced the amp board. I bagged the old amp board and placed it in the manufacturing RCA box. I set Edmundo as owner. doghouse cables were updated per CO-11402 to prevent overheating described above. | Underspeced dog house cables burn up under adverse conditions. Potentially damaged AMP board in the process. Further evaluation of board and cables still required.
2769 | Faulty battery charger | closed | lmcneil | Battery - Charger | fixed | 2015-11-03 17:34 | 2015-11-03 | 221883 | 170 | FLM - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Determined Unknown | Issue - Maria FLM reported issue with one of their battery chargers. Charger LED wouldn't light up when battery was attached and wasn't charging battery any longer. | Sent via SWO-1529 | Charger not being returned for evaluation.
2763 | Error #121-LH | closed | kmettler | System | fixed | 2015-10-29 14:38 | 2015-10-29 | 393924 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue - Maria Customer reported Error #121-LH. It occurred with 2 incomplete patients whilst walking in Max and Adaptive. Matteo from EMAC will go to the customer on Monday to check the device. | Replaced 102286 cca via SWO-1531. | Determined that CCA 102286 was not functioning properly and replace this component.
2762 | Left shank brace rotating | closed | kmettler | Lower Leg | fixed | 2015-10-29 14:24 | 2015-10-29 | 125557 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Customer Use | Issue : Maria- Colin at Elite Gym reported that the left shank brace is rotating. See video attached. He advised that they had a spare left shank brace because they had the same issue with the right shank brace back in July, so he went ahead and replaced the part. He said that the Ekso is now fine. | Device was repaired, damaged parts replaced on SWO-01288 | Customer use anomaly, excessive force applied (>200 lb). brace did not separate but rotational stop was broken.
2761 | Battery Chargers | closed | kmettler | Battery - Charger | fixed | 2015-10-28 21:23 | 2015-10-28 | 137911 | 59 | Amanda Boxtel | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Wear | Issue: Able Bionics (Aspen) One of the chargers will not light up indicating that there is a charge and we are under the impression that it's not charging at all | Sent new battery chargers via SWO-1518 | Know wear item
2755 | R Upper DDL broken | closed | kmettler | Don-Doff Link | fixed | 2015-10-23 10:59 | 2015-10-23 | 84328 | 185 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Design | Issue - Maria "We have a very serious malfunction of the EKSO SN 021501002. The right hip has broke during operation. It has occurred approx. 45 minutes ago. During sit-to-stand transition, a loud cracking noise was being heard. After that, alarm sound popped up (error#25). Patient was in the ekso but was fine." | Hip actuator assemblies changes on both sides of unit via SWO-1523 | See CAPA #2489
2752 | Hip Joint Cap Detached | closed | kmettler | Actuators | fixed | 2015-10-22 22:47 | 2015-10-22 | 180232 | 220 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | Issue: While walking with a patient in the Ekso at 9am this morning we heard a sudden loud cracking sound. The right hip bolt became loose/detached and we needed to sit the patient immediately. I have attached a picture in this email of the loose right hip component. Please advise for how you would like to proceed with this issue. | MPF, 10/27/2015: Removed RH axle assembly components. Found damaged inside cap. Bagged parts and delivered to Aaron J. Unit fixed in house via SWO-1510 | Part not designed sufficiently for uneven loading
2751 | CI has intermittent errors | closed | kmettler | UI - Crutch Interface | fixed | 2015-10-22 22:38 | 2015-10-22 | 124302 | 134 | Chuck Slyngstad | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Use | Issue: Chuck was receiving intermittent errors with his CI | Sent Chuck a new CI via SWO-1500 | The CI was sat on
2745 | Batteries not holding a charge | closed | kmettler | Battery | fixed | 2015-10-19 18:28 | 2015-10-17 | 245248 | 42 | Montreal - Canada | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Issue: We are currently experience a problem with a pair of batteries. Apparently, we get an error message ‚ÄòMISS MATCH BATTERIES‚Äô when using them. The other pair of batteries works well. However, this problem reduces our capability to train individuals with SCI at the moment. We have interchanged all batteries to access the ABOUT submenu. Battries 1A, 2A and 2C are 98% charged according to Ekso. We can not access the submenu "ABOUT" when battery 1B is in the ekso (either on the L or R port). We then only get the "Miss Match" message without the possibility to stroll the menu. The menu tells us though that the port corresponding to battery 1B holds a battery with a low charge. | Sent new batteries via SWO-1492 | Additional analysis to be done with Ticket #2384 then updated here.
2743 | Right ankle clicking | closed | kmettler | Ankle | fixed | 2015-10-19 12:15 | 2015-10-19 | 165892 | 163 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | Issue - Maria Customer advised right ankle is clicking. | Monitored issue and determined this noise was within normal ranges and nothing was needed to be done to address issue. | Known sound that comes from this area of the Ekso.
2742 | Noise from torso | closed | kmettler | Torso | wontfix | 2015-10-19 12:11 | 2015-10-16 | 165892 | 163 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | Issue - Maria Customer advised ‚ÄúThere is a strange noise coming from the left and right sides of the device. Not yet preventing walking but it is not right and I wonder have you heard it before and can we fix it? It happens when the right heel strikes then the left hip/rear torso makes the noise. And, when the left heel strikes the right hip/torso makes the noise.‚Äù | Explained to customer these noises are commonly heard with devices. | Known sounds coming from the hips of the Ekso.
2741 | Broken left hip | closed | kmettler | Don-Doff Link | fixed | 2015-10-19 11:15 | 2015-10-16 | 652589 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | TRUE | 10 - Field Service & Loaner | Major | Investigation Needed | Design | Issue - Maria Left hip broken whilst patient was in Ekso. See pictures attached. There were 3 clinicians around when this happened, so they were able to safely sit the patient and nobody was hurt. MDR #7 | MPF, 12/21/15: Michael reports that he replaced both DDL assemblies with new per SWO-01520. | RCA done with CAPA #2489 | INV-0002
2740 | Foot plate split open | closed | kmettler | Foot Assy | fixed | 2015-10-16 02:22 | 2015-10-15 | 248000 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Issue: PT‚Äôs were walking patient into an elevator, during this process the toe guard hit the lip of the elevator floor breaking the guard and separating the foot plate as seen in the picture. We told the customer to stop using the unit and are working on an action plan to get this device back up and operational. | I replaced both foot plate assembles and test walked the unit per SWO-1480 | Customer walked patient onto elevator, tearing the foot plate open.
2738 | CI elastic band | closed | kmettler | UI - Crutch Interface | fixed | 2015-10-15 19:30 | 2015-10-15 | 355432 | 58 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue - Maria Matteo reporteed that the elastic band that holds the crutch interface is broken (not sure about the exact name) | SWO-01486 | Known Issue
2737 | Broken foot sensors | closed | kmettler | Foot Sole | fixed | 2015-10-15 19:15 | 2015-10-15 | 355432 | 58 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Issue - Maria Matteo @EMAC whilst visiting Pietra Ligure noticed that the foot sole was damaged and he replaced it. | Distributor used spare sensors in their stock that was shipped to them via SWO-1006. SWO was already closed in Omnify so note was added in comments section to account for sensors used. | Known Issue with Foot Sensors. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2735 | UI demagnetized | closed | kmettler | UI - Controller | fixed | 2015-10-15 16:33 | 2015-10-15 | 355432 | 58 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Issue - Maria Matteo at EMAC reported that the UI is demagnetized so is no longer holding to Ekso | Sent new UI via SWO-1485 | Magnets were possibly dislodged
2734 | Hard DDL | closed | kmettler | Don-Doff Link | fixed | 2015-10-15 13:12 | 2015-10-15 | 161553 | 133 | Uppsala - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue - Maria Customer reported : In sitting we still have some problem with the DDL release, sometimes it¬¥s hard to push the release and it¬¥s also harder to move the leg into abduction than on the other side. | Replaced DDL's in field via SWO-1490 | Was unable to further examine removed parts to perform further analysis.
2733 | Loose hip abduction | closed | kmettler | Adjustments - Hip | fixed | 2015-10-15 13:09 | 2015-10-15 | 161553 | 133 | Uppsala - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue - Maria Customer reported that the hip abduction is loose (on both sides) and although it doesn't happen every time, it should be checked. "The hip abduction value isn¬¥t consistence during a training session, could change one or two steps during a session." | Replaced hip abductions assemblies in field via SWO-1490 | Was unable to further examine removed parts to perform further analysis.
2732 | Batteries giving error messages | closed | ebabaran | Battery | fixed | 2015-10-14 21:13 | 2015-10-14 | 371719 | 19 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue: They received an error message of missing left side battery. They received a second error message saying charge left battery and reboot | Sent new batteries via SWO-1482 | Additional analysis to be done with Ticket #2384 then updated here.
2731 | Lumbar pad not inflating | closed | kmettler | Soft Goods | fixed | 2015-10-14 13:49 | 2015-10-14 | 115533 | 138 | Nottwil - Switzerland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Wear | Issue - Maria Customer advised - The valve on the lumbar pad went out of function, we cannot ‚Äúclose the valve‚Äù and therefore cannot adjust the amount of air in it. | Sent new lumbar pad via SWO-1471 | Known issue
2730 | Hip width adjustment - loose screws | closed | kmettler | Adjustments - Hip | fixed | 2015-10-14 13:44 | 2015-10-14 | 115533 | 138 | Nottwil - Switzerland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue - Maria Customer reported: The screw to fix the hip-width (left side) started ‚Äúloosening‚Äù. If we fix the width to 10, it is shifted to 7-8 during the training. We can continue with the sessions anyway, because our current patient has good sensibility in that area, but I don`t feel comfortable with it, regarding the risk of additional pressure. | Replaced hip abductions and tube via SWO-1489 | Dimension on upper tube was slightly under spec reducing the clamping force of the hip width lock.
2720 | Epoxy Failure- Ankle rotating | closed | kmettler | Ankle | fixed | 2015-10-08 17:52 | 2015-10-08 | 114652 | 188 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Process Not Available | Issue: Another issue has cropped up as well - the left ankle is freely rotating around the part where it is attached. Began while walking with a pt 175# and 6' tall. Ambulating on carpet in PT clinic. | MPF, 11-18-15: The damaged ankle was delivered to Eng. for inspection. The unit was repaired using a shank assembled according to the Rev. H of the MPI. Work was done via SWO-1488 | Not enough resolution available in the gluing process allowing for the possibility of this failure occuring
2716 | Ripped Foot Sensor | closed | kmettler | Foot Sensor | fixed | 2015-10-07 23:03 | 2015-10-07 | 18267 | 200 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | Customer: San Raffaele Unit: TBD Parts have been shipped to EMAC and they will be received tomorrow 10/8. Hi Josh, attacched the situation of San Raffaele foot sensor. Now they have to stop the therapy. Please can you give me an upgrade about the shipping time? Thanks | New foot sensors were sent via SWO-1465 | Design/wear issue. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2713 | Cracked Ankle | closed | kmettler | Ankle | fixed | 2015-10-07 22:34 | 2015-10-07 | 117998 | 163 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: Sandy - I just had a phone call from Mark Pollock: He was walking in his GT today and heard a crack. On examination, there is a crack at the right ankle (Mark wasn‚Äôt able to specify where or what it looked like as he himself obviously can‚Äôt see it) I have asked him to get someone to take a photo and email ASAP so we know what we are dealing with. I will pass this on as soon as I receive it. No info on step count/unit number etc. Mark‚Äôs contact number is: +353 87 223 3372 ‚Äì once we know what is wrong, please contact him directly. | We replaced both ankles with the latest version with SWO-01474. Solid PF wall version of shank mount is now used in production. | Part underdesigned for loads, test stand doesn't adequately test plantarflexion.
2712 | Ekso is walking Irregular | closed | ebabaran | Foot Sensor | fixed | 2015-10-07 22:25 | 2015-10-06 | 300000 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Materials | Issue: this is to inform you, that our ‚Äûfrequently repaired‚Äú Ekso shows again failure during therapy. Our Ekso was returned last week after last repair and this week again presents failure during therapy. The incident is present over RT hip, during walking pattern. Due to frequent walking disturbance during walking patients refuse undergo therapy on our Ekso. Once we execute Ekso therapy on loaner Ekso patients don't¬¥t complain about disturbance and the therapist can see and feel different walking pattern without asymmetry, with fluent movement and without small excessive movement. For the safety and therapeutic comfort of our patients we have to suspend Ekso therapy on our device, until all this issue is resolved and we will finally receive Functional Ekso without technical problems. I also talked to Jakub on the phone and he was telling me the R side of the Ekso seems very asymmetrical... The L side is very fluid and the R side is very rigid. We are bringing this device back to the German workshop for further troubleshooting. The customer will continue to use their loaner device. Katie- Do we get data from this device to help troubleshoot potential issues? | MPF, 1/18/16: See TRAC 2654, fixed per SWO-1654 | See TRAC 2654. Signal chain issue found with right flat flex cable led to intermittent signal from right foot sensor
2711 | Foot Sensors | closed | kmettler | Foot Sensor | fixed | 2015-10-05 16:20 | 2015-10-05 | 76702 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: The foot sensors were replaced to the latest revision | We replaced the sensors with the latest version. SWO-01431. | Heel sensor slid forward. This is a low risk problem. The defect will be corrected in the next version of the device.
2710 | Stepping Issue | closed | kmettler | Foot Sensor | fixed | 2015-10-02 18:31 | 2015-10-02 | 424793 | 36 | Hobbs - UK | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | Issue: Hobbs - I have a message from HOBBs rehabilitation in England saying they are having a stepping issue while walking. It sounds exactly like a bad foot sensor. This is the first call I have taken from a customer and I have not been trained in the proper methods of response and reporting so if someone can assist I would appreciate it. | Customer walked through an auto foot calibration and that fixed the issue. Verified that fix was implemented via SWO-1541. This was added to the comments section of the the SWo as it was already closed in omnify. | Values assigned to foot sensors can drift over time due to wear. This can be addressed by re-running the foot calibration
2709 | Foot Sensor Issue | closed | kmettler | Foot Sensor | fixed | 2015-10-02 17:55 | 2015-10-02 | 404386 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | Issue: Sandy - I just received the following email from BG Tuebingen: The long and the short of it is that they have a delayed response/shorter step on the right and occasionally have to do an exaggerated weight shift to the left to trigger a right step. This has been happening on and off for two weeks and seems to be irrespective of the patients they use it with. No error messages have been reported. It is Fri afternoon there now, and they might probably won‚Äôt be with/near Ekso now until Monday morning, but it is probably worth Michael giving them a call then to run them through a foot cal or other diagnostics. Contact details are in the forwarded Email and the best Person to speak to is either Margit or Annette. | We replaced the foot sensors with the latest version. Verified completion of this change via SWO-1680 | Old version foot sensors heel shifted.
2707 | Battery no longer charging | closed | kmettler | Battery | fixed | 2015-10-01 17:54 | 2015-09-30 | 58770 | 122 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Issue: One of the batteries is no longer charging. When plugged in to the charging docking station, the battery never registers and the light stays green even when the battery level is depleted. When plugging in the batteries to the Ekso GT, we get an error that reads "Battery Mismatch Error". Our second set of batteries works fine. | batteries sent via SWO-01458 | Additional analysis to be done with Ticket #2384 then updated here.
2703 | Battery charger - Bent Pins | closed | kmettler | Battery - Charger | fixed | 2015-09-30 15:25 | 2015-09-30 | 110104 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Customer Use | Issue - Maria Customer advised that one battery charger is not working. The batteries work fine when charged with the second charger. See picture attached. Some pins are bent and some of them are missing. | New charger sent via SWO-1457 | Some of the pins were bent while trying to charge the batteries.
2701 | Left Hip Adjustment | closed | kmettler | Adjustments - Hip | worksforme | 2015-09-29 17:02 | 2015-09-29 | 229482 | 165 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Specification Not Correct | Issue: Josh - Our rehab aides have been noticing that the (L) hip width adjustments and the ability to turn the tool has been getting progressively harder. We did mention it to the tech/engineer who was here a bit ago for the other adjustments, but it is getting worse. Is there anything we can do to correct this? | From Mark Frye: I visited the site with a complete LT hip abduction assembly. Replacement of complete assembly did not improve ease of adjustment. Removed key from both assemblies and noted that adjustment stiffness remained the same with either unit, with increasing stiffness as unit approaches smaller settings. Replaced entire control bar assembly and key within existing housing assembly to address excess wear in existing control bar assembly. Torque required to move assembly from set point 5 to set point 3 is 20 in/lbs. RT assembly torque required to move from 5 to 3 is 15 in/lbs. PT's are using red and black palm driver to make adjustment. I suggested that they get a blue handled tee-wrench like the one I use which can also be used to rap on the adjustment screw. One PT mentioned that she had had a tee wrench but couldn't find it in the fit kit. I bagged the control bar assembly and key and placed it in the manufacturing RCA box. Remaining issues: * I understand that the adjustment stiffness is in-spec when 30 in/lbs or less. But, PT's report that mechanism is too stiff to adjust using red handle driver when it is more than 15 in/lbs. * Technicians are using more robust tee handle wrenches to assemble/service/test adjustment while PT's are using palm drivers * Specification of allowable adjustment torque does not seem to align with actual use case All applicable work was done per SWO-1462. | May need to reevaluate as to what is acceptable in terms of torque required to use the hip width adjustment or need to evaluate our training to show what is an acceptable level of torque.
2697 | Right step not triggering - Foot sensors issue | closed | kmettler | Foot Sensor | fixed | 2015-09-24 13:12 | 2015-09-22 | 186977 | 156 | Kerpape - France | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Maria - Issue: Customer reported that the right step is not triggering properly. I have given them instructions on how to perform a foot calibration. They have tried this twice, once it gave #170 and the second #171. I have asked them to repeat the test ensuring they firstly walk in Free Mode for 100 steps and see if this time something changes. | Replaced foot sensors, SWO 1431 | foot sensors faulty. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2696 | Hip joint caps broken | closed | kmettler | Actuators | fixed | 2015-09-24 12:38 | 2015-09-21 | 129204 | 57 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major |  | Design | Issue - Customer reported: Broken hip whilst patient was walking. The patient is fine. They heard a loud crack and managed to sit down the patient safely. | New joint caps installed via SWO-1522 | Previous version of joint caps can break during regular usage. Latest fix not yet implemented on this device.
2695 | Error 27-LH | closed | kmettler | System | fixed | 2015-09-24 09:24 | 2015-09-23 | 435747 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Design | Issue - Customer advised: ‚ÄúWhile I was installing the patient into the device and putting the settings into the software, the ekso gave me (two times/different sessions) the message ‚Äòsensor error LH-27‚Äô. This messages appears after the ‚ÄòHOLD FOR FREE‚Äô in the safe mode (moving joints). After rebooting the ekso (switch off/switch on) the message didn‚Äôt appear anymore. ‚Äú ‚ÄúThe error occurred working with two different patients. Sometimes I notice that the step length of the left foot is a bit taller, but this is not standard.‚Äù | Complete - Maria - to guide customer through joint and foot cal. Michael went on site to verify that unit was properly functioning. No additional parts required and no SWO created. | Unit functioned as designed. Series 20 errors can be thrown during regular usage.
2690 | Error #121 RK | closed | kmettler | System | fixed | 2015-09-21 23:20 | 2015-09-21 | 230648 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue: I just wanted to touch base with you regarding an error message we received while training today. It read error #121-LK that occurred when the patient mis-stepped with her R leg (short step) and then the L leg (free) kicked back with knee flexion. The device then went into safe mode with some amount of knee flexion in both legs (about 45 degrees). I was safely able to sit her down, but just wanted to loop you in to see if there is anything that needs to be done or if this is an issue given the other error that occurred recently with the device locking with significant knee flexion as well. I have cc'ed Raquel and Susan, other PTs in the lab on this email as I will be out of the office starting tomorrow. Katie- can we pull their data to determine if their device is having issues? | Melissa talked to customer and talked to them about changing the settings on the device to prevent this from occurring again. | Patient was able to cause the device to be put into safemode due to resistance provided during his walk.
2689 | Foot sensor Peeling | closed | kmettler | Foot Sensor | fixed | 2015-09-21 23:09 | 2015-09-21 | 551183 | 12 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Issue: Peeling Foot Sensors, Customer is down | Sent foot sensor via SWO-1426 http://exo-omnify/omnify5/Apps/Desktop/Quality/Object.aspx?id=4559&pg=2 | Known wear part. This is a low priority, low risk problem. The defect will be corrected in the next version of the device
2687 | UI light not working | closed | kmettler | UI - Controller | fixed | 2015-09-18 14:49 | 2015-09-18 | 733750 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Materials | Issue - Maria Matteo at EMAC reported: The light of the UI is not working | Sent new UI via SWO-1419 | Known issue.
2686 | Right knee not flexing | closed | kmettler | Actuators | fixed | 2015-09-18 13:47 | 2015-09-18 | 727466 | 147 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 40 10 350 w. 38 | Poland | no injury | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Major |  | Materials | Issue - Maria Issue: Jakub at Technomex reported that this morning before starting the first session with a patient they noticed that in free legs the right knee didn‚Äôt flex. The left one works fine. When the Ekso is switched off, the knee flexes if pressing the knee release button. The device is down. I discussed the issue with Michael and Greg and they advised it might be an issue with the ‚Äútorso to hip cable‚Äù which could be loose or there could be an issue with the cable itself. Jakub will be back in the clinic on Monday so he will look for any damage to the cable. | SWO-01450 was resolved by Greg | Damages knee external cable had to be replaced
2681 | Error 27-LK | closed | kmettler | Actuators | fixed | 2015-09-17 10:24 | 2015-09-17 | 44207 | 112 | Heliomare - Netherlands | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue - Maria Customer reported: Ekso threw error 27-LK twice with the same patient. Both times they were walking in pro-step. The first time happened when the PT adjusted the fixed setting from 80% to 75%. They sat the patient and rebooted the Ekso and carried on walking but the error reoccurred a second time. Height and weight are 1.76m and 80kg. | Tech went to customer site and calibrated leg and foot sensors. No SWO created as no parts were required to perform this service. Walked 750 steps Replaced 9v batteries Replaced ankle screws Done 100 points check. The Ekso is working. | Requirement for leg calibration is safety that functioned as designed.
2679 | Fracture Inside Joint Cap | closed | kmettler | Actuators | fixed | 2015-09-16 09:58 | 2015-09-16 | 292884 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Major |  | Design | Issue: Maria Customer reported: During therapy Rt hip screw went out, see attached photos, we were able to sit patient safely on the chair. I don¬¥t know what to say. Can one of your engineer explain how this could happen after recent complete check up? Our Ekso went thru technical check up recently and technical inspection from your engineer at Kladruby last week. Besides this we need to have Ekso functional for life TV shov on Nov 9. | We replaced both inside and outside joint cap and the bolt. We put the unit back together using MPI 102205. Updated V2 joint caps installed per SWO-1654. | Down rev version of joint cap known to fail during use.
2674 | Foot Sensor Issue | closed | kmettler | Foot Sensor | fixed | 2015-09-15 22:28 | 2015-09-11 | 60292 | 135 | GF Strong - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue: Melissa - I just received the following note from the current lead PT at GF Strong. Melissa, the Ekso was doing a bit of twerking when we used it about an hour ago and I don‚Äôt have time to do a foot sensor calibration. Can you warn your colleague? Mike Glover is leading an ITW at GF Strong with this Ekso starting promptly at 7:30 am PST on Monday Sept 14. Is there anything we can look at remotely to determine what he might be seeing, and be proactive so that Mike‚Äôs Day 1 of ITW isn‚Äôt compromised? | Sensors on device seemed to function properly but likely not within usable range for customer patients. Foot sensors changed i filed via SWO-1407 | Passable foot sensors can work not as well given customer use issue. Placing sensors with more sensitivity can absolve this issue
2673 | Ekso Stall | closed | kmettler | Foot Sensor | fixed | 2015-09-15 21:53 | 2015-09-15 | 105955 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Unassigned |  | Wear | Issue: Customer called and has stated they are stalling while trying to initiate a left step. The pt is hitting their targets but the beep is delayed causing the leg to stall before it will step. Does anyone see what the issue may be? I‚Äôve walked them through a foot calibration‚Ä¶ it passed. Please let me know what you think so I can get back to the customer. Josh | Confluence received new foot sensor via SWO-1423 | Known Issue
2672 | Ankle Epoxy Bond Failure | closed | kmettler | Ankle | fixed | 2015-09-15 21:45 | 2015-09-15 | 130000 | 173 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | TRUE | 9 - Loaner | Major | No investigation needed - see Justification | Process Not Clear/Complete | Issue: While walking today, the PT was turning a corner with a patient and noticed the foot was internally rotated. They sat the patient and called the red line‚Ä¶ Chad spoke to the PT, after running them through some testing it was determined they broke the epoxy bond on their R ankle. This unit was inspected last March and passed our test. We are sending them a loaner and will be getting their Ekso back by Thursday morning. | Shank/housing replaced per SWO-01140. Shank with broken bond placed in Manufacturing RCA box. | Several contributing factors such as inadequate surface preparation and insufficient epoxy applied led to an incomplete bond upon assembly. Bond then failed under cyclical loading. Additional root cause will be performed in CAPA #2278 Pictures located in M:\Projects\Ongoing Engineering\Shank Epoxy bond failure\Trac #2672 - Unit 173\Pictures | previous investigation performed ‚Äì see #2723
2662 | Error #137 | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-09-13 07:02 | 2015-09-09 | 111969 | 43 | Sunnaas - Norway | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Issue - Maria Customer reported having issues during walking on Wednesday 9th: ‚ÄúTwo times we got an error with error code 137(?): Ekso must sit. Reboot required. After sit down we turned the Ekso off, and then on again and was then allowed to start walking again. This was in ProStp with max assistance. With the next patient walking in ProStp, bilateral adaptive there was some very high/long steps with the right leg. The display showed in the bottom: L66 LIFT>R95 -No error code The first problem happened again during training today error code 137. This time in ProStp, Adaptive.‚Äù | took off the left brake hub, cleaned and reinstalled it. Fixed via SWO-1410 | Accumulated usage can allow for loss of holding force on break. Maintenance of cleaning and regapping can help to address this.
2655 | Axle for ankle shank mount is missing flange | closed | kmettler | Ankle | fixed | 2015-09-09 15:52 | 2015-09-09 | 148745 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue - Maria Axle for ankle shank mount is missing flange. Please see pictures attached and apologies for the lack of information. Jakub at Technomex wasn‚Äôt able to provide more details, he will give us more information tomorrow. | Sent replacement ankle assembly per SWO-1405 | Stress concentration on flange.
2654 | Foot sensors issue | closed | kmettler | Foot Sensor | fixed | 2015-09-09 13:36 | 2015-09-09 | 286280 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Materials | Issue - Maria Customer reported foot sensors issue. When patient is shifting on the right the left step is not triggering. This occurred today with 2 patients. | MPF, 1/18/16: I replaced the flex cable in the RT knee which was cutting out the foot sensor signal when in certain positions per SWO-1654 | Signal chain issue found with flat flex cable
2651 | Sit to Stand | closed | kmettler | Customer Use | fixed | 2015-09-08 23:09 | 2015-09-08 | 213000 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Hi Josh, see the video attach of the feet popping up when it goes into the "Begin to Sit" mode. (I Believe they mean begin to stand). Please advise. Thank you. Josh | Replaced foot sensors per SWO-1406 | No unique cause identified as of yet. Will continue to monitor in the future. See engineering ticket #2954 for any technical updates.
2648 | Broken foot sensors | closed | ebabaran | Foot Sensor | fixed | 2015-09-08 13:53 | 2015-09-03 | 125802 | 57 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | Issue - Maria Customer advised: I contact you because we keep having problems with the foot sensor in both feet. Ekso is down. | New foot sensors installed per SWO-01394 | Wear item and most of the time does not stop use of device. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2644 | Foot Sensor Issue | closed | kmettler | Foot Sensor | fixed | 2015-09-08 09:11 | 2015-09-08 | 22912 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Wear | Issue - Maria Montecatone have sent the following videos to EMAC. https://www.dropbox.com/sh/iv35vs6382rxo3m/AABHIcoVSfhG1oNvvWumWEeCa?dl=0 They have reported a noise and an abnormal movement of the left knee when the right leg is taking the step. The knee reaches complete extension but it clicks in the final phase and it seems there is something preventing the knee from extending smoothly. However it doesn‚Äôt happen with all patients. Could you please review the videos and advise. | Foot Sensors replaced via SWO-1662 | known wear item. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2642 | Error #21-LK | closed | kmettler | Actuators | fixed | 2015-09-07 12:54 | 2015-07-24 | 22912 | 197 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Determined Unknown | Issue - Maria Issue: Error #21-LK Customer reported error #21 multiple times but informed EMAC only now. 24/July patient X error #21LK twice at the end of the session 29/July patient Y error #21LK at the beginning of the session 30/July patient X error #21LK during initial set up 30/July patient Z error #21LK at the end of the session 28/August patient Y error #21-LK after 30 minutes into session | When servicing unit via SWO-1554 was able to verify that this issue was no longer present. | Unable to reproduce issue while working on device.
2639 | Ankle axle placed backwards | closed | kmettler | Ankle | fixed | 2015-09-03 10:17 | 2015-09-01 | 888700 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Customer Training | Issue - Maria Technician at Technomex has placed the ankle axle backwards. | Ekso Field Service Technician arrived on site to verify that axle was installed correctly. Extra parts were allocated per SWO-1357. | Technician at Technomex accidentally replaced axle in backwards while changing a shank mount captured in a previous trac.
2638 | Ankle adjustment stuck | closed | kmettler | Adjustments - Ankle | fixed | 2015-09-03 10:10 | 2015-09-01 | 715461 | 147 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue - Maria EKSO # 147 still has problem with right ankle. Right ankle stuck and screw that is in the back can not be moved, so therapists can not adjust dorsi and plantar flexion. | New slider and latch assembly was replaced in affected (R) ankle. Replaced with SWO-01342 | Retaining ring (102370) in ankle adjustment carriage was broken, due to a less than ideal mating surface against a brass insert. Chamfered edges on insert exerted a stress against the edge of the ring (causing breakage), rather than the face of the ring. Re-design is in process. 103428, 102321, 103028, 103292, 103199, 102323, 102370, 103029, 102325 (x2)replaced.
2637 | Loaner @ RegiBase: Errors 21-LH & 27-LH | closed | kmettler | Actuators | fixed | 2015-09-03 09:29 | 2015-09-02 | 11770 | 146 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Issue- Maria Ekso is throwing errors #21-LH & # 27-LH Even after rebooting it, the Ekso doesn‚Äôt work. | Torso test and re-calibration fixed the issue. No W/O associated with this action, as parts were not issued. | Customer was moving unit while turned on; resulting in sensors to "lock" the machine as a safety.
2632 | Power Down | closed | kmettler | System | fixed | 2015-09-01 20:30 | 2015-09-01 | 364000 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Design | Issue: PT called today and said on Friday they were walking in the Ekso and it powered off on them. PT said they‚Äôre hands were nowhere near the power button‚Ä¶ They were able to sit the patient and continue the session. They have had no issue‚Äôs before or since this happened, it was a one-time occurrence. If we can pull data this happened on 8/28 and was pt ID 0007. Is there any way we can determine why this happened through data? - No | Monitor the device for additional events, no further action required. | we can see from the data that the event happened and it can be can determined that it was not caused by software, as there was no crash log or data written before safe mode. Other than this, we cannot tell why it happened. We are not able to reproduce this event. Root cause speculation is ESD though this cannot be confirmed at this time.
2631 | 137 LK Errors | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-09-01 19:19 | 2015-09-01 | 283420 | 115 | NRH - Med Star | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Available | Issue: PT called today, told me they got 3 137 errors w/ a patient today. Patient was around 120lbs 5'5. They also received a 137 a few weeks back but didn't report it to us. The device has worked fine since, today is the first time it's happened again. | The brake's and motor shaft's mating components were properly cleaned and the brake was reinstalled and reburnished according to current manufacturing process. This work was completed per SWO-1663 | There was a greasy substance between the two flats of the brake which prevented proper initial installation of the brake. CO-11400 identifies the proper process for cleaning and installation.
2627 | Hip blocked in flexed position | closed | kmettler | Adjustments - Hip | fixed | 2015-09-01 10:29 | 2015-09-01 | 331939 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Wear | Issue - Maria Email from the lead physio: I was adjusting the settings for the hip width and the hip abduction while the Ekso was switched off. When I putted him back on the chair the left hip stayed in a blocked flexed position (see pics). Switching on makes no difference. We can‚Äôt move the hip joint. There were no issues when walking patients before‚Ä¶ | Flexure was inspected, linear rail was inspected, joint was inspected, pins were inspected, and bearings were inspected. All appeared in functional order. Movement was restored by installing 2 new floorstock hardware bolts (C100313) to specifications. No SWO was created, as no new parts were installed. Replaced the screws that are loosening (M3x6mm) with a new shorter version (M3x5mm) on CO-10975 in Feb 2014 | Dissassembly revealed that the 2 bolts that attach the flexure to the ball screw were loose, just enough "play" was present to occasionally work without end user noticing. Once the flexure became bound, the joint would no longer move. The 6mm long screws could bottom out in the carriage before they effectively clamped the flexure down
2626 | Left foot rotating | closed | kmettler | Ankle | fixed | 2015-09-01 09:05 | 2015-08-31 | 284396 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Process Not Clear/Complete | Issue - Maria Issue: As per pictures attached, the left foot rotates excessively when patient is walking and the feet clip together interfering with the swing phase. This happens mainly when walking patients with spasticity, 10 minutes into the session, however it doesn‚Äôt happen in right foot. I asked the customer to stop using the device. They have loaner #146. | Unit was in house per SWO-1654 and it was verified that the epoxy bond on the left shankle mount/ankle is no longer an issue. | Previous process allowed for variability in final results. New process is validated on a per batch basis and provide the extra clarity to remove these variables.
2625 | Right Leg Adjustment | closed | kmettler | Adjustments - Leg | fixed | 2015-08-28 23:44 | 2015-08-20 | 15841 | 222 | Kennedy Krieger Institute | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: Shantel - I received this email today from KKI. I want to make sure all parties are aware they are still having challenges with Right Leg adjustments. | We replaced the shank collar with SWO-01329 | The collar was rubbing against the shaft causing undue friction.
2624 | Loaner @ RIM: Foot Cal Fail & Pot Encoder Mismatch | closed | kmettler | Foot Sensor | fixed | 2015-08-28 23:25 | 2015-08-26 | 522006 | 12 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: ‚ÄúSo I was able to video tape the noise that we are noticing with the left knee during gait. I think it is a double click and the knee doesn't get full extension. I hope you can see and hear it on the video. I think I noticed it more towards the end of the short video. Just to give you a quick background, the patient's dx is cauda equina and she is incomplete and is able to walk outside of Ekso with 2 small base quad canes without assistance. She has very minimal tone. It happens with multiple clients to varying degrees. See the note from Tracy Rogers..""The error on 8/25 was abort step, stand still, then double arrow to return to standing menu. This error occurred during step phase left both times. It happened during first two trials (once each)"" This was with a client that is 6'5" with some spasticity that is decresed with a wave plate application beforehand. He is incomplete in variable assist. He is NOT in the video but just giving you written info about him (ANDY).‚Äù Chad and I walked them through a leg cal and it failed both times they attempted it. They received Errors 25 and 171. We will be sending a FST out next Tuesday. | When debugging in house it was found that the flat flex cable signal was intermittent. Replaced cable via SWO-1491. | Intermittent connection of flat due to constant move and wear seen on effected component.
2617 | Right rotating shank | closed | lmcneil | Adjustments - Leg | fixed | 2015-08-26 14:41 | 2015-08-26 | 89222 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | Issue - Maria Customer reported the right shank is rotating when patients walk in Ekso. Sometimes it happens with patients without spasticity. They have already added 2 washers to the bolts, but the issue is now reoccurring and this is causing the PT to stop the session and realign the shank. This also happened back in July and a field technician went on site to fix it. | cleaned tubes, added additional washer to set screws, replaced set screws in field per SWO-1373 | Adjustment screws can bottom out on the back wall of adjustment collar making it so that the holding torque is not fully achieved
2616 | Broken foot plate's screw | closed | kmettler | Foot Assy | fixed | 2015-08-26 14:19 | 2015-08-25 | 469375 | 54 | Prosperius - Italy | [5 empty] | Same as contact listed | ekso 1.1 | TRUE | 8 - Field Service | Major | Investigation Needed | Design | Issue: Maria - : Screw in the foot plate broke. This occurred when patient was in the Ekso, however there was no injury or loss of balance. See pictures in the dropbox link. https://www.dropbox.com/sh/m3cricj98w8xp2n/AAA-LoxMjFr3Yo9y1tGnEeSRa?dl=0 | We replaced the left ankle housing to SWO-01496. | See CAPA #2856 for complete RCA | INV-0001
2613 | Noises from Hip Joint | closed | kmettler | Don-Doff Link | fixed | 2015-08-20 20:07 | 2015-08-20 | 21222 | 190 | Mt. Sinai | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Materials | Issue: Today I received a text from one of the PT's at mount Sinai. He is concerned about noises coming from the hip joint of their Ekso GT. I have included the videos he sent me. He did mention this person has some hip tightness, but nothing extreme. If this noise sounds familiar to anyone please advise what may be causing this issue, if it's serious and how we can go about potentially fixing. | The loud popping was the DDL band. They were able to re-attach the band. | The loud popping was the DDL Band
2612 | Error 1241- Joint Board Offline | closed | rrea | Actuators | fixed | 2015-08-20 01:30 | 2015-08-19 | 265730 | 116 | Good Shepherd - IP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Process Not Clear/Complete | Issue: Good Shepherd called and they are having multiple motor failure bit set errors, L and R Knee (Error 121.) They are also having errors #1057, #1065 and #1241. We will be bringing this Unit back to HQ and will be providing them with a loaner. Update: The error log for this device shows 1241, 1063, 1061, 1240, and 1052. There are no 121 errors in the log. The errors began with the RK joint board never coming online on boot. | Replaced damaged RT knee joint board per SWO-1387. Knee cable had not been installed per manufacturing procedure, cable wires caught under separation plate crushed inductor on board. Need to link this TRAC to existing TRAC for this known issue. | There were several mfg issues with the knee boards that were corrected on Rev G6r on 2/1/15. This release also included a powerup test to ensure functionality before mfg. CO-11257 Inductor was also crushed and this was due to cable routing incorrectly (see resolution description)
2609 | Issue locking left leg in extension in R Aff Mode | closed | kmettler | System | fixed | 2015-08-19 14:37 | 2015-08-19 | 121633 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 0 - Not Required | Minor |  | Wear | Issue - Maria Customer advised: Yesterday the Ekso had a problem locking the left leg in extension when it was in right affected mode, so the left leg was free. We had high stance support selected and the Ekso gave way when standing on the left which almost resulted in a fall. We reproduced the issue in our lunch hour and I have filmed it and attached the video. I have a video for both legs as it seems to be unlocking on both legs, however more pronounced on the left side. We will continue to use the Ekso in bilateral mode until we have solved this issue. | Customer was run through a foot calibration to help absolve this issue. No SWO created as no additional parts needed to address this issue. | Loss of sensitivity in foot sensor on the left foot contributed to the left leg not locking in support when the uneffected leg.
2608 | Squeaky noise | closed | kmettler | Adjustments - Hip | fixed | 2015-08-19 12:18 | 2015-08-19 | 76078 | 46 | Guttmann - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single |  |  | Wear | Issue - Maria Customer reported: ‚ÄúSince the last two weeks aprox., there is a strange noise (e.g. plastic squeak). I send you a video in this link. http://1drv.ms/1NCXIhq. Is this important?‚Äù | New screws and loctite were replaced, restoring normal function to the machine. No W/O associated with this action, as parts were not issued. | Gusset cover has small features which wear over time, causing the screws to loosen due to increased clearance.
2606 | Bent Pins | closed | kmettler | System | fixed | 2015-08-18 17:23 | 2015-08-18 | 124000 | 161 | Glenrose - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Issue: See attached picture. Don't know how this happened. Only noticed cause battery wouldn't lock in. Please advise. Action Plan: Sending a FST to service the unit in the field tomorrow 8/19. | Replaced LT battery board per SWO-01345 | Customer accidentally docked the battery incorrectly and damaged the battery pins. Placed damaged battery board in manufacture Eng. RCA box.
2605 | Shark fin rubber band out of position | closed | kmettler | Soft Goods - Upper Leg / Pelvis | fixed | 2015-08-18 15:25 | 2015-08-18 | 109464 | 43 | Sunnaas - Norway |  |  |  |  | no injury | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Customer Use | Issue: Maria - Customer reported: We have noticed that one of the small rubber bands in the hip is out of its position (I cannot remember if this is one the right or the left side..) | replaced rubber bands - no SWO needed. Floorstock replaceable item. | The rubber band somehow moved out of position
2603 | Noise from Right Hip | closed | kmettler | Actuators | fixed | 2015-08-18 15:11 | 2015-08-18 | 109151 | 43 | Sunnaas - Norway | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 3 - Next Service | Minor |  | Design | Issue - Maria We have also noticed some clicking in the right hip. I hope the following information may help. (I am sorry that we do not have a video of these sounds. It is not every time these sounds appears). The right hip.. every time the right step is taken (this may occur with the short candidates as well) | replaced and torqued gusset cover screws. This work was completed while performing work outline by SWO-1410. This has been added to the comments sections of this SWO. | Original design called out a lower torque spec and medium loctite which can back out during regular usage.
2602 | Deep clicking sound from back of Ekso | closed | kmettler | System | fixed | 2015-08-18 15:09 | 2015-08-18 | 109151 | 43 | Sunnaas - Norway | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 3 - Next Service | Minor |  | Design | Issue - Maria Customer reported: Deep clicking sound from back of the Ekso - The deep clicking sound in the back may occur during training with the tall candidates. It appears when the right step is being taken. I can hear the sound is more on the right side of the back. | Replace gusset cover screws with and used newly defined loctite and torque spec. No SWO created as no stock parts required to complete this service. | Gusset cover screws that were replaced could previously back out with originally specified loctite and torque spec
2601 | Loose left knee joint | closed | kmettler | Actuators | fixed | 2015-08-17 14:39 | 2015-08-17 | 298640 | 31 | Ekso Bionics EU | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Process Not Followed | Issue - Maria Unfortunately there is a new problem with the loaner device. After using it until Friday, we used it once today to let rest our device. As you can see in the attached video and photo there is a problem on knee joint level at the left side. There is the possibility to move the lower leg with reference to the upper leg inside and outside. | New joint caps were installed,(B100162-001, B100361-001) and other associated parts. Remedied per SWO-01344. | The bolt and joint cap had become loosened, and the joint cap fell out. This led to the rest of the captive parts becoming loose, and "walking" out of position.
2598 | Error 41, persistent after new foot sensor change and calibration | closed | duanee | Foot Sensor | fixed | 2015-08-17 10:31 | 2015-08-16 | 245125 | 125 | Technomex - Poland | Jakub Maliszewski | bts@technomex.pl | +48 32 40 10 350 w. 38 | Poland | No injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor | No investigation needed - see Justification | Materials | Customer replaced foot sensor, and calibrated the foot sensors. Error 41 returned shortly after calibration. | Foot sensors replaced per SWO-02063. | Foot sensors replaced would not recalibrate without issue. | INV-0006
2597 | L. Hip has a cracking / pop sound | closed | duanee | Actuators | fixed | 2015-08-17 10:24 | 2015-08-16 | 245125 | 125 | Technomex - Poland | Krzysztof Cygon | kcygon@technomex.pl | +48 32 231 66 60 | Poland | No injury | (same as contact listed) | ekso 1.2 | FALSE | 8 - Field Service | Minor | Investigation Needed | Design | Reported during an on site interview @ Technomex: During walking, an occasional loud popping sound comes from the left hip. No hesitation, just noise. | New hip actuator housings have been installed. These arrived at customer's site partially built from USHQ, with everything needed except for the motor and the electronic PCB. Resolved per SWO-02835. | During deflection of mechanics at times of high stress when patient is learning to gait train, there is considerably more "sway" to the patients step when new. The housing has a machined step in the design that becomes an interference point with the DDL when the right conditions (new patient, more "sway") are met. Age of unit, may or may not be a determining factor of this issue. | (see Root Cause Analysis) No further action is necessary. (D. Edmonds ‚Äì 12 Oct 2017)
2595 | Crack in the Hip Housing | closed | kmettler | Actuators | fixed | 2015-08-14 04:57 | 2015-08-07 | 899406 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Materials | The second set show that cracks have started on the inside hip housings but not on the upper DDL link of their higher step count unit #129. The third set shows that cracks have not yet started anywhere on their lower step count unit #125. We will look to schedule a site visit next week to both fix their units and retrieve damaged parts. | replaced hip housing, SWO 1324 & SWO-01340 | Additional analysis done with CAPA #2489
2594 | Toe Guard was lost | closed | kmettler | Foot Assy | fixed | 2015-08-14 04:50 | 2015-08-07 | 222254 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Wear | This unit has completely lost toe guard. I‚Äôll have to replace it with new one. We will look to schedule a site visit next week to both fix their units and retrieve damaged parts. | Toe guards were shipped to the distributor via SWO-01327 | known wear item
2591 | Right toe guard broken | closed | kmettler | Foot Assy | fixed | 2015-08-13 14:40 | 2015-08-13 | 83293 | 170 | FLM - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue:Maria - Dennis reported: I'm at FLM in Madrid and have just seen, that the tip of the sole of the right leg is defect on their Ekso (see attached pix). | toe guards replaced by Adrian via SWO-1337 | Known Issue. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2590 | Short Left Step | closed | kmettler | Foot Sensor | fixed | 2015-08-13 14:10 | 2015-08-13 | 319382 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | Issue - Maria: we have a new Problem with the EKSO device. The left leg makes approximately every fifth step a short one. We have calibrated the device but the problem still exists. | Talked customer through subsequent foot calibration and that resolved the issue. No SWO was required. | Foot sensor values can drift with regular usage.
2589 | Thigh adjusmtent rotating after tightening | closed | kmettler | Adjustments - Leg | fixed | 2015-08-12 08:38 | 2015-08-11 | 384912 | 31 | Ekso Bionics EU | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service |  |  | Solution Not Implemented | Issue: Maria - PT reported: "Yesterday afternoon when we walked one of the last patients we experienced a problem with the tightening of the upper screws on the left side (loan device), see the attachment." | Clean tubes with alcohol. No SWO was written to address this as no parts were provided to perform this service. | Debris on the adjustment tubes can significantly reduce the holding force of this adjustment. Periodic cleaning of this area is needed to ensure proper functionality of device. Maintenance Checklist (104372) identifies this work to be done on a monthly basis.
2585 | LK motor is heating up during walking | closed | kmettler | Actuators | fixed | 2015-08-11 18:01 | 2015-08-11 | 34230 | 124 | Renown | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Determined Unknown | Issue: while walking with patients across the board the L Knee seems to heat up very quickly. No errors are thrown but Edmundo walked them into the back screens and the LK was much hotter than the other 3 joints. This particular patient was walking for 12 minutes and they could smell the Knee motor, this is when they called tech support. | Re-ran Elmo calibration for LK joint during field service visit. Work was done per SWO-1664. | If there is an incorrect synchronization between the commutation signal and motor position the overall control of the motor will be suboptimal and cause an inefficient state which in turn causes the motor to heat up. It is unknown as to what would cause this incorrect synchronization.
2580 | Loaner @ Good Shepherd: Ankle Knob Broken | closed | kmettler | Adjustments - Ankle | fixed | 2015-08-10 16:48 | 2015-07-14 | 47154 | 220 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: Josh - please see attached Picture. This occurred today... I have an email into their team to get more information regarding the failure. | We cleaned out the old loctite from the threads and reinstalled the ankle knob using a new bolt and loctite to MPI specifications. Verified completion of work when unit was returned in house via SWO-1510 | The loctite bond was broken and the screw backed out.
2579 | Multiple Sensor Errors | closed | kmettler | Foot Sensor | fixed | 2015-08-10 16:22 | 2015-08-10 | 118329 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: Sandy - Please see below - I'm about to hop on a plane. We had to discontinue with the Ekso session because we had 3 sensor errors and 1 motor error number 41 ( see enclosed photos). Could you resolve this issue please, it becomes regular. | Replaced foot sensors via SWO-1629 | Foot sensors are able to be damaged in this manner through regular usage
2578 | Errors 21 and 27 RK, 21 and 27 LH , 1054 | closed | kmettler | Actuators | fixed | 2015-08-07 20:55 | 2015-08-07 | 11694 | 198 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Wear | Issue: Maria - Customer: Negrar - Italy (during ITW) Dennis was able to certify the clinicians however the device is down. "during the training in Negrar, some problems occurs with the Ekso 021502007. Dennis was there, it happened with different patients and also with healthy people, during a normal and good walking, so probably is not a clinical issue. The errors showed are 21 and 27 RK or other time 21 and 27 LH or 1054. From the Data Dump is easy possible find them, in all the last sessions. Actually is not possible use the device, because during every session one of this error is showed and is necessary sit the patient, reboot the device and restart. Considered that the failure happened during the training week, so during the first week of use of the device, is very important give a quickly and clear answer to the customer and make in mode that he can work with the device as soon as possible. We need to know how Ekso want proceed to solve this problem. I'll wait your feedback, thanks!" | When device is serviced by SWO-1871 confirm that issue is resolved and no further action is needed. | Calibration errors caused the joints to not read the full potential of movement.
2576 | Gritty Noise from L Knee Motor | closed | kmettler | Actuators | worksforme | 2015-08-06 22:59 | 2015-08-05 | 34000 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Sandy "Gritty noise‚Äù left knee motor when in free leg and entering stance support. In all other modes the motor sounds ‚Äúclean and smooth‚Äù, as does the right side. | Device fixed via SWO-01351. Confirmation of fix determined when touched via SWO-1766. | Not determined - unable to reproduce.
2575 | Left Hip Sticky & Stiff | closed | kmettler | Adjustments - Hip | fixed | 2015-08-06 22:52 | 2015-08-05 | 34000 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: - Left hip width sticky and stiff, especially when moving from 0 to larger (no video, but there seems to be an initial resistance/catch that needs to be overcome. Then the adjustment is easy) - Left hip abduction moves by 1 after walking | We replaced the left hip abduction assembly. Replaced per SWO-01351. | Brass insert that is guiding lead-screw was worn, possibly from trying to move the adjustment while the lock was still applied.
2573 | R DDL Opens too easily | closed | kmettler | Don-Doff Link | wontfix | 2015-08-06 22:38 | 2015-08-05 | 34000 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Sandy - Right DDL opens too easily (without requiring to add any hip adduction or flexion ‚Äì has never opened during standing or walking, but as seen in the video it is very loose) | Device functions as designed. Device inspected via SWO-1351. No further action taken. | Device functions as designed. No further action taken.
2572 | Black Rubber on Foot Sole Peeling | closed | kmettler | Foot Sole | fixed | 2015-08-06 22:32 | 2015-08-04 | 34000 | 206 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Sandy - Black rubber on foot sole coming off on both feet (this has partly to do with the fact that they were walking on carpet and has been changed, as they have a smooth floor now, but as seen in the photo it has come off in places) | New foot sensors installed, per SWO-01431 | Carpeted flooring in clinic catching on rubber soles
2570 | Screws connecting shank brace broken | closed | kmettler | Lower Leg | fixed | 2015-08-06 20:52 | 2015-08-06 | 900000 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue (occurred today): Jakub reported that the screws connecting left shank brace to lower shank broke. See pictures attached. We have requested additional pictures and we will gather more details how this incidence occurred. | CO-11404 released as CA for Trac 2570 Screws connecting shank brace broken on Technomex device with 900,000+ steps. This is the only know failure of this type. RCA found C100166 screws failed in bending due to 22 in-lb screw torque being too low and loctite buildup between parts along bolt line. This is documented in Trac 2570. Increase screw tension with higher torque (60 in-lb), add washers under screw heads, remove material on shank brace arm along center of two screw bolt line to increase joint bending stiffness. REV B100041-001 ARM, SHANK BRACE TO D FROM C - ADDED .015 DEEP RELIEF CUT AROUND BOLT PATTERN FOR BETTER LOAD TRANSFER AND MOMENT RESISTANCE REV B100049-009 ASSY, SHANK BRACE TO J FROM H - UPDATE HR REV 102220 & 102230 ASSY, LEG, R & L TO G FROM F - INCREASE ATTACHMENT SCREW (C100166) SEQUENTIALLY TORQUE TO 60 IN-LB WITH 243 APPLIED TO SCREW ENDS ALREADY IN BRACE ARM - ADD 4X (2X PER SIDE) C100725 WASHER, FLAT, 0.206 ID, 0.310 OD, 0.020 T, SS 18-8, PLAIN REV B100000-009 ASSY, eLEGS, REHAB, TOP LEVEL TO V FROM U - INCREASE ATTACHMENT SCREW (C100166) SEQUENTIALLY TORQUE TO 60 IN-LB WITH 243 APPLIED TO SCREW ENDS ALREADY IN BRACE ARM - ADD 4X (2X PER SIDE) C100725 WASHER, FLAT, 0.206 ID, 0.310 OD, 0.020 T, SS 18-8, PLAIN - ADD 2X (1X PER SIDE) C100725 TO THIGH BRACE CLAMP SCREW. V&V WAS PERFORMED BY REUBEN USING TEXTBOOK FASTENER DESIGN & FEA COMPARING THE CHANGES. SEE CO-11404 Verification and Validation.docx ATTACHED TO EACH AFFECTED ITEM. FEA AND TEXTBOOK DESIGN STANDARDS VERIFIED THAT THE CHANGES LEAD TO SUBSTANTIAL INCREASE IN JOINT STIFFNESS AND RESISTANCE TO BENDING THUS KEEP THE ALLOY STEEL SCREWS UNDER THE FATIGUE STRESS LIMIT. TEST BUILD WAS PERFORMED BY REUBEN INCORPORATING THE CHANGES ON A GT AND 1.1 DEVICE AND CHANGES IN FUNCTION WERE DETECTED. DOCUMENT CHANGES REVIEWED AND ACCEPTED BY EDMUNDO PRIOR TO ROUTING. | Specified screw torques that were too low for cyclical loading that occurs during use and while inside shipping case.
2569 | Ankle Adjustment sticking; difficult to adjust | closed | rroberts | Adjustments - Ankle | fixed | 2015-08-06 19:24 | 2015-08-05 | 84156 | 137 | VA Palo Alto | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Wear | Ankle adjustment knob will stick in a setting, or not adjust fully to another setting (-3, -2, -1, 0, 1, 2) | Retaining rings in both ankles were replaced, as well as solid wall shank mounts. Screws (C101484) were replaced on both footplates to ankle box for durability and reliability. This work was addressed during SWO-1137 | Small retaining ring was found to be broken due to repeated stress from the chamfered bronze feature in the ankle adjustment. Other ankle was inspected, and found that retaining ring was also broken, but not yet interfering with operation.
2567 | Foot plate screw sheared | closed | kmettler | Fasteners | fixed | 2015-08-05 22:42 | 2015-08-05 | 267000 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | TRUE | 9 - Loaner | Unassigned | Investigation Needed | Design | Issue: please see attached picture. No one was injured and they were able to safely sit the pt. | Installed new ankle box housings on both sides of the device per SWO-1339. | See CAPA #2568 | INV-0001
2563 | noise from the hip/knee area | closed | kmettler | Adjustments - Hip | worksforme | 2015-08-04 19:04 | 2015-08-04 | 187000 | 9 | Craig | Julie Hartman | jwaldie@craighospital.org | (303) 789-8000 | USA |  | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Unassigned |  | Wear | Issue: please listen to attach video, they are getting a noise from the hip/knee area‚Ä¶ they are unable to tell which joint it‚Äôs originating. Please advise if you have any ideas for the customer to check/ trouble shoot. | While working on unit in house via SWO-1553 was not able to replicate error. | Shark fin wear
2562 | Toe Guards Broken | closed | kmettler | Foot Assy | fixed | 2015-08-04 18:57 | 2015-08-04 | 18537 | 169 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Issue : Toe guards broken. Please send replacement to EMAC Genova, so that technicians can replace them. | Sent parts to EMAC via SWO-1319/SWO-1325 | Known Issue
2559 | Ankle is broken | closed | lmcneil | Adjustments - Ankle | fixed | 2015-08-03 18:31 | 2015-08-03 | 366000 | 19 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | TRUE | 9 - Loaner | Major | Investigation Needed | Design | Issue: While walking in the Ekso today they heard a sound... they sat the patient and noticed the ankle had broken from the footplate. I plan on sending them a loaner today and getting their Ekso back to Hq. | Footplate and ankle joint assemble replaced during service SWO-01395 | Evaluated in CAPA #2568 | INV-0001
2557 | Toe guards broken | closed | kmettler | Foot Assy | fixed | 2015-08-03 16:35 | 2015-08-03 | 3706 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issues: Maria - Toe guards broken need to be changed. FSE to go on field next week. Left shank also started rotating but it has been resolved adding 2 washers. " Just to feedback, we have had another issue with the lower leg internally rotating. This time it was the left leg but was resolved with changing and adding to 2 washers on both bolts. Please see attached pictures, we have noticed that the padding/plastic on the front of the footplates has broken off on the right (DSCF1356) and broken on the left (DSCF1357). Would this affect the sensors? If so, would they need to be replaced?" | Toe guards were replaced with SWO-01319. Trac ticket not listed in SWO, however parts issued to SWO and resolution shows work completed. | Normal wear and tear. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2548 | Sensor Error #41 | closed | kmettler | Foot Sensor | fixed | 2015-07-29 21:13 | 2015-07-29 | 114725 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: The Ekso decided to have a few errors again yesterday. It said ‚ÄòSensor error #41‚Äô. All I had to do was sit it and reboot it and then it was fine for the rest of the session. Please let me know if there is anything I can do to fix it and what the prognosis is for Colbey Junior (that‚Äôs the name I have come up with for the Ekso, I think it has a nice ring to it). | New sensors replaced by Michael B utilizing SWO-01312 | Normal wear and tear.
2542 | CI is not reacting to appropriate commands | closed | kmettler | UI - Crutch Interface | fixed | 2015-07-28 21:19 | 2015-07-28 | 52410 | 196 | Children's Hospital of Atlanta | Amber Richards | arichards@chla.usc.edu | 323.361.2568 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Unassigned |  | Design | Customer: Atlanta Children‚Äôs Hospital Unit # 196 Issue: The thumb button the CI seems isn‚Äôt reacting to the appropriate commands, I spoke to Darrell and he‚Äôs said all of the customer PT‚Äôs including himself have had issue‚Äôs with this CI. We are sending a FST to this site so we can make the swap at this time and make sure we get the faulty CI back for analysis. | Provided customer with new interface via SWO-1303 | It allowed for buttons to break and not react appropriately. No further investigated needed. This is a low priority, low risk problem.
2541 | RK #137 Error | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-07-28 17:25 | 2015-07-28 | 51590 | 196 | Children's Hospital of Atlanta | Amber Richards | arichards@chla.usc.edu | 323.361.2568 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: PT called during advanced training (7:19AM PST) with Darrell and reported a LK #137 Error. This was the first front screen #137 Error seen since ITW but Katie has been monitoring this device and there have been multiple back screen errors. As this issue continues with the device and is now being thrown as a front screen error we should come up with an action plan to fix this device. | Inspection during site visit showed that brake gap was greater than .010". The brake disc was removed, inspected and re-installed with a .080" gap. Work done via SWO-1303. Since SWO was already closed when attempting to close this TRAC ticket note of this TRAC was added to the comments section of SWO-1303. | Brake gap has ability to walk given various use patterns.
2539 | Difficult Adjustments | closed | kmettler | Adjustments - Leg | fixed | 2015-07-27 16:32 | 2015-07-24 | 194568 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: I hope your Summer is going well. I have a quick question for you. Is there a particular lubricant that we can use on the Ekso's upper leg shafts, lower leg shafts, and hip width shafts? After repetitive sizing adjustments, it has become much harder to adjust the Ekso. I have cleaned each of the shafts. It seems like we should be using some form of lubricant to allow the parts to move easily. Also, I have our old adjustment tools that had expired calibration dates. Should I just hang onto those until someone from Ekso comes to our site next? | Replaced all clamp components with SWO-01346 | Early production unit. Tolerance ranges have been refined for better fit.
2522 | 195 Errors | closed | kmettler | Software | fixed | 2015-07-24 17:18 | 2015-07-23 | 109393 | 145 | Royal Hospital - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 7 - Clinical Visit | Minor |  | Determined Unknown | Issue: Sandy - I just had a phone call from the head PT at the Royal Bucks Hospital in the UK. Their Ekso has thrown repeated #195 (?) errors and states ‚ÄúSensor error, Ekso must sit‚Äù. They can‚Äôt remember if there was a RK/RH or LK/LH with it, and they are not even sure it was #195 (does not exist in the Ops manual). Additionally, they noticed that Ekso‚Äôs backpack (where the motherboard is housed) got very hot. Motors and batteries did not increase in temp. They have two patients booked in tomorrow at 10am and 2 pm so it would be great if someone could give them a call today until 5pm (UK time) or tomorrow at 9am so they can run through some diagnostics and cancel the patients if need be. The contact phone number is +44 1296 330575 and the best person to ask for is either Eva, Heather or Colbey. | Sandy went to customer site and was unable to reproduce error. | Unable to recreate error.
2521 | Foot Binding Wear | closed | kmettler | Foot Assy | fixed | 2015-07-23 18:55 | 2015-07-23 | 356000 | 19 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Wear | Issue: Please see attached picture. PT reported this was a wear issue and had gotten worse over time. | Sent customer new foot bindings, tools washers and screws via SWO-1295 | Known possible issue. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2520 | Rotating Shank Issue | closed | kmettler | Adjustments - Leg | fixed | 2015-07-23 17:57 | 2015-07-23 | 110682 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Use | Issue: Maria - Technomex advised that one of their devices has a rotating shank issue. They have already 2 washers in each bolt. Jakub believes the tools are no longer calibrated. He couldn‚Äôt provide more details as he was driving. Could you please send to Technomex directly (FAO Jakub) a new set of tools. | Customer sent new tool kit via SWO-1293 | Tools out of calibration.
2514 | Mismatched Batteries | closed | kmettler | Battery | fixed | 2015-07-22 18:17 | 2015-07-22 | 301919 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | Issue: #9 Mismatched Batteries : Error only with battery pair 2A and 2B. The error occurs a few times but not always. I do think that one of two is not fully charged and this causes the problem. Batteries charge up to 98 % | Sent new set of batteries via SWO-1296 | Additional analysis to be done with Ticket #2384 then updated here.
2512 | Error #25LH | closed | kmettler | Foot Sensor | fixed | 2015-07-22 18:14 | 2015-07-22 | 301919 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 10 - Field Service & Loaner | Minor |  | Determined Unknown | Issue: ERROR #25LH during setting set up at the beginning of the session: When reviewing the settings at the start of a session we had a few times a sensor error : #25 LH. After executing a leg test the problem was solved temporarily. I really do think that the intensity of 7 - 8 sessions a day with a session duration of 1 to 1.5 hours a session is too much‚Ä¶ Also patients with a weight of 90 ‚Äì 95 kg is hard to deal with. It happens with Multiple patients on FirstStep and ProStep during SetUp. Usually after they‚Äôve walked some heavy patients (92kg the heaviest). ‚ÄúWe only have this error after using the Ekso during some sessions, not when he‚Äôs cold..‚Äù First time they noticed it was on Monday the 20th of July and again today. Because of this error they stopped using the Ekso. They feel the motors on the left are overheating. We will be sending them a loaner. | Replaced left knee, both foot sensors, encoder and cable (left hip) via SWO-1287. | Signal chain issue unable to determine root cause when fixing this type of issue in the field.
2511 | Shank Rotating | closed | ebabaran | Adjustments - Leg | fixed | 2015-07-21 22:30 | 2015-07-21 | 3706 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Design | Issue: Customer reported right shank rotation, they have already used 3 washers in each bolt and the issue occurs with most patients. Adrian going on site tomorrow to troubleshoot and fix the issue. | Washers installed in the field per SWO-1373 | Adjustment screws can bottom out on the back wall of adjustment collar making it so that the holding torque is not fully achieved
2510 | Right rotating shank | closed | kmettler | Adjustments - Leg | fixed | 2015-07-21 18:06 | 2015-07-20 | 73460 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Issue: Maria - I have received the attached email from Nathan last night: Hi Maria I'm writing from Helpful Steps in Cork, Ireland. On Friday last a member of your ekso bionics staff came to sort some of the problems we are having with our exoskeleton. He replaced some of the sensors in the feet. He then went on to recalibrate the exoskeleton and it worked much smoother than it had beforehand which was great. Two people walked in the exo since. However when I returned today it was worse than ever. 1. Sit to stand was poor. 2. Stepping was slow to respond to weight shift. 3. Step was jerky. 4. Landing of the step was heavy and went into a near lunge position. Over all experience was uncomfortable for both user and therapist. I hope these issues can be resolved promptly as we have a high usage of the exo users on a weekly basis. Please see Dropbox link with videos received from Elite Gym. https://www.dropbox.com/sh/01tya3mwz5jmusw/AACtMEgTiTW3RpqWKA9laFo8a?dl=0 | No SWO was required as no stock parts were needed to address the rotating shank issues. Other mentioned issues were handled through clinical evaluation. | Adjustment screw hit shank collar and doesn't keep lower leg tight.
2502 | Crutch Interface Throwing Errors | closed | kmettler | UI - Crutch Interface | fixed | 2015-07-17 18:09 | 2015-06-11 | 253128 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Unassigned |  | Design | Issue: Josh- Confluence called and said they‚Äôre Crutch Interface is throwing errors, we‚Äôll need to send them a new one. | Customer was sent new CI via SWO-1206 | Down rev CI. Known issue.
2499 | #27LK Error | closed | kmettler | Electrical Interconnect | fixed | 2015-07-16 23:09 | 2015-07-16 | 30033 | 124 | Renown | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: Lily - Renown called and they just had a #27 LK Sensor Error. This happened when they were putting a patient into sit. The Error will not clear Note: A year ago the same issue occurred and they remember Greg going out there to do the fix. #1919 Let me know if we need any additional information. | We changed the main cable harness for both thighs as well as changing the encoder on the left knee per SWO-1281. | Damaged components in the signals led to the issues seen.
2495 | Hip Width Difficult to Adjust | closed | ajulin | Adjustments - Hip | wontfix | 2015-07-15 18:55 | 2015-07-14 | 813666 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Materials | Issue: Aaron I found evidence on the returned hip abduction assemblies that some of the roll pins on the large hip tube assembly were not pressed in fully and they were dragging on the Igus bushings. Service did not (I don‚Äôt think) press the roll pins in further, so we may see future problems of the same nature (hip width difficult to adjust) | FIXED WITH SWO-1263 Device has been decomissioned due to an unrelated failure (broken upper DDL) per SWO-1368. | Roll pins not pressed into adjustment tube far enough
2494 | Repeated #168 Errors | closed | kmettler | Foot Sensor | fixed | 2015-07-15 18:31 | 2015-07-15 | 18735 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Wear | Issue: Sandy - Please see below email from Bad Aibling. Their Ekso (1.1) has been throwing repeated #168 errors. I have no further info but believe they have a video showing a dragging left step. Michael - are you able to get in touch with them? The contact number is in the email signature (Olaf Krapps) | talked customer through the foot test | Foot Calibration Needed: Foot calibration periodically required as routine maintenance. This is a low priority, low risk problem.
2491 | Right Foot rotating 180 degrees | closed | kmettler | Ankle | fixed | 2015-07-15 17:45 | 2015-07-15 | 69622 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Major |  | Process Not Available | Issue: Maria - Customer Elite Gym in Cork #171 SN.021410008, reported issue with the right foot which is rotating 180 degrees. They‚Äôve realized today before the patient got in the Ekso. Also the distributor forwarded me the below email highlighting there are a few more issues than just the rotating foot. They have several patients booked in for this Friday and they require someone to look at this issue at the earliest. "We have had trouble standing the ekso only with 2 people it happens when they push backwards off the crutches and the back of the Ekso gets caught on the waist of their pants. More pressing concern at the moment is the left knee holder seems to allow a bit of movement so people's left knees are drifting inwards. Also when the left foot lands there is a delay in the ekso realising that it has landed and sometimes the right foot has already started its step so you end up with a long pause until it catches up. The lright foot turns in a lot on certain walkers. We have added an extra washer and changed the screws on Maria's advice which did help but on a few walkers it twists very easily. It turns on walkers who are paralysed." | Ankle adhesive failure - replaced both ankle assemblies per SWO-1062 per CO-11291 | Process previously used allows for to much variation that can lead to failure in bond line. New process developed by CO-11445 and presented in CAPA #2723 addresses this.
2490 | Noise coming from knee L side | closed | kmettler | System | fixed | 2015-07-15 17:35 | 2015-07-15 | 296794 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Determined Unknown | Issue: Maria - Please see below email from To Walk Again in Belgium (#53 - SN.011302001). Could you please advise on plan of action? Do we need to guide them through some tests? "As you can see in the video we have some issue (again) with our Ekso. You can hear another sound while moving the knee joint at the left side than on the other side. There is no difference in gait pattern. While moving the knee joint (extension-flexion) in a seated position (knee brake release) there is a hard resistance in moving. We have 6 to 8 sessions a day (duration one session = approx. 1 hour). I hope that the intensity of the sessions is not a problem?" | Replaced left knee, both foot sensors, encoder and cable via SWO-2490 | Unable to investigate part for further RCA.
2487 | Loaner @ Sheffield: DDL Opening Issue | closed | kmettler | Don-Doff Link | fixed | 2015-07-14 17:18 | 2015-07-14 | 268192 | 61 | More Rehab Ltd - UK | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Wear | Issue: Adrian - Chris at More Rehab has the right DDL opening when in walking and going from sit to stand. The patient is about 14 stone 90 kg he is concerned this will happen again tomorrow, so I am going to site to check the ‚ÄúT‚Äù block of movement. He checked to see if the DDL was clean and sent photos, see attached. | changed both T-blocks and the Ekso was walked by the customer, no new issues, customer is up. Verified completion of work via SWO-1651. | Known wear item. This is a low priority, low risk problem. The defect will not be corrected as it is a wear item that is replaceable as needed.
2486 | lower leg shank screw lose or worn out | closed | kmettler | Fasteners | fixed | 2015-07-13 21:16 | 2015-07-13 | 78438 | 113 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | TRUE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: I just got a txt from Vivantes Spandau - their Ekso has a "screw lose" (their words, not mine) I have no further info as I'm now on the road and they have gone home for the day. Can someone call them in the morning 0049 1715958033 (Bettina Quentin, head PT, speaks English) to see what needs doing. | Sent spare adjustment screws to customer to fix this issue and subsequent occurrences. NO SWO created as parts required are not inventoried. | Regular wear seen on adjustment screws. We now include spares of these upon initial shipment.
2485 | #137 Error | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-07-13 20:55 | 2015-07-13 | 83408 | 131 | ULM - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue: Sandy - sind eben mit Hr. Bleher gelaufen. Gegen Ende der Zeit kam die oben genannte Fehlermeldung mit dem Zusatz "EKSO Bionics anrufen - Neustart n√∂tig". Haben ihn dann hingesetzt, Patient ausgepackt und danach noch einmal "leer" aufgestanden. Keine Fehlermeldung. Nun die Frage, ben√∂tigen wir einen Techniker ( der ja eh bald kommen sollte wegen einem Check-Up ), oder sollen wir einfach weitermachen mit der Hoffnung dass es nicht mehr auftritt? Google Translate: are flat with Hr . run Bleher . Towards the end of the time came , the above error message with the words' EKSO Bionics call - necessary reboot " . Have then sat him unpacked patient and then again " empty " got up . No error message. Now the question is , we need a technician (who anyway would soon come for a check-up ) , or should we just go ahead with the hope that it does not recur ? | Unit functioned as designed. Brake slippage was due to customer use. | Use allowed for brake slippage. Unit functioned as designed.
2483 | Error 133 and 121 | closed | kmettler | Electrical Interconnect | fixed | 2015-07-13 16:41 | 2015-07-09 | 167713 | 114 | Halle - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Process Not Available | They had error 133 and 121 RK and their Ekso is down. I have no further info as I'm in the middle of doing a certification. This unit just returned from a retrofit. Based on error log for 7/10, the problem is Error 121 code 64 Right Knee. This means there is a continuity or short to ground issue with the RK motor hall sensor signal wires. The cable clamps and connectors must be checked. | We grounded the on/off switch per SWO-1375. | Torso on/off switch was not fully grounded. At time of original build this fix was not yet available
2482 | The hip joint cap fell out during pt session | closed | kmettler | Actuators | fixed | 2015-07-13 16:29 | 2015-07-08 | 7000 | 207 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Followed | Customer: Good Shepherd Unit # 207 Step Count: 7,000 Issue: We are having major issues with Ekso # 3, our new unit. We had a super user in it this AM. It started only taking a partial step on the left then after that there was a clunking noise that appeared to come from the left hip. Charlene was the therapist working with the patient so they turned around to go back to the gym. The plastic piece that covers where the ball bearings on the L hip popped out and landed on the floor. We then put the patient in stand still and sat them down. This is only the 4 user we had in the device. All the prior users were paras. | Creation of TDO-00100 which requires supervisor approval to ensure that step was properly completed. On this particular unit fixed in field via SWO-1262 | Found that screw C100129 was not installed with loctite 243. This unit originally went through a shipping test in a crate that was subsequently discovered to be not up to spec. The unit was shipped in a proper crate to GS after full op qual of device.
2474 | Thigh adjustment is difficult | closed | kmettler | Adjustments - Leg | fixed | 2015-07-08 20:13 | 2015-07-07 | 735000 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Process Not Followed | 1. Today morning therapists reported me that during walking screw that is responsible for stiffness of dorsi and plantarflexion of ankle is moving during walking and also during swings in standing. Please see movie showing this situation: https://dl.dropboxusercontent.com/u/23745927/20150707_092136.mp4 2. What is more they complain about adjustment of thigh. No are visible but black paint is coming out. Please see attached pictures. | Send ankle box/foot to distributor to repair on unit per SWO-1265. Taught customer process to periodically clean adjustment tube to ensure proper functionality. | Shoulder bolt inakle can break when old version is in play. Process to periodically clean adjustment tube is required to keep proper functionality.
2473 | #164 Error occurs on Ekso when trying to use CI | closed | ckealoha | Other | fixed | 2015-07-08 20:07 | 2015-07-06 | 67000 | 33 | WakeMed | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Materials | Customer: WakeMed Unit # 33 Steps: 67,000 Issue: Device was sent back to HQ for a VAS upgrade. When they received the device back their old CI would no longer work. We sent them a new CI and they had the same result, the Ekso throws a #164 error and won‚Äôt recognize either of the CI‚Äôs they have. We‚Äôll be sending them a loaner and bringing their device back to determine root cause. | Wires were seated into connector housing, and then gently tugged to confirm proper engagement. Cable was reseated and tested per SWO-1282 | Wiring harness assembly was not properly assembled by manufacturer.
2472 | Lower leg is slipping during pt sessions | closed | kmettler | Adjustments - Leg | fixed | 2015-07-08 20:04 | 2015-07-06 | 3706 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Please see pictures from Sheffield regarding Ekso SN.021410003 - #166. The customer reported right rotating shank. Adrian will be shipping new washers to them as they have only one left in their kit. | Washers sent to customer. No SWO was created as parts are freestock. | Adjustment screws can bottom out on back wall of adjustment collar preventing proper holding force from being achieved.
2471 | Lower leg is slipping during pt use | closed | kmettler | Adjustments - Leg | fixed | 2015-07-08 20:00 | 2015-07-06 | 123000 | 138 | Nottwil - Switzerland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Design | Please note Nottwil in Switzerland reported: ‚ÄúThere is a plastic on the footplate, which is loosened and also the two screws at the lower leg (left side) don`t fix against rotation anymore.‚Äù I have requested additional information, pictures and step count. I will update you as soon as I have an answer. Create new ticket for toe guard. | Washers sent to customer. No SWO was created as parts are freestock. | Adjustment screws can bottom out on back wall of adjustment collar preventing proper holding force from being achieved.
2470 | Left ankle sheered from foot plate | closed | ebabaran | Ankle | fixed | 2015-07-08 19:52 | 2015-07-06 | 585000 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | TRUE | 9 - Loaner | Major | Investigation Needed | Design | Hi Adrian, this morning during the walking with the patient, the left foot of Villa Beretta‚Äôs Ekso is gone. Fortunately no consequence for the patient. Please let me know how you intend to proceed. | Fixed by installing new foot ankle assemblies per SWO-01258. | Evaluated in CAPA #2568, insufficent holding strength with previous torque specified. | INV-0001
2469 | Free movement of intrarotation from the hip | closed | kmettler | Adjustments - Hip | fixed | 2015-07-08 19:45 | 2015-07-06 | 336000 | 58 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Wear | While I'm writing you because the Ekso in Pietra Ligure s/n 011305002 has a problem on the right leg. The PT said that it seems like there is a bit of free movement of intrarotation from the hip. Alessandro next week will be there to check it, there is also some other upgrading that they need to do? Because in case we'll do all at the same time. The step count is 336000. | When unit was serviced via SWO-1286. We cleaned the adjustment on both halves of the upper leg adjustment during service and were not able to replicate the issue. | Normal build up in adjustment region requires periodic cleaning to maintain proper functionality.
2468 | #1241 RK Error thrown on Ekso device | closed | kmettler | System | fixed | 2015-07-08 19:37 | 2015-07-01 | 153745 | 114 | Halle - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Process Not Followed | I just had a phone call from BG Halle (Bergmannstrost) in Germany: SN unknown (PT was not with the device) They have had a repeated #1241 RK error thrown this morning. They tried rebooting Ekso, but that didn‚Äôt work and the error comes up straight away. They said that according to the display "Ekso requires a leg test‚Äù , but I don‚Äôt have the instructions, so couldn‚Äôt talk them through it. As the PT was about to go home and wasn‚Äôt with the device, I have arranged the following with her: ‚Ä¢ A tech to call her tomorrow between 12.00 ‚Äì 12.30 German time (UTC +1) to run through some diagnostics ‚Ä¢ I have asked them to have step count and last settings ready for this to help with diagnosing the problem ‚Ä¢ The contact phone number is +49 345 1326359 and the PT is either Jule, Matthias, Rene or Nadine (Matthias has the best English) | We grounded the torso switch and the issue did not reappear. This fix was conducted via SWO-1530. | Torso shroud on/off switch was not fully grounded.
2466 | Ekso having step length issues during pt session | closed | kmettler | System | fixed | 2015-07-08 19:08 | 2015-06-30 | 283000 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | In attach you can find a video concerning the difference in step length. As you can see the steplength on the left side is much bigger than the other side. Our patient also has the feeling that the initiation of the right swing phase lasts longer (maybe due to the bigger left step length?). Our patient feels uncomfortable compared with earlier sessions due to the instability of the different step length. | Walked the customer through leg & foot calibration, no more issues. No SWO required as fix was conducted over the phone. | Foot sensor values can drift due to wear. Calibration of sensor needed periodically.
2460 | Step Lagging | closed | kmettler | Foot Sensor | fixed | 2015-06-25 23:18 | 2015-06-25 | 288220 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: Josh - Customer: Good Shepherd Unit # 132- Internal #3. Issue: I just wanted to let you know that we are having issues with Ekso #3. Today at 0800 we had a regular user in and the Left leg would not step right away. The user had to over shift to the right for the left leg to advance. Then the Left leg got stuck in extension and did not step. (I have a video that I will forward) Then Troy got in the device, the problem did not occur with him in it and he recalibrated it. We had another patient use Ekso #3 and it worked fine. Then the same pt got back into the device at 15:00 and the same thing occurred. No errors were thrown. Please let me know what follow up questions should be asked, this unit has constantly given them trouble so I'd like to get back to them ASAP. | Foot Sensors on the unit have been refurbished to new design specifications. SWO-01289. | Recent updates to foot sensor design have improved reliability.
2452 | Error #42 | closed | kmettler | Foot Sensor | fixed | 2015-06-24 22:29 | 2015-06-24 | 50262 | 135 | GF Strong - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Issue: Customer had Error #42 occur 3 times today while walking with patient. They occurred around every 50 steps taken. Have asked customer to perform foot calibration. | Changed out the foot sensors for the new design under CO-11331. This work was done per SWO-1230. | The heel and toe sensors drift underneath the foot plate causing errors.
2451 | 121 Error Code and Electrical Smell | closed | ebabaran | Electrical Interconnect | fixed | 2015-06-24 22:22 | 2015-06-18 | 246920 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | Issue: 2. Notable: During training today (after walking >3300 steps in the session), with the device at B ~35% Fixed with slow swing complete, the R knee appeared to buckle upon loading. The device was put in Stand and the participant was re-positioned to begin again. We attempted to walk again, but the L LE appeared not to step. We tried again at ~50% Fixed without success. We tried again in B Max‚Äîthe L LE would not step, but, with external assistance to initiate the L swing for the first ~ 3 steps, the device was able to resume ambulation. After ~ 25-50 steps, however, the device stopped with a 121RH code. The participant was assisted to sit. There appeared to be a slight electrical smell. Session was ended. ~ 3 hours later, I retried device with able-bodied pilot for 172 steps in B Adapt without a problem. I am not here tomorrow, and we have no participants scheduled to be in the device tomorrow. We have 3 participants planning to come in to use the device on Monday, starting at 10AM. Please advise. | Unit was evaluated per SWO-1260 and no further action will be taken. | Design of the device allows for the occurrence of 121 errors with regular customer use
2446 | Multiple Error Codes | closed | kmettler | Electrical Interconnect | fixed | 2015-06-23 16:55 | 2015-06-23 | 239145 | 63 | More Rehab Ltd - UK | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 10 - Field Service & Loaner | Minor |  | Materials | Issue: Maria - Please see email below from More Rehab. Adrian is driving to them today and will bring with him a loaner. He will also take their device back to the workshop. Adrian and Chad to provide more details about the incident once available. I have asked the PT to confirm that that him and the patient are fine. "Hi it's Chris from More Rehab in Sheffield. So the Ekso wo do fine last week. Then it worked for an hour today with a client although there was a marked squeaking which was different to the normal squeaking. Towards the end of the hour we pressed hold for walk, then take first step at which point the Ekso safe moded and began to sit without any prior notice, leaving the client sitting on my knee. We have 2 people booked in from 1.45 tomorrow afternoon and 2 on Thursday so the sooner we can fix it the better." | Replaced soft pot on device via SWO-1233 | Bad soft pot
2445 | Battery Not Charging | closed | kmettler | Battery | fixed | 2015-06-23 16:19 | 2015-06-23 | 392830 | 10 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Issue: Maria - Mark Pollock‚Äôs layspotter has reported one Ekso battery is no longer charging. ‚ÄúOne of the batteries won't charge at all. It shows green on the adapter but has no charge when plugged into the Ekso. I've re-tried several times, making sure that the plug is properly connected but can't get anything out of it. The adapter works fine with the other batteries.‚Äù We are working to send him a replacement and we will arrange collection of the faulty battery. | Send Mark two new batteries via SWO-1223 | Additional analysis to be done with Ticket #2384 then updated here. Batteries will not be coming back to HQ
2442 | Error #41 | closed | cmeadows | Foot Sensor | fixed | 2015-06-22 03:43 | 2015-06-18 | 299498 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Issue: Lily- Yesterday (6/18) Good Shepherd had an Error #41. This morning, the PT did a foot Cal and walked around 200 steps and it passed. They had two patients walking in this device today. The first session everything ran smoothly. The second session, The PT said that around 3:00 PM today (6/18) the device threw an Error #41 after around 2 steps. They canceled the rest of their sessions for the day. They are not down as they have two other devices however this has been a persisting issue with this particular device. | Customer Replaced the Foot Sensor and walked through a Re-Cal for 20 steps. Foot Sensor was sent via SWO-1220 | Heel sensor power lead shorted to signal lead. Most likely due to heel sensor physically drifting so contact could be made between the toe sensor solder joint and heel sensor solder joint.
2441 | Possible foot sensor issue | closed | kmettler | Foot Sensor | fixed | 2015-06-22 03:35 | 2015-06-17 | 36590 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Issue: Maria - Matteo from EMAC sent me the attached video. This is from Villa Adria - SN.021406002 - #141. The right step doesn't trigger properly. Matteo will guide them through a foot cal. Any other suggestions? | Replacement sensors were sent per SWO-1225. | Sensors were never returned for root cause analysis.
2438 | Device will not turn on | closed | rsandler | Electrical Interconnect | fixed | 2015-06-22 02:54 | 2015-06-19 | 105109 | 134 | Chuck Slyngstad | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | Issue: Josh - When changing the batteries the unit wouldn‚Äôt power back on. Al has spoken to the PT and believes it to be an Amp Board problem. It appears we‚Äôll have to send them a loaner unit and bring his unit back to HQ. | AMP board has been replaced with SWO-01254. Old amp board is in Engineering for RCA. | Dan - No root cause yet but some additional information. AMP board powers on with a power supply but NOT with a battery. Not obvious shorts found anywhere on the input circuit of LEFT or RIGHT battery connectors. Inspected for clamp diodes or MOSFETS broken and found none. Regen circuit components appear to be working normally, VREF and comparator outputs working okay. Dan - NCR-11542 pulled against this unit. Discovered that Q16 (Motor Power Circuit breaker MOSFTET) failed short which suggests why the board powers with a PS but not a battery due to the inrush current of the motor power capacitor bank and the soft starting capability of the power supply versus the low output impedance of the battery. AMP board is working and ready to go back to production. Dan - There is a design issue with this circuit that is in progress to be resolved (see ticket #2464). The corrective action is most likely to be a revision to the board.
2435 | Weird static pattern on UI | closed | kmettler | UI - Controller | wontfix | 2015-06-17 19:43 | 2015-06-16 | 104068 | 161 | Glenrose - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: ‚ÄúHi. No error report but at ends of session screen doing weird flashy pattern and pushing any buttons does not change it. Here's video. Also have still. Unable to get session data from patient who was in unit at the time. 3-4pm MST.‚Äù I told them to turn off the device and turn it back on. Edmundo also suggested they can just pull out the UI and plug it back in and it will return to normal. | This was due to static discharge and does not affect the function of the device | Static discharge on the UI can still allow this to happen.
2433 | Errors #133 and #121 | closed | kmettler | Motor | fixed | 2015-06-17 18:50 | 2015-06-17 | 246129 | 13 | Mark Beack - South Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Design | Issue: Maria - Please see email below from Mark Beack - #13 - 011202001. I have requested additional information, including a video. However it seems the Ekso is down. "All has been going very well with the EKSO suit. However yesterday apparently it showed a motor error. According to Rob & Avi, the left top hip motor is sounding not too good. The suit will not walk, Brandon had to stop his walking session yesterday. Strange how Murphy pops his head out, just the other day you asked me how the suit was doing. I will get Rob or Avi to contact you about this Error message." | Change out phase cables on this unit per SWO-1354. | Wires for previous cable design can melt leading to the reporting of this error.
2431 | Loaner @ EMAC Italy: Unusual Sound | closed | kmettler | Adjustments - Hip | wontfix | 2015-06-16 05:01 | 2015-06-11 | 124595 | 55 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: Maria - EMAC have notified that their Ekso #55 SN.011303001 currently on loan to Fondazione Padre Pio is making an unusual noise. The noise comes from the right side near the hip and it happens with all the patients when in ProStep+. Dropbox link of the video: https://www.dropbox.com/l/nMAJNp28pliiqvEgkpd9cu They want to make sure the Ekso is safe for use. Could you please advise course of action and whether you require additional information. I have requested the step count. | Customer inspected hips and noted it was normal noise expected to come from a robotic device. | Noise is expected to come from this design of the hips. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2430 | Lower Left Leg Rotating | closed | kmettler | Lower Leg | fixed | 2015-06-16 04:57 | 2015-06-11 | 15867 | 195 | Rehab Institute of Ohio | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: PT‚Äôs reported Lower L leg was rotating during a patient session. This is the first time they‚Äôve had an inward rotation issue with their device. I told their PT‚Äôs to double up the washers on both sides of their device and to report back if the L leg continues to rotate during use. | Had customer add another washer to both legs and this has appeared to fix the lower leg from slipping. | known issue
2429 | Sensor Error #41 | closed | ebabaran | Foot Sensor | fixed | 2015-06-16 04:48 | 2015-06-15 | 6390 | 205 | Barrow Neurological Institute | Susan Wheeler | suzanne.wheeler@dignityhealth.org | (770) 867-3400 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Unassigned |  | Design | Issue: Mike - What Happened: Mid- ITW, we received a Sensor Error #41. It occurred during a Left Step. The Step appeared to initiate, I heard the ‚ÄòStand Still‚Äô Sound and the Left foot returned to the ground and Ekso was placed into Stand Still. A controlled Sit was required and then Reboot Req‚Äôd. I used this as an opportunity for the team to call Tech Support line and Josh walked us through the process. After a reboot, I was able to perform both First and Pro+ without issue and it seemed clear. We received 2 additional Sensor errors. The second appeared in the same context with one of the site PTs handling, the patient (post-tumor resection who appears with L hemiparesis and is 213pounds) was too far forward and the trailing limb sensor was triggered in error. I then jumped in, did not allow excessive forward lean and still received the same error. We‚Äôre breaking for lunch now. If anyone has thoughts/suggestions, I am fully available by my cell phone at 614.296.2881 until 1pm PST before training resumes. | We replaced the foot sensor via SWO-1209 | The heel portion of the foot sensor slid forward and caused the errors. Design change to foot sensors will be required for a permanent fix. Changes were implemented in CO-11331.
2427 | Foot sensor issue | closed | kmettler | Foot Sensor | fixed | 2015-06-15 22:25 | 2015-06-11 | 43817 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Issue: Maria - Sheffield Teaching Hospital reported issue with their Ekso SN.021410003 - #166 (step count 43‚Äô817) The PT reported that the right step doesn‚Äôt trigger easily as the left one and the movement is jerky when walking all patients in ProStep+. Adrian advised them to do a foot calibration, however they have walked even over 60 steps but according to the physio the Ekso didn‚Äôt give any #170 or #171 message at the end. Adrian is unable to follow up with the customer as he is in Czech Republic. Can I have one of the technicians to call the customer to troubleshoot further? Contact: Suki Wong on +44 1142715673 | Re-calibrated foot sensors per SWO-1231 | The customer which was using the device had very poor walking technique which resulted in a bad foot sensor calibration.
2424 | Broken Foot Sole | closed | kmettler | Foot Sensor | fixed | 2015-06-15 21:37 | 2015-06-15 | 16807 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost | Major |  | Materials | Issue: Maria - Please see email below and pictures attached from Chris at Technomex. Customer : Kostancin Jeziorna SN.021411005 - #176 ‚ÄúIn the attachements you wil lsee brand new, echanged approx 1 week ago foot soles. Any hints what could cause a damage like this ? Customer claims that they did not bend it nor stepped on something sharp.‚Äù | Distributor used foot sensors available in their stock that were delivered by SWO-1204. | Variation in material used to produce cap sole by vendor has led to premature failure in extreme cases.
2422 | Left L Lower Leg Slipping | closed | kmettler | Lower Leg | fixed | 2015-06-15 15:24 | 2015-06-12 | 114196 | 165 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Unassigned |  | Design | Issue: Left lower leg is rotating, they reported this issue on Tuesday and we told them to double their washers. They performed this task and are still having the same problem with their lower leg slipping while walking. This unit is also due for a retrofit so we are sending them a loaner unit on Monday and will have their unit back in house next week for further analysis. | New shank collars were installed on unit to remedy the issue per SWO-1213. | Known design issued - Holding force on the lower leg rotation reduced due to screw bottoming out.
2420 | Error #40 Foot Sensor | closed | kmettler | Foot Sensor | fixed | 2015-06-11 23:16 | 2015-05-26 | 342900 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Issue: Lily - Attached is email/picture from one of our therapists at Madonna. Since getting our EKSO back from maintenance we've had an issues where the machine gives us an #40 error code during walking and tells us to sit down and reboot EKSO before we can walk again. This has occurred with multiple patients, however not always happening with those patients, maybe every other session. In addition, as you can see with the picture, we are losing pieces again to the foot sensor. Let me know if we should keep using Ekso as we are able.‚Äù Attached are pictures of the R foot plate of the Ekso, after our session from 11-12 today. A piece came off, as shown in one of the pictures and it looks like another junk is missing from some other session. We did have one episode of "sensor error #40" and it made us sit and reboot. We did regular pivot turns and not veers today, on the carpet, but they seemed normal and the pt was doing quite well in the device, so not sure what the issue is that caused this piece to fall off.‚Äù | Sent new foot sensor via SWO-1176 | This happened because they were walking on carpet.
2419 | CI buttons not working | closed | kmettler | UI - Crutch Interface | fixed | 2015-06-11 21:22 | 2015-03-10 | 68630 | 46 | Guttmann - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Determined Unknown | Issue: Maria - The 164 is happening all the time even if the CI is plugged. It¬¥s seem that buttons of CI are not working. Need to send new Crutch Interface. | Sent new crutch interface via SWO-1047 | Buttons on CI were not working all the time, also it was throwing error 164
2418 | Label not put back on the Ekso | closed | kmettler | Documentation / Labeling | fixed | 2015-06-10 19:42 | 2015-04-22 | 15672 | 195 | Rehab Institute of Ohio | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Workmanship | Issue: CR Inbox - Hopefully, you won‚Äôt be hearing from me too many times‚Ä¶ we just had our Ekso training last week at MVH in Dayton OH. There were some electronic motherboard (?) issues with our device and the tech person Chad came out to fix it on Friday, April 17th (awesome ‚Äì thanks! ‚Äì red eye flight and everything!) We were all sure there was a sticker on the Ekso unit where the service number and other contact info was listed. We couldn‚Äôt find that information today. Are we crazy or could Chad have replaced the part that had the contact info on it when he was here fixing things? If so, could we get another sticker or could you bring it with you when you come back for phase 2 training ‚Äì I am afraid down the road it will be harder to find our written phone number should we need it and it would just be easier if it were on the actual unit. | New label sent via SWO-1183 | Process not followed to replace labels.
2417 | Batteries draining quickly | closed | kmettler | Battery | fixed | 2015-06-09 21:21 | 2015-06-09 | 49785 | 49 | Able Bionics - Canada | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | Issue: Mitch Brogans batteries aren‚Äôt holding a charge. ---- | Sent out 4 batteries via SWO-1199 | Additional analysis to be done with Ticket #2384 then updated here.
2383 | Broken Torso Shroud | closed | kmettler | Torso | fixed | 2015-06-02 23:04 | 2015-04-21 | 34328 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Issue: Matteo - More over, the Ekso GT ‚Äì 021406002 in Ancona, few days ago is falled down from the chair. They needed the torso calibration and the device works well. But there is the plastic part that you can see in the attached picture, that is easy to lift. How I can fixed it or I have to replace it? | Sent new torso shroud via SWO-1113. Fixed in the field by Russ. | Customer dropped unit onto floor.
2381 | Foot sole slightly damaged | closed | kmettler | Foot Sole | fixed | 2015-06-02 17:44 | 2015-06-02 | 112336 | 131 | ULM - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Issue: Maria - Please see pictures below from RKU. The foot sole of their Ekso is slightly damaged. How can we resolve this? | No SWO was created as adhesive was sent out to the customer and they were walked through repairing the issue for this instance and for future use. | Regular usage of device can create this wear, particularly in when walking on high friction surfaces.
2378 | Issue with Turning Mode | closed | ebabaran | UI - Controller | fixed | 2015-05-29 23:08 | 2015-05-27 | 14057 | 178 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue: Maria - Please email below and video from Technomex. The Ekso at Kamien Pomorski (SN.021412001 #178) seems to be unresponsive to turning mode. "Please let me know if we require further details from the customer. One of our therapist during second training reported us strange problem in EKSO sn: 021412001 steps no: 14057. Turing mode is not working in 8 in 10 cases. Device do not react like it should. From what I know it looks like turning mode is off when it‚Äôs on. Please look for video for more details: https://docs.google.com/file/d/0B9z4LRDwfiC7SmVwZXdkVElOUzg/edit?usp=docslist_api Waiting for suggestions from your side." | New UI was sent via SWO-1208 as preventive action. Informed customer that in right or left affected mode the turning mode shouldn‚Äôt work. | This is a training issue.
2377 | Stiff DDL | closed | kmettler | Adjustments - Hip | fixed | 2015-05-29 22:49 | 2015-05-28 | 239145 | 63 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | Issue: Maria - Please see email below from More Rehab - #63 SN. 011307003. Could you please advise on plan of action? "The hip clip on the Ekso has gone very stiff. Any ideas? Can we spray with Wd40 or silicone spray to ease it off?! Or anything else?" | Chad has contacted customer and provided some troubleshooting techniques to help resolve issue. Fix implemented with parts and work done via SWO-01233 . | Down rev version of T-Block B100972-001 allows for more wear which can result in particulate getting into the bullet to housing interface,causing stiffness/interference seen by the customer. This can be resolved through periodic cleaning but longer term by replacing wear components.
2376 | ankle resting angle broken | closed | kmettler | Adjustments - Ankle | fixed | 2015-05-29 22:08 | 2015-05-28 | 15697 | 168 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Eva- I'm in a training and after the second patient of the day the screw of the ankle resting angle has fallen down. We have just found the pieces at the attached picture. Is it safe to continue?? We need to finish the training | Re-install existing ankle adjustment knob and used loctite 263. No SWO was created at time as no stocks were required to complete this work. | Original loctite used (243, medium strength) was not sufficient for all units.
2374 | Left hip motor issue | closed | kmettler | Motor | fixed | 2015-05-29 21:33 | 2015-05-29 | 110029 | 57 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Determined Unknown | Issue: Maria - Please see below email from Joan Romero at Step by Step unit 57- SN. 011305001. I have asked for additional information and for a video. The issue happens in every walking mode. "I'm contacting you because our ekso is starting to give us problems again with the left hip. When doing the step with the right leg, the left hip seems to be block and it's not going forward, making difficult to do the movement on the right. The left motor makes a "zzzz" noise when it happens, too. The thing is that next Wednesday June 3 we have an important conference and our president was going to use the ekso. Do you thing that is it possible to have a technician in Barcelona next Tuesday June 2 (because Monday is holyday in Barcelona) to solve the problem? We would really appreciate that." | Actuator assembly replaced on device via SWO-1320. SWO already closed so added to comment section in Omnify. | Unable to replicate issue as implementation of fix negates ability to do any root cause analysis
2373 | Battery Charge Issue | closed | kmettler | Battery | fixed | 2015-05-28 18:06 | 2015-05-28 | 3706 | 166 | Sheffield - UK | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Issue: Maria - A physio at the Sheffield Teaching Hospital reported issue with battery 1A. She said the battery doesn‚Äôt always charge properly and they get a mismatch error when they place the battery in the Ekso even just after charging it. The chargers work and the light is green to indicate the battery is charged, however when they place it in the Ekso, it shows the battery charge is only 42%. In order to charge it fully they have to charge the battery twice. | New batteries sent via SWO-1180- all 4 batteries were replaced. | Additional analysis to be done with Ticket #2384 then updated here.
2372 | UI not working properly | closed | kmettler | UI - Crutch Interface | fixed | 2015-05-27 20:24 | 2015-05-27 | 115271 | 6 | Thevenaz - Switzerland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Determined Unknown | Issue: Benoit Thevenaz reported that when booting the Ekso (#6 SN.011201001) sometimes they get LCD error #1180. When this happens they have to reboot the device and swap the batteries, which seems to resolve the issue. What could be causing this? | Sent new UI via SWO-1178 | Could not determine root cause.
2369 | 1.1 Tool Kit out of Calibration | closed | kmettler | Tools | fixed | 2015-05-27 06:08 | 2015-05-21 | 262747 | 18 | RIC | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Use | RIC's Tool kit was out of calibration | New tool kit was sent out via SWO-1167 | It is not a problem unless the torque capacity has decreased and then that leads to another.
2368 | Battery Not Working | closed | kmettler | Battery | fixed | 2015-05-26 17:11 | 2015-05-26 | 44022 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Issue: Maria - Colin from Elite Gym reported that one battery of the Ekso is not working. See picture of the SN at the bottom of the email. The Ekso throws a missing battery error, and it doesn‚Äôt work it in either side. When trying to charge the battery, the light of the charger goes green after a few seconds. Please let me know if you require additional information. | Send one new battery via SWO-1175 | Additional analysis to be done with Ticket #2384 then updated here.
2365 | Loaner @ To Walk Again: One Screw missing from shank | closed | kmettler | Adjustments - Leg | fixed | 2015-05-20 15:50 | 2015-05-20 | 185520 | 26 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Followed | Issue: Please see email below from To Walk Again who have on loan unit #26 - 011204003. There is a one screw missing in the shank. Kris confirmed over the phone that they've used the Ekso and the shank is not rotating but they are monitoring it. Adrian is sending them spare screws today. Could we also send them a new tool kit? "As you can see below we are missing one screw. Can we go on like this? There were initially no extra set screws delivered with our Ekso, so we do not have any." | Adrian sent missing screws to customer | The process as found in the box check, PN 102131, clearly states to ensure that all external screws are installed. This process was not followed.
2363 | Unlocking Hip Rotation Issue | closed | kmettler | Adjustments - Hip | fixed | 2015-05-20 00:46 | 2015-05-19 | 342676 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Workmanship | Issue: Forwarded from Maria - The device has problem with unlocking hip rotation. Please see attached photos. This situation is present on both sides. On bolt there is no locking ball. Please send us spare locking pins ASAP and I‚Äôll replace them. What is more I asked months ago to send me spare foot sensors. At this moment we do not have any of them on stock. In case of any problems we will be in big troubles. | Effected component replaced on site by distributor technicians with parts supplied via SWO-1085. | Run off of loctite caused locking mechanism to malfunction
2362 | Hip Width Screw Release Issue | closed | kmettler | Adjustments - Hip | fixed | 2015-05-20 00:22 | 2015-05-19 | 28394 | 17 | Penn Partners | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: Lily - We are having some small issues with our Ekso. The left side hip width screw needs to be unscrewed a great deal to release for adjusting hip width. | We performed CO-10949 to fix the loose bolt issue via SWO-1579. | Hip inserts w/helicoils allow the hip adjustment bolts to loosen during usage
2361 | Double Handle Walker not aligned correctly | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-05-20 00:20 | 2015-05-19 | 28394 | 17 | Penn Partners | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Issue: Lily - Also our new walker is not aligned correctly, it veers to the left and patients are having increased difficulty with walker management. | Sent new walker via SWO-1165 | Walker was aligned incorrectly
2360 | UI magnent not holding to dock | closed | kmettler | UI - Crutch Interface | fixed | 2015-05-19 23:53 | 2015-05-18 | 555022 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Issue: Doing advanced training in Bad Aibling with their 1.1 The left magnet that holds the UI to the torso is not working. The safety handles are tight and Ekso is working fine apart from that. Is there a quick fix I can do while I'm here or do we need some new parts? I will be here until tomorrow (Thur) 12.00 so any plan of action is much appreciated. | Customer was sent new UI via SWO-1159 | The left magnet was not holding to the UI
2357 | Knee Issue | closed | kmettler | Adjustments - Leg | fixed | 2015-05-16 00:21 | 2015-05-14 | 85193 | 123 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Internal Training Insufficient | Issue: I have just received a call from Matteo at EMAC. He reported that our Ekso loaner #123 which arrived at Villa Beretta yesterday, has an issue with the knee. The knee remains in extended position when the Ekso is switched on , legs free and also when trying to press the knee release button. Apparently it just doesn't move. Matteo has requested a video and will send it across at the earliest. We need to resolve this as soon as possible. We don‚Äôt have another loaner to give to Villa Beretta and their device is at the UK workshop (not started yet). | Replaced screws that mate the carriage to the ball screw per specifications. Work was done per SWO-1232. | Screws holding the ball screw to the carriage were torqued to specifications, however was not visually checked for a complete mating of the parts. Technician most likely heard the torque wrench "click" indicating that proper torque had applied, but was unaware that screw was not flush. Removal of screw did not find errant debris in hole, but screw had a small burr on threads possibly creating a binding situation.
2356 | Unusual Step | closed | kmettler | Customer Use | fixed | 2015-05-15 23:18 | 2015-05-15 | 87201 | 149 | SDSU | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Training | Issue: Please see below video- this patient is having an issue with one of his steps. This has only occurred with this patient, PT did mention she has seen this with someone else once before but hasn't been an issue with their device until recently. https://www.youtube.com/watch?v=xbshlMwcQ94 | Melissa informed the customer of proper clinical settings and patient strength interactions. Fix was implemented on site through customer training no SWO created. | Issues with training with one patient
2353 | Rotating Shank | closed | kmettler | Adjustments - Leg | fixed | 2015-05-13 22:57 | 2015-05-12 | 38241 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Elite Gym in Ireland unit #171 - SN.021410008, 38238 steps, reported rotating shank during walk. They will send me a picture of it, in the meantime I asked them to add an extra washer and reminded them that they should be alternating tightening. | Customer issue was solved when the customer added the second washer. | Holding force on the lower leg rotation reduced due to screw bottoming out.
2352 | Broken Crutch Interface | closed | kmettler | UI - Crutch Interface | fixed | 2015-05-13 22:20 | 2015-05-13 | 6643 | 120 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Use | Issue: Sandy - Please see below email from Vivantes Spandau: Essentially the CI cable got stuck on the motor during sit and the cable ripped. I have no further info and I am with a customer so can't help right now. I will be visiting them next week though so if need be I can take a new CI out if it's sent to me. | Sent new Crutch Interface via SWO-1144 | Customer broke UI during use.
2350 | Stepping Issue | closed | kmettler | System | fixed | 2015-05-11 22:52 | 2015-05-11 | 192000 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Determined Unknown | Issue: Josh - Our unit experienced a total failure this morning. We started walking in B Adapt with the participant that had been describing that ‚Äúfree swing‚Äù sensation. I was prepared to hit the yellow button as an event marker (as we‚Äôd discussed last Thursday), but the participant did not experience the sensation. We progressed to B Fixed 60% without event and then progressed to 55%. Shortly thereafter (after about a total of 19 minutes of walking), without warning, the machine appeared to ‚Äústumble‚Äù (the participant was promptly caught and did not fall or sustain any injury). Once I was able to ensure the participant was safe, I hit the yellow button. I wish I could describe the event better, but my attention was focused on ensuring the participant‚Äôs safety. After the ‚Äústumble‚Äù (while in Stand mode), the R knee did not lock for several minutes (while my assistant obtained a w/c), then finally it appeared appeared to lock, perhaps due to the participant shifting his weight. We cautiously attempted to resume ambulation, but the R LE would only lift and lock up in flexion. We took a video of that event, which I will attempt to forward to you in a separate e-mail. The participant was than assisted to sit using the stand-to-sit program, however, the transition was not smooth. We attempted sit-to-stand, but the unit did not perform. We did not power down and try again, because I believed I‚Äôd asked enough from the participant, and there was obviously something wrong with the device. | We replaced the knee motor. SWO-01154 | Determined unknown, replaced all parts that could possibly be the issue.
2349 | Loaner @ Health South: Foot Sensor Issue | closed | kmettler | Foot Sole | fixed | 2015-05-11 22:18 | 2015-05-11 | 44079 | 182 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Workmanship | Issue Please see attached picture. I have a call into their PT's to get more info regarding this failure. | Chad field serviced this device and brought spare parts with him via SWO-1135 | Mating nylon piece was not properly sanded.
2348 | Error 168 | closed | kmettler | Electrical Interconnect | fixed | 2015-05-11 17:41 | 2015-05-07 | 3102 | 45 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Issue: Maria - Please see below email from To Walk Bionic in South Africa - Ekso #45 SN.011211002. Their Ekso is throwing error 168. We are waiting for further information. "We are experiencing an error that we can't find in the book- error 168? Any ideas?" | Customer performed foot calibration and passed test. | Foot calibration periodically required as routine maintenance. This is a low priority, low risk problem.
2345 | Error 41 | closed | kmettler | Foot Sensor | fixed | 2015-05-08 22:56 | 2015-05-07 | 3123 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | Issue: Duisburg‚Äôs device #152 SN.021407007 is down. They will send me the step count via email. The Ekso is throwing error #41. They tried to perform a foot calibration but after about 50 steps, error #41 comes up again and when they switch on the device the error appears each time. We need to look into it as soon as possible (today) as we don‚Äôt have a loaner available. | Foot sensors replaced via SWO-1729. | Foot sensors can fail with regular usage
2337 | Ekso Struggles to Stand | closed | kmettler | Customer Use | fixed | 2015-05-06 15:30 | 2015-05-06 | 29606 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Issue: Our exoskeleton, at times, struggles to stand some users in one swift motion. Colin has to help the exo up to reach its full standing position. Can you please advise on what the problem may be? Barry requested additional information and in particular: Patient (weight, condition, spasticity) does this happened with. Does the Ekso report an error code Video of the issue | Clinical team consulted with customer and explained they needed to spot patient to ensure enough force was being used to help the Ekso stand. | Ekso didn't have the power to stand patient, spotter wasn't supplying enough assistance.
2336 | Motor Error: 121 RK Code 45056 (motor short) | closed | kmettler | Motor | fixed | 2015-05-05 23:59 | 2015-05-05 | 172000 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Determined Unknown | Issue: Today during training in a bilateral low Fixed value (~40%) with slow swing complete, the Ekso had a Motor Error: 121 RK Code 45056 (motor short) The device went into Safe Mode, and we lowered the participant into a wheelchair. We freed the legs and were able to get up and resume training. We did not get any more errors, however the device appeared to intermittently complete some right swing cycle as if it were in Adaptive Mode (vs. ~40% fixed with slow swing complete). We continued the training in Adaptive and higher Fixed values. We trained our subsequent participant in the same fashion. Both participants say the right leg definitely ‚Äúdoes not feel right‚Äù (no pun intended). There were no further errors. | The knee motor was replaced.SWO-01154 | Determined unknown, replaced all parts possible to be issue
2333 | Rotating Left shank | closed | kmettler | Adjustments - Leg | fixed | 2015-05-05 17:46 | 2015-05-05 | 67559 | 155 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Design | Issue: ‚Ä¢Left shank is rotating slightly with a couple of patients (both of them tall and heavy). I have asked them to send pictures if it happens again. Also they have confirmed they were not alternating tightening the top screw and the bottom screw. They are using new calibrated tools. | Customer added double washers from their spare tools kit. | Holding force on the lower leg rotation reduced due to screw bottoming out.
2332 | Right Knee Battery Depleting Quickly | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-05-05 17:30 | 2015-05-05 | 67559 | 155 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Use | Issue: ‚Ä¢Right knee battery depleted quickly. It had been changed at the end of March. Customer confirmed nothing is accidentally pressing the knee release button causing the battery to deplete fast. | 9Volt was mailed to customer and replaced by them in the field. | Customer was depressing button while fitting patients.
2331 | Device down LK issue | closed | ajulin | Electrical Interconnect | fixed | 2015-05-04 16:38 | 2015-04-28 | 39718 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Workmanship | Issue: Russ - Ood Morning all, Attached you will find an error log of interest. There are several of the below errors in the log, Error 25, ----------------------------- LK, Pot/enc mismatch And Error 21, RK, and LK most of the time, an occasional RH and LH----------- Current limit exceeded There is no error 27 associated with the log. Barry tried recalibration of feet, not of joints. Torso test is OK Brakes seem functional ‚Äì OK SYMPTOMS: Left leg is affected, will freeze motion or vibrate. | Re-seated connector. | Found that a cable was preventing the brake cable to fully seat in the connector, providing an intermittent connection.
2328 | Walker does not lock into place | closed | cmeadows | Cane / Walker / Crutch | fixed | 2015-04-30 16:53 | 2015-04-30 | 6529 | 120 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Issue: Sandy - Please see below email ‚Äì essentially the walker at the clinic in HUK Berlin is broken and it does not lock into place at all now (apparently they have had this problem for a while). They had fixed it successfully before, but now it is beyond repair. As they have a lot of patients they will need a new functioning walker ASAP if possible. Maria ‚Äì I have asked for the correct shipping address. The best person to contact regarding this is Bettina Quentin: Bettina.Quentin@vivantes.de | Send customer new walker via SWO-1121. Design fixed with CO. | Pin inside walker is too big so would not drop down and lock into place.
2317 | Stiff Hip Width Adjustments | closed | kmettler | Adjustments - Hip | fixed | 2015-04-27 23:00 | 2015-04-27 | 124809 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Materials | What: Sandy - They also mention the following issue they are experiencing with their device: Stiff hip width adjustment They have not reported any error codes or other issues, this all seems to be purely mechanical. Can we get a plan of action together to address these issues? I guess at 124000 steps they could do with a service too. Maria ‚Äì Are you happy to communicate with them, their English is ok. If not let me know. I have replied to their email telling them we‚Äôll get back to them with a plan | Installed new hip width keys PN 103195 per SWO-1116. | Found that the hip width keys PN 103195 were not made to specifications and were causing the hip width adjustment to be stiff.
2316 | Very Stiff DDL | closed | kmettler | Don-Doff Link | fixed | 2015-04-27 22:58 | 2015-04-27 | 124809 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | What: Sandy - They also mention the following issue they are experiencing with their device: - Very stiff DDL They have not reported any error codes or other issues, this all seems to be purely mechanical. Can we get a plan of action together to address these issues? I guess at 124000 steps they could do with a service too. Maria ‚Äì Are you happy to communicate with them, their English is ok. If not let me know. I have replied to their email telling them we‚Äôll get back to them with a plan | Reworked device to include grease in affected areas. Work was verified completed per SWO-1472 | Early units weren't built with grease. This issue has been addressed and resolved per the changes identified in CO-11155.
2315 | Two broken toe guards | closed | kmettler | Foot Assy | fixed | 2015-04-27 22:54 | 2015-04-27 | 124809 | 144 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | What: Sandy - They also mention the following issue they are experiencing with their device: 1. Two broken toe guards They have not reported any error codes or other issues, this all seems to be purely mechanical. Can we get a plan of action together to address these issues? I guess at 124000 steps they could do with a service too. Maria ‚Äì Are you happy to communicate with them, their English is ok. If not let me know. I have replied to their email telling them we‚Äôll get back to them with a plan | Installed new toe guards as outline in CO-11239. This work was completed per SWO-1116. | Known issue where the previous design created a stress concentration which in turn caused premature toe guard failure. This issue has been addressed and resolved per to the changes in CO-11239.
2314 | L toe guard broken, R is cracked | closed | kmettler | Foot Assy | fixed | 2015-04-27 15:50 | 2015-04-27 | 19699 | 153 | Ekso Bionics EU | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service |  |  | Design | What: At Rehab Karlsruhe, the Ekso's toe guard on the left has broken off and the right has a crack. SN 021407008 It's ok to use, but will need fixing sooner rather than later. If you send some toe guards we can fix it unless it's due back in Adrian's workshop anyway? | Sturdier new version of toe guards have been installed on the affected unit. work was done per SWO-1124 | Known issue where the previous design created a stress concentration which in turn caused premature toe guard failure. This issue has been addressed and resolved per to the changes in CO-11239.
2312 | Broken Toe guard | closed | kmettler | Foot Assy | fixed | 2015-04-23 02:48 | 2015-04-16 | 13360 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service |  |  | Design | Issue: When transporting the Ekso at the end of the session the footplate got caught on the rollator and the front decorative edge came off of the L foot plate. It doesn‚Äôt affect the unit, but aesthetically the black piece is not uniform when comparing the L to the R foot plate. | Parts mailed to and replaced by customer via SWO-1261. | This is a known issue. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
2311 | Broken Right Shank Mount/ Plantarflexion Wall | closed | kmettler | Ankle | fixed | 2015-04-20 17:40 | 2015-04-20 | 549480 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | From Customer at Technomex - Poland: One of our EKSOs has encountered a broken right ankle joint‚Ä¶.. during the patient sit-to-stand transition. EKSO SN is 021404004 EKSO step count is 549480 Photo attached. | Unit repaired in field via SWO-1115 | Design allows for cracking on plantar flexion wall of shank mount. Part revised to solid wall version via CO-11291.
2309 | Knee doesn't move in Stand Up | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-04-16 17:22 | 2015-04-16 |  | 57 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Determined Unknown | Where: Step by Step Unit: 22 What: Maria - Today we have tested our esko. I'm sorry to say that we keep having the same motor problem. We tested the ekso with a T4 ASIA A client (no spasticity) and with expertise in the use of ekso. I'm surprise to say that we had the same motor problem (I think it's the right knee). We only could do a few steps. We sat down and now we cannot get up again. Please fins attached the videos. Let me know how to proceed. | Actuators replaced on device via SWO-1320. SWO was already closed in omnifiy so added to comments section | Unable to perform root cause on parts as repair eliminates opportunity to perform this task.
2308 | Ekso threw error 121 | closed | ckealoha | Electrical Interconnect | fixed | 2015-04-15 19:09 | 2015-04-15 | 230000 | 115 | NRH - Med Star | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Wear | Customer: NRH Unit #115 Steps: 230,000 Issue: While walking a patient the Ekso went into SafeMode. They got Error code 121, they also noticed a very coggy R Hip motor. The error is just looping now and the hip is making an unusual sound, Al spoke to the PT on the phone and believes the hip motor has shorted. They are currently down. | I changed the amp board and reprogrammed the ELMO's and the cogginess on the motor was gone.Work was done via-SWO1091 | I tested the motors by changing the legs from left to right and the issue stayed with the same side. I determined the issue to be with the amp board.
2302 | Ekso turned off completeley | closed | kmettler | Torso | fixed | 2015-04-13 16:55 | 2015-04-10 | 26258 | 133 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Solution Not Implemented | Where: Uppsala - Sweden Unit: 133 What: Maria - Please see below email from Uppsala in Sweden. I have asked to confirm they experienced the following when the Ekso turned off completely: -no power to hips -knee brakes -no power to LCD -no indication of power on the power button Katie, can we retrieve the data they requested? "Something strange happened during a training session on 7 april. During the session the Ekso turned off completely when the patient walking in the Ekso had to stop due to the opening of a door in front of him. Is it possible to see from the transmitting data what happened at that specific time, around 2.25 pm (Swedish time, probably 1.25 pm in the UK). Is this a problem that occurred before or maybe it¬¥s because we made something strange. The problem has not occurred again and the Ekso is working as usual. If possible we would like to get the data from that specific training session, lasting between 2.00 - 3.00 pm Swedish time, because when the Ekso turned off we lost the data for the first part of the training session." | Issue was not replicated after this initial instance both by the customer and while in Ekso possession. Grounding of power switch fix was implemented via SWO-1563 | Grounding for power switch fix via CO-11171 was not implemented.
2297 | Vibram heel cup peeling | closed | kmettler | Foot Assy | fixed | 2015-04-08 22:05 | 2015-04-07 | 124973 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Where: EMAC - Fondazione Padre Pio Unit: 127 What: Broken part on foot plate. Adrian is going to the customer on Thursday for other repairs but I am not sure he has parts to do these fixes. | Vibram reglued on device when serviced via SWO-1119 | Improper adhesive spec'ed out for this particular application.
2295 | Broken Left DDL Skirt | closed | kmettler | Don-Doff Link | fixed | 2015-04-08 18:36 | 2015-04-07 | 124973 | 127 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Where: EMAC - Fondazione Padre Unit: 127 What: Maria - Broken Left Skirt Adrian is going to the customer on Thursday for other repairs but I am not sure he has parts to do these fixes. | This issue is being verified per SWO-1119 | Material chosen can crack after extended regular usage. When the thigh brace is set too low on patients this can cause a crack due to the bend placed on assembly.
2287 | Phone stopped sending data | closed | kmettler | Pulse Data | fixed | 2015-04-07 23:41 | 2015-04-01 | 121834 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Where: Burke Unit: 142 Issue: phone locked up and stopped sending data. | Sent new phone via SWO-01099 | Known issue, phone's lock up.
2273 | Loaner @ Halle: Missing screws on right side cover | closed | kmettler | Fasteners | fixed | 2015-04-06 19:29 | 2015-03-31 | 137000 | 109 | Halle - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service |  |  | Process Not Followed | Where: Halle - Germany Unit: 109 What: Customer in Germany Halle has on loan #109 SN.021401002 whilst their device gets retrofitted (I expect this to take another couple of weeks before we can send it back to the customer) and has reported some issues: please see pictures attached and email Google translated: At our loaner we have found the following defects today and now raises the urgent question whether it is still operational. See photos. -Missing two screws on the right side of the cover. Is the loaner safe for use? | The missing screws were pushed back in by hand. Ultimately the device will need cylindrical loctite applied but this will be addressed next time the unit is touched. The walker locks were pushed down into place when aligned with the hole on the frame. Verified completion of work to replace missing screws via SWO-1157 They were informed that if the walker was left open it would remain safe for use. | They were missing two screws on the right side of the cover.
2271 | Loaner: Foot Sensor ripping | closed | ebabaran | Foot Sole | fixed | 2015-04-06 17:31 | 2015-03-25 | 30000 | 188 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Workmanship | Unit: 188 What: Josh - Attached are some pictures of the left foot of the loaner Ekso that we are using. The therapists using it this morning noticed this when they were turning. We wouldn't be able to use it any more I don't think, I doubt the sensor will work much longer. | New foot sensor was installed per SWO-1068. | Nylon top piece not sanded fully to edge where primer and adhesive are applied. Right heel and toe sensors installed upside down in left capsole.
2270 | Broken Walker @ RIC | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-04-03 18:25 | 2015-03-24 | 213368 | 18 | RIC | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Materials | Where: RIC Unit: 18 What: I have pictures attached. However they do not explain the problem well. The first walker that has high and low handles, the back poles on the bottom no longer will hold in place. On the other walker the bottom where the caps are have worn through and no longer moves in a straight forward path. Could you all help us with this? | Sent out a new walker via SWO-1072 | Bottom of walker caps were worn through as well as the back poles of the walker no longer hold into place.
2269 | Broken Walker | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-04-02 22:59 | 2015-03-31 | 137000 | 114 | Halle - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Customer in Germany Halle has on loan #109 SN.021401002 whilst their device gets retrofitted (I expect this to take another couple of weeks before we can send it back to the customer) and has reported some issues: please see pictures attached and email Google translated: At our loaner we have found the following defects today and now raises the urgent question whether it is still operational. See photos. Deficiencies in their own walker: The lock on both sides after unfolding the device is no longer fixed. Can the walker be repaired or do we need to send them a new one? | New walker sent via SWO-01075 | The lock on both sides after unfolding the walker was no longer fixed.
2265 | Broken Walker | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-04-02 17:47 | 2015-04-01 | 391302 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Where: Good Shepherd Unit: 108 Issue: Just wanted to let you know that one of our rolling walkers broke. The one side just keeps swinging open and does not lock in place. | Sent new walker via SWO-01088 | Known failure.
2264 | Lower shank slipping | closed | kmettler | Adjustments - Leg | fixed | 2015-04-02 16:38 | 2015-04-01 | 22553 | 155 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Where: Linkoping - Sweden Unit: 155 What: Maria - ‚ÄúThis is happening with 2 patients, the leg is rotating inwardly during the walking (not from the start but after a short walk) even after they have chosen the middle ( of 3) - neutral and Double Checked that the 2 screws are tighten. This is causing then the feet to kick together and causing a bad quality on the walking. We have been trying to figure this out, when you are adjusting the length on the upper thigh/femur is there a certain flexibility in the rotation also in that adjustment. That adjustment haven¬¥t moved after the screws has been tightened but we believe that is of importance according to the position concerning the rotation it will be in when you lock it. This was not at all on the table during the training ‚Äì probably because this is maybe not where you are suppose be able to adjust rotation? It is a matter of only a few degrees inwards or outwards but we experiencing anyway that it is significant for the walking pattern here. We tested our hypothesis today and gave the patient on the pictures the small inwards rotation that the upper thigh/femur permits and after that we could walk significantly longer before we had to make a stop to adjust the rotation ‚Äì especially on the left foot. Could this be a hint? Think, however that the construction should be powerful enough to be able to retain the settings that are made before walking. Now it seems that if ‚ÄúWrong‚Äù rotation in the upper thigh/ femur it causes a tension/stress in the Ekso skeleton which it will give into?!‚Äù | Sent customer two new T-Wrenches via SWO-1102. The customer also received double washers to put on at the shank screws. This was able to solve the issue. | Tools expired - wear item. Washers - Holding force on the lower leg rotation reduced due to screw bottoming out.
2255 | Battery Draining quickly with patient | closed | kmettler | Battery | fixed | 2015-03-29 21:50 | 2015-03-28 | 6222 | 179 | Care Techniques - Saudi Arabia | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Design | Where: Care Techniques - Saudia Arabia Unit: 179 What: Barry - I am in Saudi Arabia using the distributors / customers device and noticed a 34% power loss with walk time 8.09 steps 374. Dale is using bilateral max. Batteries in question 2a 0124EC651700008D 2 0162C8EA16000045 I have swapped for the 2nd set and will observe and update accordingly 2nd set of batteries 8.40 walk time 439 steps with 14% battery loss. | Sent customer new batteries via SWO-01082. | Additional analysis to be done with Ticket #2384 then updated here.
2248 | Battery is loose on the charger | closed | kmettler | Battery - Charger | fixed | 2015-03-26 23:30 | 2015-03-26 | 3123 | 152 | BGU Tubingen - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Process Not Clear/Complete | Where: BGU Tubingen Unit: 152 What: Sandy - I got an email from BG T√ºbingen stating that one of their batteries is siting loosely on the charger. I have asked for more info including: ‚Ä¢Does it still charge? ‚Ä¢Are the pins ok? ‚Ä¢Is it just the one battery? ‚Ä¢Is the charger loose on the other three too? ‚Ä¢Step count I guess we‚Äôll wait to hear back, but I assume they might need a new charger or battery depending on the answer to the above questions. Anything else I should have asked? | Sent a new battery charger to customer via SWO-1106 Routed CO-11326 to improve MPI. | (see follow up email with attached pictures and videos) Affected assy is: B100909-009 ASSY, BATTERY CHARGER Initial investigation indicates the a few contributing Root Causes for two unique issues. '''Issue 1''': charger is loose on battery Workmanship: Two of the three screws holding B101828-001 HOUSING, MIDDLE, BATTERY CHARGER GEN 2 not being tightened enough to hold the part firmly against the lower charger housing. This allowed the connector to shift away from the Battery when docking and that allowed slop between them that was noticed by the customer. Materials: The part B101805-001 HOUSING, LOWER, BATTERY CHARGER GEN 2 .105 dia holes are not as deep as the solid model specifies them. They should be .325+/-.015 in deep but the drawing doesn't indicate their depth since the part is cast Urethane. Solid Concepts does not do detailed part inspections so adding the depth would not prevent this issue. Process not complete: 4.16.2 states "Hand tighten screws until just snug and do not use Loctite." The screws were snug but B101828-001 was not, and had noticeable slop. The MPI should indicate tighten the self tapping screws until B101828-001 is snug AND the board/connector can still shift laterally without resistance. Perhaps we should add a final QC (NOT QA) to assemble the charger to a battery to ensure it attaches without slop. '''Issue 2''': cable can rotate inside charger housing Process not complete: Similar issue to Trac#1925 in which the lip on the cable strain relief (74125K81) was carved away. The issue found with this customer battery charger was that one of the flats on the strain relief was carved away so the battery charger housing didn't prevent cable rotation as designed. This could lead to cable failure over time.
2246 | Software Error Codes | closed | ckealoha | System | fixed | 2015-03-26 00:39 | 2015-03-25 | 199105 | 132 | Good Shepherd - IP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner |  |  | Determined Unknown | Customer: Good Shepherd Unit: 132 Issue: Lily - Regretfully I must inform you that our detestable #3 device had another major issue this afternoon similar to the last event it was in the shop for. The device threw 3 error codes all at once; Software error #1154 RH , Software error #1241 RH, and Sensor error #20 RH. This time however, it occurred about 6" into the sitting (decent) process vs walking. The patient then had to be lowered the rest of the way manually and we were unable to free the legs. After reboot the errors just kept coming and then finally it went into safe mode and said "joint test req". The device had been operating without issue during the walk. As for the patient profile he is a regular user, a very light weight complete paraplegic with zero tone. He is a different user than the last 2 incidents. I will have time at 8am or 3pm to do whatever you need me to test or suggest. We do not have double bookings for the Ekso tomorrow however Thursday & Friday we do. | The fan cable was replaced and installed according to the current manufacturing standards which help mitigate this problem. This work was done per SWO-1069. | Upon receipt of the device we found that the right hip fan cable had been damaged. The jacket of the red line which provides the current going into the fan had at one point in time was perforated. Based off of how the circuitry works, the fan cable was constantly drawing power from the 5v line and discharging the electricity into the carbon fiber plate which the damaged cable was in contact with. When the charge of the plate had reached a critical value, a direct path from the fan cable to the chassis would be created and in turn cause a short. The errors which were seen are the same as in ticket 1763 which was found to be a shorted fan cable.
2240 | Loaner @ Step by Step: Error #137 | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-03-23 16:43 | 2015-03-20 | 193360 | 22 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Wear | Where: Step by Step - Spain Unit: 22 What: Maria - Please note that Step by Step are currently using loaner unit #22 and they‚Äôve reported error #137. The clinician advised: When it happens, the ekso stops walking and the message is : Error 137. Then we have to sit and let ekso rest for a while. If we restart quickly, the ekso use to give the same error n¬∫ in a few minutes. Usually, after a 30 min break, we can go for another workout. The customer‚Äôs device is with Adrian awaiting to be repaired (sidebar cracked). Please advise how this issue could be resolved. | Reset the brake gap according to the current manufacturing process. Work was done per SWO-798. | It was found that over time that the brake gap had widened and in turn caused the 137 error.
2239 | Irritation from thigh pad's | closed | lmcneil | Soft Goods | fixed | 2015-03-23 16:26 | 2015-03-20 | 3123 | 164 | VA Minneapolis | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Testing Insufficient | Where: Minneapolis VA Unit: 164 What: Josh - Additionally, I just wanted to let you know that we had some skin issues with the Ekso thigh supports‚Ä¶patient has areas of irritation where the thigh support ends on both thighs(about mid quad). He is 209# and has some quad activation, larger quads than an average patient we see here, so it has been more difficult to fit him comfortably within those thigh supports. They are ‚Äúout‚Äù as far as they can go, we had this issue before with another patient that has larger quads. The patient said he wasn‚Äôt even going to mention it because ‚Äúit isn‚Äôt a big deal‚Äù but I wanted to let you know. We will take extra measures to ensure the skin remains intact, and make sure the patient is properly aligned within the device at his next session so maybe it won‚Äôt happen again ÔÅä | Sent customer new thigh pads via SWO-01051 CAPA #2016. | Patient sample size was not enough to detect failure.
2238 | Right lower shank is rotating | closed | jwitherspoon | Adjustments - Leg | fixed | 2015-03-23 15:44 | 2015-02-23 | 232152 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Where: Villa Baretta Unit: 140 What: Maria - Villa Beretta reported the right lower shank of their Ekso is rotating. Looking at the video it seems this is not safe for use. | Completed by distributor. | Holding force on the lower leg rotation reduced due to screw bottoming out.
2235 | Rotation of the lower shank when walking | closed | kmettler | Adjustments - Leg | fixed | 2015-03-19 23:50 | 2015-03-11 | 8408 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Where: EMAC - ITALY Unit: 141 What: However the customer also reported: - Rotation of the lower shank when walking | This change was done per SWO-1055. | The length of the screws which are used to keep the lower shank from rotating are at times to large due to tolerance issues. This extra length prevents proper clamping. We have in turn added washers underneath the screw to decrease the overall height of the threaded section to increase clamping force.
2234 | Hip width changing sizes while walking | closed | kmettler | Adjustments - Hip | fixed | 2015-03-19 23:50 | 2015-03-11 | 8404 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner |  |  | Materials | Where: EMAC - ITALY Unit: 141 What: Matteo from EMAC resolved replaced both skirts in the Ekso of Villa Adria. However the customer also reported: - Hip width changes when walking | Found that hip width key, pn 103195, was not made to spec. The piece in question was reworked on site so as to conform to specifications. This work was done per SWO-1055. | Found that hip width key, pn 103195, was not made to spec.
2231 | Battery Fit Issue | closed | kmettler | Battery | fixed | 2015-03-19 23:17 | 2015-03-19 | 9836 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Customer Use | Where: Technomex's Customer Site Unit: 176 I was at a one if Technomex's customers today to look at the ankles of units 176 and 177 and they complained about how none of the batteries easily fit on the left side of the Ekso. I took a look at the batteries and I found that the customer had placed a large label on top of our label on all of the batteries. The added height in conjunction with our Ekso serial number label caused this interference. We removed their labels and the problem was gone. I instructed the customer to place their label on the exterior surface of the battery to avoid this issue. No further action is required. | Additional Label was removed. | Technomex's customer had a large label over the Ekso original label making it difficult to fit on the Ekso
2229 | Client was dropped in Ekso | closed | kmettler | System | fixed | 2015-03-18 20:22 | 2015-03-18 | 41264 | 135 | GF Strong - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Mishandling | Where: GF - Strong Unit: 135 What: It‚Äôs not a panic. I can speak to someone tomorrow, one of us dropped a client while walking and the Ekso screen tells us a torso test is required. I think it might be due to the swing complete feature confusing the client /clinician but I am not sure since the display was turned off and back on. Your customer support is not available at the 888 number. I have called and left messages using the 510 number too. Hope you had a good flight. Ian 604 737 6410 Then, after you have spoken to me:- 604 734 1313 and ask for local 2201 (that is the phone number in the room where the Ekso is) | Recalibration. | dropping the exoskeleton caused the sensors to go out of range. The torso tests are required if we ever see a motion we do not expect from the exo at the torso. The knee pot likely got shifted in the fall and got an out of range error which was fixed with recalibration
2228 | Loose Shank Issue | closed | lmcneil | Adjustments - Leg | fixed | 2015-03-18 17:48 | 2015-03-17 | 18556 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Wear | Where: To Walk Again - Belgium Unit: 53 What: Maria - Please see pictures of the loose shank from To Walk Again, device #53 - 011302001. I requested the step count. Please advise if this is safe for use. | Cleaned off excess material using IPA. Work was done per SWO-1142. | Found material build up on the shank and inside the tube which adversely affected the clamp's grip.
2226 | Noise coming from the knee when pushing release button | closed | kmettler | Brake / Release / 9V Battery | fixed | 2015-03-18 16:52 | 2015-03-17 | 18002 | 155 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Customer Use | Where: Linkoping Unit: 155 What: Maria - Linkoping have reported a noise coming from the knee when pushing the release button. They have changed the knee battery but the noise persists. | The correct battery was replaced with the correct Energizer Alkaline 9V battery. | The battery was replaced with the wrong type of Energizer Alkaline 9V battery.
2223 | Stiff Left DDL | closed | kmettler | Don-Doff Link | fixed | 2015-03-16 16:51 | 2015-03-11 | 8408 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Available | Where: EMAC - ITALY Unit: 141 What: Matteo from EMAC resolved replaced both skirts in the Ekso of Villa Adria. However the customer also reported: - Stiff left DDL Would Matteo be able to fix these? | Added grease to ddl per SWO-1055. | The initial installation process did not require for grease to be added into the ddl. Over time, the lack of lubrication caused excessive wear, which in turn caused the ddl to become very stiff.
2222 | Loose ankle joint | closed | apoultney | Ankle | fixed | 2015-03-16 16:21 | 2015-03-16 | 150000 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Where: BGU - Duisburg Unit: What: Maria - Here is our Problems with the Ekso. The Joint Play in the foot is a Little Bit to much. You can see it. in the video. Left side. We have another Problem with one of the Motors, it makes a Strange Noise when we move the leg up and down when the Ekso is shut down. | Sent replacement ankle to field via SWO-1049. | Known failure point fixed with CO-11230 and 11291
2221 | Rust on Screws Area | closed | kmettler | Fasteners | fixed | 2015-03-16 15:57 | 2015-03-16 | 22968 | 171 | Rolling Ball - Ireland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Materials | Where: Rolling Ball - Ireland Unit: 171 What: Andy - Colin O'Shaughnessy pointed out to me that the fixing photographed is not adequate. It appears to be made of lightweight plastic. The two other photographs show what appears to be rust on the surface of the screw areas. Perhaps you would raise thus with the appropriate people and let me know the response. | Reuben explained that we have not had a failure on this part before but that we appreciated the input and will consider using a stronger buckle in the future. Upon testing several of this buckle to failure in tension, we found their capacity to be 120-180 lbs. Consider a 220 lb Ekso patient falling forward and caught at 45 deg by the safety tether without arm/crutch support. Assuming the torso is 50% of the body weight AND the anterior pad and side pulls are for some reason not attached, the maximum load that would be applied to each shoulder strap buckle is 94 lbf. Since the anterior pad and side pulls are always attached while in use, the actual load applied to the buckles is much less. As an added safety, at 100 lb, the webbing SLOWLY slides out of the buckle to absorb energy without snapping. | The customer was concerned that the fixing photographed was not adequate due to the fact it was made of lightweight plastic.
2217 | Crack on the right knee joint plastic cover. | closed | kmettler | Lower Leg | fixed | 2015-03-13 21:39 | 2015-03-10 | 322426 | 45 | Justin Smith - So. Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Materials | Where: Just Walk Bionics Unit: 45 What: Our right knee joint has cracked a little, right by the screw. It's just the outer casing so not affecting the walking but we obviously need to fix it. Can we do it ourselves somehow? Or what would you suggest? | Will confirm completion of work via SWO-1885. Knee covers will be sent together with the needed soft goods and the Updated Phone to the customer. The Customer already received the tools to exchange the covers, but not the parts. Parts will be sent to the customer together with the Phone. | Plastic part is capable of cracking through regular usage.
2214 | No seat pad included in shipment | closed | kmettler | Soft Goods | fixed | 2015-03-12 20:18 | 2015-03-12 | 7107 | 178 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Followed | Where: Technomex - Poland Unit: 178 What: Maria - There was nor pillow for the chair inside the box ‚Äì so missing pieces during the packaging. | Seat pad was sent to customer via SWO-01044 | Not able to confirm via packaging Photo since it wasn't taken. If not included in shipment, process not followed.
2213 | Ekso shipped with without training packets | closed | jstoll | Documentation / Labeling | fixed | 2015-03-12 17:41 | 2015-03-12 | 7107 | 178 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Process Not Clear/Complete | Where: Technomex - Poland Unit: 178 What: Maria - Second time, there were no clinical materials( so-called yellow envelope ‚Äìwe had to print it on-site, like in Konstancin) | Shipped new documents with SWO-01048 | Training packets were only shipped for US training. Process changed to ship for all units.
2212 | Lower shank tube epoxy bond failed on Right Ankle | closed | kmettler | Ankle | fixed | 2015-03-11 16:55 | 2015-03-11 | 9519 | 178 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Clear/Complete | Where: Technomex - Poland Unit: What: Krzysztof - This a very urgent case: on the third day of training in Kamien Pomorski, on a brand-new device, that was a subject to inspection by Chad last week, see what happened to the right ankle joint ‚Äì the pipe support has slipped out of the unit. Question is, is it possible to fix it on-site ? the whole pipe went out just a minute ago! | Unit fixed in field via SWO-1029 | Workmanship SOP didn't define training requirements and methods. Training was not provided to tech prior to performing work. Workmanship standards do not cover epoxy use and preparation rules. Not enough resolution provided in assembly method to provide for control and repeatability of process needed.
2209 | Error #42 came up on screen | closed | kmettler | Foot Sensor | fixed | 2015-03-10 17:06 | 2015-03-10 | 406571 | 147 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Where: Technomex - Poland Unit: 147 What: Jakub - I am really sorry to report that unit SN:021407002 Steps: 406571 has problem. After unit stand up and make first step error #42 shows on screen. We tried it few times and it happen every time (with and without patient). Please advice what to do. Can we fix this problem or I should return forwarder with loaner unit?? | Replaced with SWO-982 | See Trac #2140
2205 | Issues with Right Knee | closed | kmettler | Battery | fixed | 2015-03-06 21:07 | 2015-03-06 | 18556 | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | Where: To Walk Again - Belgium Unit: 53 What: Maria - I have received a call from To Walk Again and their Ekso won‚Äôt sit up. This occurs with and without patients. Device #53 SN. 011302001 When standing up the right knee goes up to an angle of 45 degrees and it won‚Äôt progress any further, then gets timed out and it goes in sitting position again. The device doesn‚Äôt throw any errors, however the right knee makes a strange noise. I have requested a video of the issue and they will send it to me at the earliest. This is very urgent because they have a press conference on Tuesday and they need the Ekso up and running. | During the inspection of the device it was found that a cable leading to the battery had been pinched in a previous battery replacement done by the customer. It is believed that when the customer switch which side the battery was on, that this pressure was relieved, and in turn caused the device to function as expected. All work was completed per SWO-1142. | Customer accidentally pinched brake battery cable during a brake battery replacement.
2197 | Device throwing error 168 | closed | ebabaran | Foot Sensor | fixed | 2015-03-02 17:17 | 2015-03-02 | 3123 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.d | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Unassigned |  | Materials | Where: Duisburg - Germany Unit:152 What: Maria - Manuel received a call from Ralf Lanzmich, Lead Physio in Duisburg. The device is throwing Error 168. It walks 3-4 Steps and stops, this happens when walking all patients. Could you please advise course of action? | It was determined through further analysis that this customer required new foot sensors. New sensors were sent to customer via SWO-1064 | Worn foot sensors required replacement
2195 | Battery no longer charging | closed | cwhite | Battery | fixed | 2015-03-02 16:53 | 2015-03-02 | 317002 | 6 | Thevenaz - Switzerland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Where: Switzerland Unit: 6 What: Benoit Thevenaz reported 1 Ekso battery no longer charges and when the Ekso is switched on with that battery on, it throws error ‚Äúmismatched batteries‚Äù. I have asked him to send me the serial number of the battery in question. Also he advised he is regularly using both sets of batteries but he cannot remember if this battery is part of the new set we shipped to him at the end of October. | Sent customer new batteries via SWO-1030 | Additional analysis to be done with Ticket #2384 then updated here.
2186 | Left Hip squeaking badly | closed | kmettler | Adjustments - Hip | fixed | 2015-02-27 17:59 | 2015-02-27 | 122819 | 10 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service |  |  | Wear | Where: Mark Pollock - Ireland Unit: 10 What: Maria - The left hip is squeaking badly. It is not unlike a sound that we heard last year on one of the devices that eventually turned out to be a major problem. | New gusset cover has been installed, greatly reducing the "squeaking" sound. Small squeaks still exist, due to the mating part (the hip abduction) not being replaced. Functionality and safety are not compromised in this repair. All fasteners have been torqued to spec. Work was done per SWO-1110. | Gusset cover had worn down contact points where repeat adjustments and / or loosened bolts had allowed minute flexing of the hip to torso assembly. Points of contact had worn enough to begin abrading the anodization from the surrounding areas, causing a powder like debris that accentuated the squeaking sound.
2185 | Multiple errors 121RK (code 12288), 133 (code ??), 1030, 1031, and burning smell | closed | rsandler | Electrical Interconnect | fixed | 2015-02-26 22:50 | 2015-02-26 | 108182 | 8 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Design | Original Reported Issue: Maria - I just spoke to Matteo at EMAC who has been told by the physiotherapists at the University of Pisa that device #8 - SN.011201003, threw errors 121RK & 133. This happened only once when walking a SCI patient complete. '''5/22/15: Summary of errors and disassembly''' Katie on 2/26 121(Elmo motor failure bit set) error code was 12288(undervoltage, all joints) Katie on 2/26: Frequent Intermittent IMU data issues: 1030(IMU response echo mismatch), 1031(IMU returned with error errors) over multiple boots. New issue we haven‚Äôt seen before. There may be noise on the SPI line or something else is going on. They get a low battery warning after 42-48 mins as well. Maria on 2/27: The customer reported they can smell something is burning in the Ekso. Maria told them to stop using Katie on 2/27: Ironically, there were no 1030/1031 errors today‚Ä¶ they had a few instances where the left hip was pulling a lot of current, Katie on 3/11: After about a week without any, the 1030/1031 errors returned today. Greg on 5/21: When I received the unit back from engineering I removed the amp board and CPU for replacement. I noticed some heat damage to the doghouse cable connectors. The RK connector (same as initial 121 error) was so melted I was unable to get it out of the amp board. Pictures attached. Reuben on 5/22: Adding to Greg's findings, all four motor connectors have melted insulation at the board connectors that expose the wire strands. No other signs of heat damage on the board. RK connector black and clear wires are melted together with some insulation compression as shown in attached Photo IMG_0772.jpg. This occurred right where there is a tight bend in the clear wire and black wire. Affected cables are Rev G of B100154-109/209 (thigh) and B100153-109/209 (hip). Rev G should include hipot test introduced in Rev F: PERFORM 1000V HIGH POT TEST BETWEEN HIGH CURRENT PINS A1-A2, A2-A3, A1-A3. It seems probable that the melting motor lead insulation was related to the motor errors 121/133, high current draw, quickly draining battery, and burning smell. But, it is not clear what root causes led to these effects. It seems unlikely the IMU errors are related to the motor errors as the high current battery and motor leads are on the opposite side of the board than the IMU. '''5/22/15: Questions''' No-Dan-Could having two motor phase wires touching at a bend cause increased resistance and heat? Not possible-Dan-Can you analyze the cables and board for abnormally high resistances? No-Katie-Did IMU errors (1030/1031) ever occur at the same time (or boot) as the Motor errors (121/133)? Not avail-Katie-Was there a back screen code for Error 133? -Katie-During the week between 2/27 and 3/11, was there any use? Maria instructed them to stop using. KS: Yes on 3/11. It turned on a few times in between but just for a bit which was likely so we could get data. --Katie-Could the 1030/1031 errors have occurred just from turning device on prior to packing to send back to HQ? --Maria, when did they get their loaner device and pack Unit 8? -All-How might the IMU errors (1030/1031) be related to the Motor errors (121/133)? KS:The IMU errors occured on mulitple boots, repeatedly. The motor error occured 1x on 1 specific boot in the middle of a storm of IMU errors. --Could high current or temp affect IMU? --Might they be unrelated? KS: Yes --Can a magnetic field cause 1030/1031 errors? --If the battery power leads were routed over the IMU could high current draw cause 1030/1031? Russ on 5/29/2015: NEW CPU PN B100566-009 B3 failed to program. Request replacement to facilitate troubleshooting and repair. New CPU and Amp board resolved issue. | New Amp board and CPU installed in Unit with SWO-01166. Changed wire size of phase wires on cables B100153-XXX and B100154-XXX from AWG 18 to AWG 14 so that wire fits properly into Hirose contacts | Melting occurred due to high resistance from poor fit of AWG 18 wires in hirose power connectors.
2184 | Ekso Error's during sitting relating to customer use | closed | kmettler | Customer Use | fixed | 2015-02-26 22:37 | 2015-02-20 | 122819 | 10 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Training | Where: Mark Pollock - Ireland Related to #2170 Unit: 10 What: Adrian - We walked in the new Ekso today for the first time this morning. As I was doing "move joints" on the first set up, a sensor error came up. I cleared it and we stood up fine after that. Mark walked for an hour and the Ekso felt smooth throughout. As we sat down at the end (2nd sit as we changed batteries halfway through), the Ekso locked in safe mode half way down with a motor error 133, I cleared that one and another motor error RH 121 came up. We had finished the hour so didn't try standing again. Any advice? | Maria told customer to slow down while sitting so this issue didn't occur. | Known issue when sitting technique is not followed as trained.
2181 | Loaner @ Step by Step: Side Bar Cracked | closed | kmettler | Adjustments - Hip | fixed | 2015-02-26 19:35 | 2015-02-26 | 72761 | 57 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Design | Where: Step by Step - Spain Unit: 57 What: Joan Romero - We have a problem with our ekso. I'm attaching with this mail a picture (just near the width selector). At this moment it's very difficult for us to work with the ekso because the width selector is almost blocked. Maybe it's being time for a check out. Let me know how to proceed. | Unit fixed in the field via SWO-1024. Additional parts identified as required and shipped via SWO-1197. | Torso side has cracked in the past after extended use
2175 | Ekso shoulder strap broken | closed | kmettler | Soft Goods - Shoulder Straps | fixed | 2015-02-25 17:35 | 2015-02-25 | 316542 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Wear | Where: Good Shepherd OP Unit: 108 What: This morning at Good Shepherd the R shoulder strap let loose while a patient was walking. It tore at the sewing point that attaches the double female end clip to the actual strap pad in the upper section. They were unable to use their Ekso due to this issue. We have already sent out a new shoulder strap and they will be up and running again tomorrow morning. | Sent customer new right shoulder strap via SWO-01015. PENDING prototype of new design to prevent these issues going forward. Waiting for Cerice/Beth to approve Zipporah's time. | see CAPA #2296
2170 | Ekso went into safe mode with motor error #133 & error RH #121/ Over voltage | closed | kmettler | System | fixed | 2015-02-23 17:15 | 2015-02-23 | 122819 | 10 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Related to #2184 What: Simon O'Donnell - We walked in the new Ekso today for the first time this morning. As I was doing "move joints" on the first set up, a sensor error came up. I cleared it and we stood up fine after that. Mark walked for an hour and the Ekso felt smooth throughout. As we sat down at the end (2nd sit as we changed batteries halfway through), the Ekso locked in safe mode half way down with a motor error 133, I cleared that one and another motor error RH 121 came up. We had finished the hour so didn't try standing again. Any advice? Maria on 2/26: The ekso sat okay on the half hour mark to day but got an error 133 and 121 on the next sit. Below is the data requested: Error Code Count 121 3 133 1 1058 3 Error String Count Elmo motor fail 3 Motot drive fau 1 CAN PDO receive 3 Error value Count 20480, RH 3 0, 0 1 58, RH 3 Eric on 2/26: Here is what the Elmo documentation says about the over voltage error: Voltage of the power supply is too high, or the servo drive failed to absorb kinetic energy while braking a load. A shunt resistor may be required. The over-voltage threshold differs with the power stage model (refer to product user manual). A probable reason for this protection is an attempt to decelerate a motor that is moving so fast that the drive fails to absorb the motor‚Äôs kinetic energy. The solution to this is to decelerate more slowly or to add a shunt resistor to the power supply. This definitely makes sense for the case where a large patient is sitting. It will be interesting to see if the other times this error occurred are also linked to a motor deceleration. | Fix iterated through various services but fix finally verified via SWO-1209 with change out of dog house cables. | Need RCA when device in house covered by #2184: The combined weight and speed causes the accelerated sit. In addition to unit functioning as designed (error being pushed out)the change of dog house cables as driven by CO-11357/11402 helps to mitigate this issue.
2169 | Don doff link skirts cracked | closed | kmettler | Don-Doff Link | fixed | 2015-02-23 16:42 | 2015-02-23 | 8408 | 141 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Customer Training | Where: EMAC - ITALY Unit: 141 What: Maria - Matteo from EMAC (Italian distributor) sent me the attached pictures regarding device SN 021406002, #141 located in Villa Adria, Ancona. I am afraid I don‚Äôt know the name of this part, but the pictures show these are broken on right and left leg. Also they scratch the patient when walking in the Ekso. Please advise on action plan. | Replacement parts sent to customer via SWO-1010. Training and documentation provided such that customer could replace parts on their own. | Thigh brace is being set to low on patients causing this issue. Plastic DDL skirt can crack due to bend placed on assembly.
2167 | Loaner @ Technomex: Speaker on unit was not functioning | closed | kmettler | Torso | fixed | 2015-02-19 23:44 | 2015-02-06 | 37578 | 123 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Followed | Where: Technomex - Poland Unit: 123 What: Adrian reported - Speaker on unit was not functioning. Speaker shroud assembly has been replaced and unit is up and running in use by customer. | Adrian plugged the speaker back in on site. No parts were required, no SWO was written up. | Speaker assembly was found to not be properly plugged in upon inspection.
2166 | Device throwing error #168 | closed | kmettler | Foot Sensor | fixed | 2015-02-19 19:53 | 2015-02-19 | 232152 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Where: EMAC - Italy Unit: 140 What: Maria - I spoke to Matteo at EMAC who reported unit #140 SN. 021406001 located in Villa Beretta is throwing error #168. This is happening when walking patients in ProStep+. The left step doesn't trigger and the device throws error #168. They have done a foot calibration but it hasn't solved the issue. Matteo has requested an additional set of foot sensors. Should the replacement of the foot sensor resolve the issue? | Sent two new sets of foot sensors to customer via SWO-01006 | Left toe sensor was not functional. Do not know if it was electrical interconnect or sensor issue.
2164 | 178: Mixture of German and Polish in initial screening form | closed | kmettler | Documentation / Labeling | fixed | 2015-02-19 18:58 | 2015-02-16 | 3807 | 178 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Process Not Available | Jeff - Update CO will need to be made Complete - In progress: Therapist at Technomex will resolve the translation issue | Therapist at Technomex translated the document. CO-11296 Update Translations of Clinical Forms - closed | The forms were not properly checked before sent out.
2163 | 177: Mixture of German and Polish in initial screening form | closed | kmettler | Documentation / Labeling | fixed | 2015-02-19 18:53 | 2015-02-16 | 3868 | 177 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Process Not Available | Where: Technomex- Poland Unit: 177 What: The pictures attached show a mixture of German and Polish in the initial screening form. The Ekso measurements form format and translation also needs correcting. I suggest we speak to Technomex and their therapists to work in resolving this quickly. | Therapist at Technomex translated the document. CO-11296 Update Translations of Clinical Forms - closed | The forms were not properly checked before sent out.
2162 | Broken Toe Guards | closed | kmettler | Foot Assy | fixed | 2015-02-19 17:51 | 2015-02-19 | 370210 | 147 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Broken toe guards found on unit. | Installed new design of toe guards which removes this stress concentration. Work was done per SWO-1096. | Previous design allowed for a stress concentration to form which in turn allowed enough stress to build causing fracture.
2161 | Broken left ankle plantarflexion stop cracked | closed | kmettler | Ankle | fixed | 2015-02-19 17:49 | 2015-02-19 | 370210 | 147 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Solution Not Implemented | Where: Technomex - Poland Unit: 147 What: Technomex physio, Gregory, reported that whilst he was demonstrating the device to a hospital in Konstancin, the left ankle broke. | Ankles from unit 174 swapped out onto unit 147 in field per SWO-1007. Ankles for 147 will still have to be swapped out in the future. | Design
2143 | Rivet popped off side closure | closed | kmettler | Soft Goods | fixed | 2015-02-18 22:03 | 2015-02-17 | 268670 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Where: Madonna Rehab Hospital, NE Unit: 119 What: Today we had a rivet pop off that attaches to the bottom Velcro strap that we tighten across the torso pad. Is it ok if we have our brace guy (who does rivets all the time) fix this? We wanted to make sure it was ok with Ekso before we fixed anything. | Replaced side closure via SWO-01004 | Had the old design rivets.
2135 | 176: Mixture of German and Polish in initial screening form | closed | jstoll | Documentation / Labeling | fixed | 2015-02-16 17:52 | 2015-02-16 | 7163 | 176 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | Unassigned | Minor |  | Process Not Available | Where: Technomex- Poland Unit: 176 What: The pictures attached show a mixture of German and Polish in the initial screening form. The Ekso measurements form format and translation also needs correcting. I suggest we speak to Technomex and their therapists to work in resolving this quickly. | done - new SOP for translations created done - new translation done of Polish forms routing through Omnify CO-11296 Update Translations of Clinical Forms - closed | The forms were not properly checked before sent out.
2133 | Ekso box damaged making it difficult to put the Ekso back in | closed | kmettler | Shipping / Handling / Packaging | wontfix | 2015-02-13 22:17 | 2015-02-13 | 3797 | 181 | VA Bronx | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Mishandling | Where: Bronx VA Unit: What: Today, at the end of the training at the Bronx VA, they asked me to help them store the Ekso back in the shipping box for the weekend. It felt much different then I am used to and was very challenging. Previously I was able to simply lay the Ekso down to position it in the box. Today I had to flex it slightly at its hips to get it to fit. If this is a new design and way of doing this, for a person my size, the old way was much easier. Then, the box didn't close/ lock as expected. It seemed to be slightly off in alignment between the top portion and the bottom. Perhaps it was damaged in the shipping process. I have included pictures. I eventually got it to close, but it was more challenging then it typically is. | The PT was able to eventually close the Ekso inside the crate. | The box didn't close/ lock as expected when PT was trying to put the Ekso in the crate. It was most likely damaged in shipping.
2132 | Errors #1054 #27 & #20 | closed | rroberts | Electrical Interconnect | fixed | 2015-02-13 21:57 | 2015-02-06 | 6802 | 132 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Determined Unknown | Where: Good Shepherd OP Unit: 132 What: You will find the attached photos of the errors ( #1054 - RH Software, #27 RH sensor, #20 RH sensor) This was the 2nd event today with 2 totally different patients in our #3 machine. After the 1st event I performed torso test, and both R & L / HIP & KNEE joint test in side-lying and in hanging position. It checked out okay. However approx. 30' into our 3pm user's session it threw the same errors , came to an abrupt stop and she had to be a controlled lowering. Our #2 machine is still here and we can unpack on Monday utilizing the loaner over here again. Just let us know whether or not someone is coming to fix or ship before I pack it. Should #3 fall out the plane somewhere over the Rocky Mountains on its way back you will not hear us complain Ha Ha! It has not been a friendly machine since new! Today it cost us 3 sessions alone. | Replaced thigh to knee cable per SWO-1018. | Thigh to knee cable (102557) was defective and has been sent to MRB for analysis.
2100 | Swing Leg Kicks Backwards | closed | kmettler | Customer Use | fixed | 2015-02-12 20:32 | 2015-02-09 | 245000 | 47 | Scripps | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | Customer: Scripps Unit #47 Steps: 245,000 Issue: Josh - I saw Melissa and Darrell last week at a conference and I was asking them about the situation where the swing leg kicks backwards. Melissa said she has never seen that occur and that I should let you know. It has happened with us on a number of occasions. It happened most recently Monday 2/9/14 with pt ID 0000. It happens most frequently when we are crossing thresholds (tile to carpet) and there is a small lip to go over. I think the patient shifts his weight forward over the stance leg, but then in anticipation of this threshold, he starts to shift the weight back a little. It might be useful for your team to check out. We usually put the Ekso in standstill when this happens and just regroup/start over. | Our PT, Melissa, shared that this leg kicking back issue was with only one patient, and also that this patient was very small in size. This issue only ever happened with the specific patient and Melissa was able to walk them through several tips to help | This was cause by the patient, and the device was being used of of intended use.
2099 | Creaking noise that occurs at the end of stance phase on the R | closed | kmettler | Torso | fixed | 2015-02-12 20:11 | 2015-02-12 | 245000 | 47 | Scripps | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | Customer: Scripps Unit #47 Steps: 245,000 Issue: Here are two videos of the noise we are hearing. It is somewhat difficult to hear the noise in the video because there is so much noise going on. However, there is a creaking noise that occurs at the end of stance phase on the R. The noise is very soft at the start of the session, but increases during the session. We haven‚Äôt been using the Ekso too much over the past couple of weeks. However, it is occurring with all of our patients. | Normal noise expected to come from the unit. | Normal wear for Ekso.
2097 | The dial on the R ankle turns a little during stance phase of gait | closed | kmettler | Adjustments - Ankle | fixed | 2015-02-11 20:08 | 2015-02-11 | 317000 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Customer Use | Customer: Good Shepherd Unit #108 Steps: 317,000 What: Josh- Please see attached video- do we have any suggestions to pass along to the customer? Please advise. Issue: I did notice on Ekso # 1 the dial on the R ankle turns a little during stance phase of gait. I did video tape it and will send that to you. I tried to zoom in on it. Let me know if you have any suggestions. | As long as the knob doesn't change between number settings, the ankle turning a little is normal. There is a small amount of play between the knob and the knob's locating feature on the ankle housing due to tolerance requirements. | Customer noticed the dial on the R ankle turns a little during stance phase of gait
2096 | Down Rev of walker sent to Mt. Sinai | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-02-10 21:24 | 2015-02-10 | 288475 | 28 | Mt. Sinai | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Where: Mt. Sinai, NY Unit: 28 (011205001) What: "We received a replacement walker several months ago after the one we were given with the VA upgrade was broken during shipping. However, the replacement one is the older model (no lower hand placement area). The one we are using now isn‚Äôt as tall as the one we were given with the VA upgrade. Is there any way to get a taller walker? We have a few patients who are struggling with the shorter one." | Complete-Replacement sent out via SWO-00989 SHIPPING FIX: We have now taped off an area for accepted and in process material and are labeling all items clearly that we are unable to keep in the inventory cage. | Complete -walkers were not being properly stored and labled as inventory while rework was being performed resulting in an old walker being pulled. We have now taped off an area for accepted and in process material and are labeling all items clearly that we are unable to keep in the inventory cage.
2094 | 2 of the 4 batteries wont charge above 20% | closed | kmettler | Battery | fixed | 2015-02-09 20:47 | 2015-02-09 | 52000 | 33 | WakeMed | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Materials | Customer: WakeMed Unit # 33 Steps: 52,000 Issue: PT called and reported 2 of their 4 batteries won‚Äôt charge above 20%. | Sent out new batteries via SWO-00984 and closed out SWO-00984 now that batteries are debugged in house. | Customer reported two bad batteries, however upon inspection we found only one bad battery that had a weak cell #2. The other three batteries that came back had lower than acceptable capacity and are now marked for R&D. 0162492015000085 - Capacity: 0 Charge percentage: 0 This is the battery with cell number 2 reporting 0 volts. It measures 0 volts with multimeter. Gave battery to Cerice to have Refurbed. Cell 2 is definitely dead. No reasonable explanation as to why. Cells go bad sometimes. 01CD70201500008A - Capacity: 1.65Ah Charge percentage: 98%. On R&D Shelf 016193201500003F - Capacity: 1.69Ah Charge percentage: 98%. On R&D Shelf 01ED7920150000CF - CapacitY: 1.67Ah Charge percentage: 98%. On R&D Shelf documented on SWO-00984
2090 | Send Customer Replacement Walker | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-02-06 19:04 | 2015-02-06 | 208672 | 19 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Where: RIC, Il SN: 011203005 What: "So, the reason I'm reaching out is that the newer walker (the one with the curves so you can push up from a lower level) that we have is broken. The walker does not lock in place when we open it up. Can we get a new one or is there a quick fix, new screws? Have you found this to be a common problem?" | Replacement sent out via SWO-00983 | Walker does not lock into place when it is opened.
2083 | Ekso went into safe mode, displayed error #1063 | closed | gloudermilk | System | fixed | 2015-02-03 21:47 | 2015-02-03 | 339294 | 28 | Mt. Sinai | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Design | What: We had another session today with second new set of battery, and yet again we still had previously mentioned issue, so we had to sit and switch the batteries to continue the session. Today Ekso also went into safe mode without any warning (battery level was ~40%). Displayed error #1063; we did the long sit, and Ekso displayed error #1057-LK. Also we were unable to go into free leg mode despite control option being displayed. Ekso was then turned off and restarted and error #1051 was displayed and was still unable to go into free legs. I know you advise me to have an able bodied person walk in Ekso; however, I will not be able to do that until this Friday. And we have a new participant starting Ekso walking from upcoming Tuesday; so would it be possible to begin the process of getting us a loaner device as starting next week it will be difficult for us to have back to back Ekso session? | Loaner sent to their facilities. We changed the shaft and ball screw as well as installing a new encoder in case of damage from the moving shaft. Work on unit was done via SWo-994 | The knee motor shaft nut did not seat properly due to supplier not threading all the way to the base. This caused the shaft to have approximately .015" of play.
2081 | The L knee motor was making funny noises | closed | kmettler | Foot Sensor | fixed | 2015-02-02 23:50 | 2015-02-02 | 150491 | 106 | RIC | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | Where: RIC SN:021311002 What: The Ekso GT began acting strangely with a patient today. The patient is a R hemi and we train in R. Affected. The L knee motor was making funny noises (almost a whirring) and bouncing when he was in L stance. We turned it off and restarted to see if that would correct it and it made no difference. I also triple checked hardware and software settings. I just got into the Ekso to see what was going on and it did the same to me. I have attached a video. The bouncing doesn't start till I weight shift onto the L. No errors appear on the Ekso. The patient typically walks over 2000 steps CGA and today he required mod A consistently. We are concerned about safety using the device to train tomorrow. What do you think is going on and what should we do? | New foot sensors replaced on unit via SWO-981. | After walking in unit found sensors on unit were hyper sensisitive and would quickly transition from reading on to off resulting in the situation the customer saw. Sensors will be given to Chris in engineering for further evaluation.
2080 | Ekso threw Error stating batteries weren't connected | closed | kmettler | Torso | fixed | 2015-02-02 16:27 | 2015-01-26 | 33860 | 161 | Glenrose - Canada | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Solution Not Implemented | Customer: Glenrose Unit # 161 Issue: After walking a patient and in a seated position the PT got an error stating the left battery was disconnected, after acknowledging the error they got the same issue on the R side. The Ekso told them to reboot and they did, after reboot the problem has not been replicated. Melissa is there tomorrow for advanced training, Edmundo spoke with Mark and Eric and they believe it could be ESD especially being in Canada. | Add star washers for grounding when unit returns for retrofit. Device is safe for continued use. Will confirm completion of of fix defined by CO-11171 with work done in the field on SWO-1251 | Star washers not installed yet.
2079 | Lower Shank gripping issue | closed | kmettler | Adjustments - Leg | fixed | 2015-01-31 00:16 | 2015-01-29 | 3123 | 151 | Glostrup - Denmark | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Where: Glostrup (Denmark) Upgrade SN: 021407006 What: The problem with lower shank gripping issue continous. We have a new person walking and the right lower leg rotates, he has no spasticity. It¬¥s the right side again so there must be something wrong with the shank, tube or screwes? We have changed the screwes and the t-wrench tool is a new calibrated one. What to do? - Lower Shank is not rotating with respect to the lower shank mount | Had customer install extra washers on the screws which control the length for the lower shank. | The screws which are used for the lower shank clamp can be at times too large due to tolerance issues, which in turn decrease the available clamping force.
2071 | Burnt smell from motor during patient session | closed | kmettler | Customer Use | fixed | 2015-01-23 23:30 | 2015-01-22 | 15000 | 155 | Linkoping - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Training | The first time the smell was noted at 200 steps and now at 750 to a small amount of smell coming from the hip motor today, it seems to be with one patient more than the rest, I have attached photos of the motor temps after the 750 steps and the patient concerned to this email. | Josh communicated to the customer that there was no problem with the device, the burning smell was from some dust or lint. It was also communicated that there needed to be several clinical changes/ more training for patient settings. | The customer reported a smell at 200 steps and now at 750 and then a small amount of smell coming from the hip motor today.
2067 | Batteries draining quickly @ TLM | closed | rrea | Battery | fixed | 2015-01-23 21:42 | 2015-01-19 | 100000 | 34 | TLM - Mexico | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Customer: TLM Unit: 34 Issue: see Below. Hi Josh, How are you? I had a new report of our Ekso, it looks that the life time battery is almost over. According to the users they full charge the batteries, but when plugging them to the device it says only 95 or 98% and it last around 20 to 30 minutes then it is over. The replacement of the batteries is under warranty or under quotation? Please let me know. | Sent new batteries with SWO-1043 | Battery was draining too quickly Additional analysis to be done with Ticket #2384 then updated here.
2050 | Mark Pollock broke his unit's upper shank | closed | jwitherspoon | Lower Leg | fixed | 2015-01-20 22:11 | 2015-01-19 | 238000 | 20 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Solution Not Implemented | . I have walked in my loaner about 10 or 12 times max since I got it. And, today the lower left leg of the Ekso came apart. Simon saw a couple of sheared nuts. | Both Upper shanks were rebuilt to our manufacturing standards with CO-10813 applied. Work was done on SWO-995 | CO-10813 was not applied to this unit before sending it out.
2049 | Feature's on device not configured for cusotmer | closed | kmettler | Shipping / Handling / Packaging | fixed | 2015-01-19 18:00 | 2015-01-16 | 3186 | 170 | FLM - Spain | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Process Not Available | FLM has used their new GT for the First Time and they only have First Step Activated. Back then we did the training with a loaner Unit as it wasn‚Äôt released from customs in time. What is the plan to activate all the other features? | Adrian arrived at the customer and turned on the correct user settings. | Do not have a detailed process for turning on features on a device when it is delayed due to customs issues. CAPA to resolve this moving forward.
2048 | Ekso pins bent when PT was putting on battery | closed | ckealoha | Torso | fixed | 2015-01-19 17:42 | 2015-01-16 | 11522 | 172 | VA Dallas | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | Where: Dallas VA. GT #? SN 021411001 Steps: 11522 What: one of the PTs tried to place the battery upside down and the post bent a few pins to dock the battery. Please see picture below. I have the PTs calling to trial our CR line process. Please let me know if this is something I can help fix remotely. | pins bent back. See CAPA #2015 to address long term fix. | See CAPA #2015
2047 | Batteries draining quickly during use | closed | kmettler | Battery | fixed | 2015-01-19 17:36 | 2015-01-13 | 330000 | 28 | Mt. Sinai | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Design | Customer: Mt. Sinai Unit # 28 Step count: 330,000 The service they are referring to below is their retrofit. Issue: We are having some battery issues. Both set of batteries, the old one and the new replacement set we received, discharges really quickly. For instance, we would start a session with fully charged set (usually around 98 to 99%) and by the time we do a lap, which is about 325 ft, the battery level would go down to 50-55%. This really started happening after we got the service done. Initially I thought Ekso motors might be working hard due to new patients we were having who just started walking in Ekso. However, this seems to be happening with both sets of batteries. We had someone who walks with minimum to contact guard assistance, and the battery level would still go down as I described above. Would you know why this might be happening? | Replacement batteries being sent out on SWO-969 | The battery level would go down to 50-55% quickly.
2045 | Torso shroud loose on unit | closed | ckealoha | Torso | fixed | 2015-01-12 17:58 | 2015-01-08 | 48508 | 137 | VA Palo Alto | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Solution Not Implemented | Torso shroud is coming loose on unit. | It is currently unknown why this change was not implemented. Upon in field inspection of the device the torso shroud was re-secured to the torso frame. This work was initially closed out per SWO-728. | Upon inspection of the device it was found that fix which firmly secures the torso shroud to the torso was not implemented.
2044 | Device threw #135 error followed by Error #136 putting it into safe mode | closed | kmettler | Motor | fixed | 2015-01-12 17:35 | 2015-01-07 | 19744 | 17 | Penn Partners | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single |  |  | Customer Use | Today at GSPP we got someone into the device who had limited ROM at their left hip. They had a restriction of 15 degrees, we stood them at 10 degrees of hip flexion and them started evaluation mode. While in evaluation mode (after only standing for about 5 minutes, we didn‚Äôt take any steps) we received error code #135 left hip. About 10 seconds later we received error code #136 and had to sit the patient via safe mode. The left hip motor was EXTREMLY hot, to the point it could burn someone. I know that with the 1.1 device there are different hip motors that‚Äôs don‚Äôt manage heat as well. I wanted to let the team know how quickly this team received this error, and the safety risk associated with it going to safe mode so quickly. Ultimately this patient was not appropriate to use the device at this time Let me know if any other information is needed. | Clinical team informed customer that the issue was due to the patient not being within the inclusion criteria. | Patient did not meet the inclusion criteria.
2043 | Broken Walker- Plastic Piece | closed | kmettler | Cane / Walker / Crutch | fixed | 2015-01-07 21:34 | 2015-01-07 | 18500 | 17 | Penn Partners | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Materials | Today the RW at GSPP broke. I have attached a picture. It is still functional, but has the issue with the pin not locking in place. Since they have the old RW, would it be possible to send them the the updated version? | Sent replacement walker via SWO-00959. | Pin on the walker not locking into place.
2042 | Battery not fully charging | closed | lmcneil | Battery | fixed | 2015-01-07 19:16 | 2015-01-07 | 18500 | 17 | Penn Partners | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | Here at Good Shepherd Penn partners and one of the batteries will only charge up to 89% at maximum. We have charges it fully on 2 separate occasions, and this has happened each time. The other batteries are charging fully . I have attached a picture of the serial number. | Returned batteries given to Rochelle. RMA 2042. Batteries replaced with SWO-00957. | Additional analysis to be done with Ticket #2384 then updated here.
2041 | Broken Walker | closed | jwitherspoon | Cane / Walker / Crutch | fixed | 2015-01-07 18:42 | 2015-01-06 | 180785 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Wear | PT from Good Shepherd reported a broken walker, the plastic stop has broken off and the walker can swing open. | Replacement sent out via SWO-956. | Walker had become damaged from use.
2040 | 1054 Error occured during setup | closed | jwitherspoon | Electrical Interconnect | fixed | 2015-01-06 22:03 | 2014-12-30 | 11286 | 173 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Workmanship | Customer unit came back to HQ due to a failure in Mid-December. It was fixed here @ HQ, upon opening the box and using the Ekso with their first patient they received an error 1054 RK. I had them reboot the Ekso and the error occurred again. They had a loaner unit already on site so they‚Äôve been using this for the past few days. This occurred on 12/30. | Replaced knee board, brake, and main cable harness. Work done per SWO-960. | A pinched brake wire was found under the CCA shield.This may have caused a short in the board so we changed both as well as the main cable (knee to torso) to be sure we caught the issue.
2039 | Torso Shroud coming loose | closed | kmettler | Torso | fixed | 2015-01-06 19:24 | 2015-01-06 | 79355 | 137 | VA Palo Alto | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Solution Not Implemented | Customer: Palo Alto Unit # 137 Issue: Customer called to report their Torso Shroud is loose and they also had been prompted to perform a joint calibration. Customer was walked through the leg calibration and it passed. Al is going to Palo Alto now to fix the shroud. | Unit serviced in field via SWO-955 by properly securing torso shroud with washers. | Upon inspection of the device it was found that fix which firmly secures the torso shroud to the torso was not implemented. See trac #2045
2038 | During patient session R hip joint cap failed | closed | ebabaran | Actuators | fixed | 2015-01-06 19:16 | 2015-01-02 | 180042 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | Attached are the pics of the R Hip joint failure on our 3# GT. It was reported that there was not a precipitating event, error or major sounds prior to the "snapping" sound that I was summoned to. The pin holding the "bearing" was sticking out 3/8 - 1/2" from the joint with smaller roller bearings visible and falling out of the joint. We then immediately sat the patient and got him out of the machine. The pin then completely fell out when we abducted the hip to get him out with more bearings and pieces with it. I have them all collected and saved for your team to look over and will send with machine. The patient was fine but it was seconds from coming apart. Let me know if you need anything else or if we can break anything else! Thanks for your help. | As per SWO-00958.I replaced both joint caps (inside/outside), the joint shaft, the spacer,and both bearings. Torqued to 22 in-lbs. Replaced all inside caps and re-torqued to 22 in-lbs on the other three joints as well. | Torque set too high for usage.
2036 | Collar is loose on lower leg | closed | kmettler | Adjustments - Leg | fixed | 2014-12-19 21:00 | 2014-12-17 | 5227 | 175 | St Joseph's Hospital | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Customer Training | What happened: it was noted that the Left lower piece that houses the 2 bolts to tighten/lock the lower leg length and labeled with 2x was slightly loose/I could jiggle it slightly at value 39 despite tight bolts. When the PTs readjusted the leg length to 41 it was completely fine. Any ideas of why this could change or is there something that could have been loosened and now tight again? Here is the picture of the piece I'm describing. This is on the FAQ and is normal behavior. Customer: St. Josheph's Barrow. | This issue falls under our spec and is within our normal tolerances. | It was noted that the Left lower piece that houses the 2 bolts to tighten/lock the lower leg length and labeled with 2x was slightly loose/I could jiggle it slightly at value 39 despite tight bolts. When the PTs readjusted the leg length to 41 it was completely fine. Any ideas of why this could change or is there something that could have been loosened and now tight again? Here is the picture of the piece I'm describing.
2035 | #21 Error RK occurs every few hundred steps | closed | kmettler | Encoder / Potentiometer | fixed | 2014-12-19 20:52 | 2014-12-15 | 10190 | 173 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Process Not Available | Customer: Warm Springs Steps: 10,190 Unit #: 173 Issue: Every few hundred steps the Ekso is throwing a 21 Error RK. This is intermittent but has happened 4-5 times today across the board with all patient types. It‚Äôs occurred during standing, sitting, mid walk etc. They are doing advanced training tomorrow with Shantel so we need a speedy plan of action to get them fully functional. | Unit will be fixed by changing and properly mounting encoder per SWO-950.A new encoder is installed now properly and the issue did not return. System fix will be released via CO-11263 | The center bore of the encoder hub was not concentric with the housing, this was causing the side to rub and generate friction.
2033 | Batteries draining quicker usual | closed | rrea | Battery | fixed | 2014-12-15 17:53 | 2014-12-15 |  | 53 | To Walk Again - Belgium | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Regarding the batteries I asked additional information and they‚Äôve confirmed: The problem occurs with both sets. Sometimes we have an error message that only one battery is ‚Äòlow of charge‚Äô (always the right one). The sets are always regularly swapped, as teached by Dennis. On average the batteries last 45 minutes of standing time (35 ‚Äì 40 minutes of walking). | New batteries shipped to customer via SWO-943. | Batteries still require further analysis prior to determining root cause for failure. Additional analysis to be done with Ticket #2384 then updated here.
2031 | Step not completing during session | closed | kmettler | Foot Sensor | fixed | 2014-12-15 17:46 | 2014-12-14 | 260000 | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | our EKSO makes again (! ) problems. Tobias describes it down pretty well. The step left is not finished , but dropped the foot before. We can therefore use limited the device. In pts. With a lot of support it is not safe to handle , among fit it comes just when it is predicted that the step is not done completely . It is the first time today and then occurred directly at 3 Pat . Consecutively . The Pat . Had been to several times without problems in the unit. Even when testing with a colleague was a problem . It can therefore be assumed that it occurs again on Monday . On the device itself is not unusual to see (screws , cables, soles , etc.). When the unit is turned off passive movement is normal. Switched on the left H√ºftmotor hears slightly differently than the right. But this can be deceiving. Error Messages There were no lt . Tobias . | Guided the customer through re-calibration per SWO-1476 | Foot sensor required re-calibration.
2029 | Foot Sensor broke during walking session | closed | kmettler | Foot Sensor | fixed | 2014-12-15 17:37 | 2014-12-09 | 269795 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Please see below email from Technomex regarding broken right foot sensor (now replaced). SN: 021403006 - #125 Steps: 269795 I have asked them to confirm whether they have any spare foot sensors of if they used them all. If they used them all perhaps we should send them a few in case this happens again. | Sent new foot sensors to customer. Replaced foot sensors superceded by work completed via SWO-1408. | The Foot Sensor broke during walking session. Sent from EU stock.
2027 | R leg swings quickly w/ one patient | closed | kmettler | Customer Use | fixed | 2014-12-15 17:16 | 2014-12-08 |  | 164 | VA Minneapolis | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Training | Customer VA Minneapolis Unit# 164 Please see below video‚Äôs, PT says with this patient his R leg swings too quickly. | Melisa was able to walk through the through some clinical suggestions. | The PT complained The R leg swings too quickly. This was a training issue
2024 | Error 21 LH Intermittent | closed | apoultney | Encoder / Potentiometer | fixed | 2014-12-08 17:50 | 2014-12-05 | 62000 | 131 | ULM - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner | Minor |  | Determined Unknown | This morning I spoke with Sebastian and went through the leg test but the Ekso would not clear the error code. The attached email gives details of the plan. I will be on site Monday to see if a repair can be done to their Ekso, then fly back Tuesday to the workshop for the collection of Dales Ekso and to receive his loaner. The customer has not yet sent the issue report details. Error 21 LH No Data now Adrian there today. | Encoder cable and hip external replaced on unit in the field per SWO-936 | Parts not evaluated - We need to have cable returned for further investigation of root cause
2023 | Lower leg not holding in alignment guide | closed | kmettler | Adjustments - Leg | fixed | 2014-12-08 17:43 | 2014-12-03 |  | 156 | Kerpape - France | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | The photos seem to show that the foot is not holding in the shanks alignment guide. | Customer installed additional washers. | Due to tolerance issues with screws which clamp the lower shank, at times they are too long and decrease the clamping force.
2014 | Torso Test Req'd | closed | ckealoha | Software | fixed | 2014-12-01 17:54 | 2014-12-01 | 54000 | 138 | Nottwil - Switzerland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Solution Not Implemented | They switched it on today and it came up with a torso error (they didn‚Äôt note down the error number). On restarting the same happened again and Ekso is only going into safe mode stating ‚Äútorso test required‚Äù | Software to be updated per CO-11173. Listed on MFG Spreadsheet to be done at next service. | IMU drifts and error test does not adequately check. This is released in 2.0.6 but this device is not updated.
2013 | Delay when using ProStep+ | closed | kmettler | Foot Sensor | fixed | 2014-12-01 17:51 | 2014-12-01 | 3500 | 152 | Duisburg - Germany | Ralf Lanzmich | ralf.lanzmich@bgu-duisburg.de | +49.203.76883613 | Germany |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Unassigned |  | Design | I just got an email from the PT in Duisburg ‚Äì they seem to have a problem with their ProStep+ as in there is a delay in the the stepping of the right leg. I have no further info but have asked them for some more detail. Adrian ‚Äì can you please contact Ralf Lanzmich the head PT by email Ralf.Lanzmich@bgu-duisburg.de to arrange a time to call him so you can run through some diagnostics. I will be there next week for advanced training but our time then is limited and we‚Äôll need a fully functioning Ekso so we can cover everything in the training schedule Adrian determined that foot sensor calibration was needed but customer didn't want to do the calibration. | Adrian walked customer through foot calibration and fixed the issue. | Foot cal needed.
2012 | Ekso turned off during session | closed | rsmith | Electrical Interconnect | fixed | 2014-12-01 17:48 | 2014-11-28 | 93000 | 116 | Good Shepherd - IP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Minor |  | Design | Customer: Good Shepherd Unit #: see below. Answer needed Monday morning. Issue: while walking a patient, they were in turn mode and the Ekso powered off. Batteries had a charge level in the 90s. They sat the patient using a long legged sit. Patient was on tether, they suspect the power button was possibly hit by the clasp. They can't remember if the power button was depressed or flush after they sat patient. They turned it back on and everything worked fine. They walked an able bodied person w/ no issue. I told them not to walk any further patients until we can look at their data Monday morning. Here is the info you need to troubleshoot the incident from earlier today when the Ekso shut off. Our SN is 021402005 (Ekso 3) The pt's ID number is 1086 and it happened on 11/28 around 10:15ish The device step count when it happened was between 89330 and 89574. | Properly grounded power switch per CO-11171 via SWO-859 | Unit was not properly grounded through the power switch.
2007 | After batteries are changed the R leg is taking a short step | closed | kmettler | System | wontfix | 2014-12-01 17:25 | 2014-11-24 | 374500 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Determined Unknown | I have walked in my Ekso 3 times since it‚Äôs return from America. And, the same problem is emerging again just like before. After the battery change the right leg is taking a short step and sorting itself out over time. This is exactly what happened since the devices return from America in May and it got worse and worse until it wasn‚Äôt safe to walk in. As you know I was out of action from late Dec into Jan this year with a device problem. Then from early May until mid June. And then from mid/late Aug until now. So, I‚Äôve missed months of training in my own device. | No resolution implemented. Unit has been taken out of service. | We were not able to reproduce issue on unit. Unit has since been taken out of use from system.
2005 | Customer has one bad battery | closed | kmettler | Battery | fixed | 2014-11-24 17:54 | 2014-11-18 | 18000 | 62 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Wear | Good morning Adrian, if he put the Right Battery the device works. When he put the Left Battery, the device doesn't switch on. So I think is necessary replace just one battery. So, how we have to proceed? Lucca- EMAC customer | SWO-00934 was generated and 2 replacement batteries were sent to the customer. | Over time the batteries will naturally degrade.
2004 | Broken left ankle plantarflexion stop cracked | closed | kmettler | Ankle | fixed | 2014-11-24 17:49 | 2014-11-20 | 320000 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 10 - Field Service & Loaner |  |  | Design | Dear Adrian. During trainig with patient the ankle joint got loose. Now it has play of approx 15 degs. I opened the left ankle and as you can see there is a tear in an elemwnt that was supposed to hold the ankle in place i point it with the finger on the pic. Also attached a pics of sn of the unit and step count. Pkease advvise ASAP if we can work and the part shipment cause we have patients scheduled daily! | The fractured shank mount was replaced with the new thick wall version per SWO-1090. | Previous design allowed for a stress concentration to form on the plantar-flexion wall, which in turn caused the material to fatigue and crack.
2000 | R Knee buckled during patient session | closed | kmettler | Electrical Interconnect | fixed | 2014-11-24 17:35 | 2014-11-18 | 20037 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner | Major |  | Workmanship | Customer: Kessler Unit #: 136 Steps: 20,037 Please see below issue. I don‚Äôt have any further information as of now, I have a follow up email into customer to obtain more details. Thanks Hi Gail, Buffy had an issue with the Ekso GT today while using it for her 2:30 patient. It has to do with the same issue with the RT knee. She said that the knee hesitated to lock and worsened over time. The motor became extremely hot and would not lock anymore. She was forced to do an emergency sit. They tried to passively move the joint, which seemed fine. They re-donned the device, tried to do a sit to stand, without success. Did a max assist sit to stand, but the right knee would not lock out. She passively extended the knee to lock it, and decided to end the session for safety. She‚Äôs asking if we can forward this information to the engineers at Ekso for them to review (I think Josh?) ? She is hoping to have this problem fixed by Thursday for the patient‚Äôs next session. Thanks! | Issue resolved on SWO-722 by replacing brake cable assembly | Pinched break cable.
1999 | R arrow pressed for hold to walk- UI screen didn't change | closed | kmettler | UI - Controller | fixed | 2014-11-24 17:33 | 2014-11-18 | 41178 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | Confluence Health- Central Washington Hospital 41,178 Ekso GT with VAS At 11:45am PST the PT pressed the right arrow to hold for Walk. The sound rang, notifying that we entered walk mode. However, Remote screen still was in regular screen for Stand Still mode. No other missed actions occurred. We were walking in Pro Step, Bil Max. | Known Issue & not critical enough to replace. Will go away when unit updated to 2.0.6. Replacement UI sent to customer as precautionary step via SWO-917. DSP was later addressedvia SWO-1138. Link to this SWO only shown in notes as was already closed in omnify when this link was made. | Trac Ticket #1667 describes root cause.
1998 | Both toe guards have broken off Ekso | closed | kmettler | Foot Assy | fixed | 2014-11-24 17:27 | 2014-11-18 | 47000 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Customer: Burke Unit #: 142 Issue: See below pics and description. Response: I have Chris‚Äô instructions on how to remove and replace toe guards. I‚Äôll get parts from MFG. Josh, Hope all is well. Amy and I noticed that a piece of plastic broke off the foot during a session last week. During closer inspection we noticed that on BOTH feet the plastic in the forefoot was broken off. Please see attached photo. Also, is it normal for the knee joint to "click" when flexion to extension in "Fixed" mode? | SWO-00916 was generated and the toe guards were fixed in the field. | Known issue.
1996 | Step Abort Errors during session | closed | cmeadows | Foot Sensor | fixed | 2014-11-21 19:35 | 2014-11-21 | 135000 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Our ekso #3 is experiencing multiple step aborts. From 9-10 today there was about 8-10 step aborts. Could this be related to the sensor issue? or a different issue? | Replacement sensors sent out and replace via SWO-923 | Bad sensors returned in house for further evaluation Right sensor failed verification test on foot stand. Right heel sensor had shifted position within capsole and overlapped toe sensor. Electrical tape became bunched from use and separated from top and bottom layers of sensor stackup. See attached pictures. Left sensor passed verification on foot test stand.
1994 | Thigh pads replaced with foam | closed | ajulin | Soft Goods - Upper Leg / Pelvis | fixed | 2014-11-17 18:00 | 2014-11-14 | 3123 | 151 | Glostrup - Denmark | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Followed | Since we have had 3 people who have received wounds of the thigh support so we have decided to remove the padding, with the sharp edge, and replace it with foam. | CO-11225 covers hook material and seam with trim. Sent replacements to customer with SWO-00915 | Insufficient space between hook tape and pilot provided in thigh pad design. Insufficient testing done on able-bodied pilots before design released.
1990 | Step won't initiate when it should | closed | kmettler | Foot Sensor | fixed | 2014-11-17 17:31 | 2014-11-11 | 179877 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Customer: Madonna Unit #119 Steps: 177,000 Katie is pulling data to determine possible failure. It‚Äôs Matt again. Sorry. We seem to have a new issue developing with Ekso. There is a video attached. Basically, the left leg is getting stuck and unable to move forward for several seconds; we cue the patient to just keep trying and eventually it steps forward. It has happened in Left Affected mode, and also in Bilateral mode with only the left leg getting ‚Äústuck.‚Äù It has happened with multiple patients, however not every patient. Typically it can happen anywhere from 5-15x per session. I checked the sensors as I did before. They all seem to be working ok; if I push on them the pressure readings increase in technician mode. | Walked customer through foot calibration. Sent replacement foot sensors with SWO-00857 | Known issue.
1989 | Fluttering sound coming from knee brake | closed | ckealoha | Brake / Release / 9V Battery | fixed | 2014-11-17 17:22 | 2014-11-10 |  | 165 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Unassigned |  | Workmanship | Hi Team, Please see the attached video. There is a winding up noise coming from the right hip that they have not experienced before so they canceled the patient they had left today and wanted to make sure this is "normal". No more patients are scheduled today. No errors have occurred and it worked fine on Friday after Chad left. | Cleaned brake hub and shaft with Acetone and rebuilt unit according to MPI. Fixed in house via SWO-918 | It was determined that uncured 641 worked between brake hub and base causing the brake to seize up. The root cause was determined to be excess 641 during hub installation.
1974 | Stiff thigh adjustment | closed | ckealoha | Adjustments - Leg | fixed | 2014-11-06 19:13 | 2014-10-28 | 31509 | 137 | VA Palo Alto | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Palo Alto VA brought their Ekso to HQ for a software upgrade and also mentioned their thigh was becoming harder to adjust over time. MFG looked into this issue and confirmed their adjustment needed to be fixed by reaming thigh structure bore. | SWO-00728 reamed bores and cleaned up buildup. The adjustment was smooth after service. Fixed with CO-11217 | Build up in tubes.
1967 | While charging battery LED won't turn amber | closed | kmettler | Battery - Charger | fixed | 2014-11-03 17:37 | 2014-10-31 | 14629 | 141 | Ancona - Italy | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | I‚Äôm writing to inform you about a little problem with the Ekso #141 in Ancona. They have a problem with the charge battery. The orange led doesn‚Äôt switch on when you attach the battery for charge. The battery is charged normally, but I think could be replace the battery charge. Because I‚Äôm afraid that it could not stop the charge when the battery is full. Moreover, next week I‚Äôll replace them the knee brake battery on the left leg, because it‚Äôs down. Let me know what do you think about the charger or if you need other information. From Rochelle 11/3: I think we should bring back the charger, battery in question, and it's matting battery (1a/1b or 2a/2b). | SWO-00894 was generated and a new charger was sent to the customer. NO battery was sent- only a charger. | did not get charger back.
1965 | Plastic DDL guard prevented locking of legs | closed | ckealoha | Don-Doff Link | fixed | 2014-11-03 17:27 | 2014-10-29 | 263000 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Training | Yesterday I had a call from Jakub at Technomex, he was saying he had an issue with the leg abduction, so I asked for a photo of the fault and he sent the attached photo of a DDL, I think this is the issue as I can see the locking pin is sticking out more than it should if it is meant to be locked. As far as a fix the DDL might have something in the joint stopping it from locking, do you have anything else we can think of? I am still waiting on more information from Jakub about the Ekso issue details. I will follow-up with him tomorrow if he is available. Customer reported the plastic guard was pinched and prevented the legs from locking. They were able to find this issue and move the plastic out of the way so it no longer interfered with the closure of the DDL. | Customer moved the DDL skirt out of the DDL locking mechanism. | Known issue. Part of the design.
1964 | Upon Boot Ekso threw Errors #133 followed by #121 | closed | kmettler | Electrical Interconnect | fixed | 2014-11-03 17:12 | 2014-10-28 | 19776 | 126 | VA Boston | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Workmanship | Customer: VA Boston No Data Response Needed: Only respond if you have suggestions for Greg who will be visiting their site on Thursday. Issue: When turning on the Ekso they got Error #133 followed by #121. This same error sequence happens every time they turn the device on. Greg's Summary of field fix 11/3: For the Boston VA failure I tried several different methods before changing the joint board. 1. I opened the covers and saw the motor phases were pinched together. I checked the cables and rerouted and the issue returned. 2. I ran a new main cable harness from knee to torso and the issue was still present. 3. I reprogrammed the joint board and routed the wires properly and the issue went away. We calibrated, cycle tested, airwalked, and finally walked with PT and no issues. Then after sitting down the error messages returned. 4. I changed the joint board and programmed. Once the unit was put back together I calibrated, cycle tested, air walked, and walked the PT with no issues. To validate again we walked the next day and no issues returned. I am not 100% that this is the proper fix. If the errors return at any point I would suggest we change out the motor itself. Greg Customer Name: Boston VA Location of Fix: Boston Unit S/N:021404001 Step Count On Device: Reason For Visit: Error #133, Error #121 Issue#1 Able to replicate? Yes Able to determine root cause?: Joint board failure Resolution#1: Installed new joint board. Issue#2 Able to replicate? N/A Able to determine root cause?: N/A Resolution#2: N/A What other work was completed on the unit (Open updates, SWO‚Äôs, Unit Maintenance Spreadsheet, TRAC, etc)? None Cusory checks completed: Were you able to check their phone connection and verify it works? Yes, checked the phone and it works fine and is collecting data. Phone works If not were you able to fix the connection? Yes/No How? Charge Phone/Reset Phone/Connect to WiFi/Replace phone/Replace SIM Card OK Were all the adjustments functioning properly? Yes Foot Sensors in good working order? Yes Soft Goods all in place and not requiring replacement? Yes If so which? OK, not : unit already had new thigh straps. | I changed the joint board and programmed. Once the unit was put back together I calibrated, cycle tested, air walked, and walked the PT with no issues. To validate again we walked the next day and no issues returned. Work completed per SWO-00887. | The motor phase wires were smashed together with the barrel connectors pressing against the EMI shield/joint board.
1958 | Battery charger pin bent | closed | kmettler | Battery - Charger | fixed | 2014-10-29 20:23 | 2014-10-29 | 46400 | 138 | Nottwil - Switzerland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Customer Use | Customer: Swiss Paraplegic Centre Nottwil SN: 021405005 Step Count: TBC I have received a call from a physio in Notwill reporting error #9 - mismatched batteries. Whilst over the phone I have asked her to test different batteries and she took pictures of the user interface. It seems that two of the batteries weren‚Äôt charged enough and there was a mismatch twice. But then mixing up the two set of batteries the issue didn‚Äôt come up. So I‚Äôve asked her to check if the pins on the charger are all straight and she advised that actually they had difficulties in placing the batteries on one of the charger. It seems one of the pins is slightly bent. Please see pictures attached. I asked her to charge again all batteries using the charger with the straight pins and try again if the issue persists. Please let me know if there is anything else we should ask to test. | Sent new battery charger via SWO-00891 | There were difficulties placing the battery on the charger.
1955 | Knee brake not engaging | closed | jwitherspoon | Electrical Interconnect | fixed | 2014-10-27 16:56 | 2014-10-23 | 16164 | 136 | Kessler | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Workmanship | Customer: Kessler Unit # 136 Steps: 16164 Response needed soon. Issue: I got this video from a PT at Kessler. They told me the knee brake wasn‚Äôt engaging until very late if at all. They have been walking with the same patient and have had no issue's, this suddenly occurred during their last session. Any suggestions on why this is happening and what we need to communicate to the customer to determine action plan? From Katie 10/23: Agreed. The problem is on the Right Knee. In VA, the current is still pushing the knee towards extension (not at max current, but 7-10 amps). However, the knee isn't really reaching the point the brake kicks on until late double stance. Towards the end they switched to max assist... the Right knee is holding around 22 deg and the current is railed at 40 amps... and it's just not moving... which makes no sense and makes me worried that there's something mechanically wrong. Can we have them move through the ROM passively? And then able bodied walk if that seems ok? It gets stuck as high as 33 degrees with current railed... so either there is something in the actuation mechanism jamming in extension or the brake is catching then? In the stand, they got to 2 deg, so it doesn't seem like a ROM issue for the patient. I would recommend not using with a patient until we figure this out. | After rerouting the wiring we were able to walk the unit with the PT and it worked fine with no issues. This Trac ticket completed with SWO-00888. | Greg removed the knee covers to check out the issue and found the temp sensor wire was pinched. I rerouted the wiring so the cables were not pinched and the error did not reoccur.
1942 | L Leg going into extension before heel ground | closed | kmettler | Customer Use | fixed | 2014-10-20 16:36 | 2014-10-20 | 187000 | 26 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Materials | We were walking on Friday and about 45 mins in the left leg of the device started acting up. It was like it was straightening too fast so was at full extension before the heel even hit the ground. There are two videos attached and it is a bit clearer on the side angle but it was strong enough to make walking difficult. Also it was noticeable on the loaner that Mark's right step was fairly even with the left. On his own device, the right foot nearly always ended up with the heel at about the level of the left toes. This had been identified before but the physio advice was that it was probably to do with the weight shift. Given the apparent power issues with the right motor, could it be that the problem had been there for a long while? | Foot sensor replaced in the field and verified no longer an issue via SWO-1376. | Failing foot sensor was not sensing on/off response properly
1927 | Stiff DDL- hard to open legs | closed | kmettler | Don-Doff Link | fixed | 2014-10-08 17:54 | 2014-10-07 | 100000 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Unassigned |  | Process Not Available | From Maria 10/7: Hi Adrian, like anticipate by phone, we have a problem with the Ekso 021406001 in Villa Beretta, Costa Masnaga. They push the lever in the Don Doff Link to unlock the leg, but to move the leg is very hard. The lock and unlock seems to work but to move the leg until the 80/90 degrees is hard, like it is stuck. This happen in both legs. Can they continue to use the device or they must stop it? There is something that they can check or do quickly? In case, can you move from them on Friday afternoon after your Italy tour? Let me know if you need other information. Thank you very much for your support. This is a known issue due to lack of MPI GPL grease was not installed until unit 142. | Unit to be fixed in field 10/23 via SWO-855. The tools needed for this fix have been damaged and the unit was not fixed on 10/23. Adrian returned with new tool and fixed on 12/2. | Earlier builds did not require additional lubrication within the DDL. CO-11155 adds in the required grease to prevent and correct this issue.
1926 | Duplicate- Torso Test Req'd | closed | kmettler | Software | duplicate | 2014-10-08 17:07 | 2014-10-06 | 13889 | 154 | Confluence Health | Jeremy Coble | Jeremy.Coble@confluencehealth.org | 509-741-7500 | USA |  | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single | Unassigned |  | Design | Customer: Confluence Health Unit # 154 Steps: 13,889 No customer response needed. Issue: PT called redline apox 8:30AM PST. They were setting up device, once they turned on the Ekso they received an error #80 ‚Äì Torso Test Req‚Äôd. Walked PT through the tech menu and they were successfully able to pass the test. | see ticket #1384 | Software bug- see ticket #1384
1921 | Walker doesn't lock into position | closed | kmettler | Cane / Walker / Crutch | fixed | 2014-10-01 17:51 | 2014-10-01 | 159632 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | The walker now swings around as the stop has been broken, also the trigger is getting stuck, so it will not lock in the walking position. | Replacement walker was shipped to customer via SWO-874 | Down rev version of walker was prone to failure
1919 | Error #27-LK Occured during patient Session | closed | ebabaran | Electrical Interconnect | fixed | 2014-10-01 17:39 | 2014-09-30 | 22965 | 124 | Renown | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Materials | Customer: Renown Unit # 124 Steps: 22,965 Answer needed: As soon as possible. Issue: PT just called the redline and reported getting a #27-LK error. This put them into safe mode, it‚Äôs a reoccurring error and won‚Äôt allow them to use the device. The first time this happened was at 3:30PST and they rebooted several times and the error kept happening. Is there any further details we need to get from the customer? | Fixed with SWO-00846. Intermittent failure was found in the torso to thigh external cable. Cable was replaced according to current manufacturing process. | Cable is known to fail when excessive strain is applied. It is believed that during normal use over time the cable became strained which in turn caused the failure.
1916 | Shroud is coming loose on unit | closed | ckealoha | Torso | fixed | 2014-10-01 17:27 | 2014-09-29 | 316542 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Customer: Good Shepherd OP Unit # 108 Steps: 316542 Response needed: ASAP Issue: I just wanted to bring it to your attention that the top component of the ekso(where the power button is) came loose. The same thing that happened to our inpatient unit. Let me know what Troy or myself can do to help you with this matter. I‚Äôve asked for a picture, I‚Äôll send it through when received. | Unit reinforced with new assembly method for Torso shroud as defined in CO-11171 and reworked in field per SWO-00860 | Securement wasn't sufficient
1914 | Ankle screws broken | closed | ajulin | Ankle | fixed | 2014-09-30 22:54 | 2014-09-25 | 411079 | 36 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Design | Customer: Gaylord Unit # 36 Steps: 411,079 Issue: PT walked a patient this morning around 11:00AM EST. When setting up the device for second patient they noticed the issue in the above picture. They don't know when this occurred. | Customer traded device for a new device. CO-10851 changed bolts to high strength Holo-Krome, and upgrades are being done fleet-wide. | Insufficient torque holding and strength on QTY 4 #8 screws.
1911 | Patient tripped while walking- PT guided patient and Ekso to the ground | closed | ajulin | Electrical Interconnect | fixed | 2014-09-24 17:50 | 2014-09-23 | 172600 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Workmanship | Reported by PT 9/22: Today we had an incident report- fall with patient in Ekso. Description- Hemiplegic patient was ambulation in Ekso with 2 physiotherapist. According one therapist the patient stumbled with one limb over the ground, lost his balance and together with therapist fell down. It was controlled fall with the help of 2 therapist. On the ground they released the patient, filled our incident report. Patient is withnout medical issue after the incident. On the Ekso there wasn¬¥t any error readings, even after the position on the floor. When we test the Ekso witnout patient after the accident we found flickering movement forward and backward on RT knee orthosis during the swing phase. There is no error message on the Ekso. Ekso is currently on hold, until the Rt knee orthesis will be resolved. Setting: RT leg Free, AF.L, Left leg- Max, Walk method- Pro, Train LATERAL. From Adrian 9/23: I have been able to talk to Jakub and we got a person in the Ekso to test it out. In first step there were no errors and it walked as it should. In Pro step max power to both legs gave no errors They then changed to the patients settings, right leg with assistance and left leg no assistance the jerky movement returned but no brake slipping issue was seen. Jakub had to go, so I sent him a video of accessing the Technical Menu as he has not been able to do this before. They have no Wi-Fi and have not been able to get the SIM to work for them. He said we can try again tomorrow.The issue is with the right hip is free but they can still feel tension when using pro step, I have asked for full details of the settings used | Unit fixed via SWO-635. Root cause for vibration seen by customer was phase cable that was not fully seated | Phase cable not fully seated.
1909 | Sensor Error #41 while walking w/ Patient | closed | kmettler | Foot Sensor | fixed | 2014-09-24 17:19 | 2014-09-23 | 17628 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Customer: Burke Unit #142 Steps: will advise Response needed: as quickly as possible, customer is down. They are connected to Wifi so we should be able to pull data with no issues. Issue: PT called redline. During patient session they received a sensor error #42. They sat patient and cleared error as instructed by manual. They stood patient again and error returned immediately. They rebooted several times but same issue appeared. Now the Ekso is throwing error right after boot and won't allow them to go any further. Ekso setup: bilateral, one leg in adapt and the other fixed at 75. While PT was using fast parameters and changing assist to 70 during session is when error first occurred. From Katie 9/23: This is a right toe pressure out of range issue. The signal seems to be floating high, so my assumption is still a short on the right toe sensor line. It's also somewhat intermittent in that it starts connected and then becomes disconnected and then seems to always be disconnected. | A new foot sensor was installed and calibrated without issue. Foot sensor was supplied per SWO-759. | Upon inspection of the unit, the unit would consistently reproduce error code 41 when pressing down on the heel while the toe was raised about the ground. On inspection of the heel of the foot sensor a small tear was found. A new foot sensor was installed and calibrated without issue. The root cause is believed to be the damaged foot sensor which was damaged during normal use.
1907 | Torn foot sensor | closed | cmeadows | Foot Sole | fixed | 2014-09-24 17:12 | 2014-09-23 | 166262 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Customer has torn a foot sensor. From Maria 9/24: The surface where they are walking the patients is just a ‚Äústandard floor that you can find in a clinic/hospital‚Äù. Please see attached pictures of the foot sole ‚Äì Jakub offered to send it to HQ along with the batteries so we can check it. The don‚Äôt have any spare foot sensors, they used them all. Barry has also stated they are trained on how to turn properly and do use the turning program. | CAPA #1259 will improve foot sensor robustness. Sensors sent out on SWO 910. | Wear issues Action - Continual use (walking and turning) on various surfaces Condition - Brittle adhesives may crack and easily propagate through material. Foot sensor material is not robust enough.
1906 | Battery not charging when connected to charger | closed | kmettler | Battery | fixed | 2014-09-24 17:06 | 2014-09-23 | 166262 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Distributor complained the battery won't charge while attached to charger. The light will only turn green and won't stay amber. | Sent new battery via SWO-1524 | Cell 8 in GT battery serial# 01b597ea160000fe has a weakened cell 8 and remains out of balance with the rest of the Cells in the battery pack.
1904 | Upper thigh hard to adjust | closed | kmettler | Adjustments - Leg | fixed | 2014-09-24 16:37 | 2014-09-24 | 133371 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Specification Not Correct | Customer: Madonna Unit # 119 Steps: 133,371 Answer needed: Katie has already pulled data- waiting for customer to return phone call. Issue: Vibrating from R hip area. This vibration has been occurring for the past 2 weeks with a variety of patients (Stroke, SCI) in Bilateral, Left & Right affected. The most recent time this happened it shook for 5 seconds so they decided to report the issue. Before this it was only happen briefly- 1 or 2 seconds. Katie was able to pull their data and it looks like their R Heel foot sensor isn‚Äôt sensing. I have called them back to walk them through some troubleshooting to see if a recalibration would fix the issue. | The right foot sensor was replaced with a new foot sensor and was calibrated without issue. The upper left thigh adjustment was altered for additional clearance for the lower don doff link. This alteration alleviated the difficult adjustment. Work was done per SWO-602. | Clearance bore in thigh structure - known issue Vibrating hip is due to damaged foot sensor.
1897 | Torso Test Req'd during setup of device | closed | kmettler | Software | fixed | 2014-09-17 17:44 | 2014-09-16 | 22254 | 105 | Warm Springs | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Design | Warm Springs- 30 day loaner Unit # 105 (internal device) Step Count: 22,254 Issue: When turning on the device an Torso Test Required was thrown. They called the redline and reported the issue. I walked them through the Torso Test and it passed with no further issues. Rebooted device and everything appears to be working appropriately. From Katie 9/16: Likely the known issue that was fixed in 2.0.6. | Walked customer through a torso calibration. | Known issue.
1895 | Thigh brace pad- created red mark on pilot | closed | kmettler | Soft Goods - Upper Leg / Pelvis | fixed | 2014-09-17 17:23 | 2014-09-16 | 3123 | 151 | Glostrup - Denmark | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Testing Insufficient | From PT 9/16: Hi Barry, We have a problem with padvelcro on the thigh support. I mean the pad against the thigh not the strap. The pad is attached to the thighbrace with Velcro. The Velcro side/edge is very sharp and the patients gets redmarks/wounds. Can you please send us a new one, it¬¥s only on the right side. Thanks Mats From Barry 9/17: Hi Team I followed up with Mats this morning and have attached a picture of the problem. The issue occurred as a result of the patient wearing shorts whilst undergoing EMG analysis. This is a single incident with just this patient. However they are adamant the pad is the causative factor. The picture, to me, also looks like the thigh strap may have been too tight. I have informed him we will be sending a new pad. | CO-11225 Update design of thigh pad 103596/103595 to add trim to cover up unfinished edge. Sent replacements to customer with SWO-00915 | This part was changed from 1/2" thick to 1/4" thick and was tested in house wearing pants. It was not tested on patients with shorts and if it had, we would have learned of the possible abrasion risk.
1892 | Noise coming from the hips during patient session | closed | kmettler | Adjustments - Hip | fixed | 2014-09-17 16:55 | 2014-09-15 | 226275 | 21 | PAC | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | Customer: PAC Unit #21 Steps: 226,275 Please let me know your thoughts so I can get back to the customer. Issue: I thought I would send you a video of one of our users using the Ekso last Thursday, 9/11 (let me know if you have trouble opening the file). Listen to the clicking sound the Ekso makes when the right leg takes a step forward. To me, I feel like the hip width adjustment may not be lining up correctly, so when the leg advances, something is rubbing against something else and making the noise. The reason why I feel it may be the hip width adjustment is because sometimes when this user stands up in the Ekso, the hip width screw (the vertical screw that we use the red socket wrench-type thing to tighten...sorry I don't know the correct terms for these things...) becomes loose somehow even though I tighten both areas (using the socket wrench thing and the torque wrench thing) twice each with the hips in full flexion. | Explained to customer noise is within expected tolerations. | Normal noise from unit- within expected tolerations.
1891 | Bent Pins on the Ekso battery plug | closed | kmettler | Battery - Connector | fixed | 2014-09-17 16:46 | 2014-09-11 | 9963 | 139 | Asepeyo - Madrid | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Customer Use | From PT 9/11: I have just received another email from Asepeyo: After talking to Josefina and watching the damaged part, I must tell you that the charger and the batteries are OK. The problem is actually in the pins of the eKso to which the battery connects (right side). Please let me know what we are going to do to solve this issue. From Rochelle 9/11: Attached you‚Äôll find a picture of the critical pins on the Ekso side of that connector. (Circled in blue) It looks like a pin may have pulled out of the charger then was lodged in the battery when they tried to plug in the battery. Have them check the charger please ÔÅä You‚Äôll also find a charger picture attached with the 2 pins that are critical for its use. Take a look at the attached drawing. If they can straighten the critical pins without breaking them, and then straighten/break off the other pins then they should be good to go. It looks like they may have bent some of the critical pins, so if they can get it up and operational without intervention, then I‚Äôd say it‚Äôs okay to keep using but we should put it on Adrien‚Äôs list to change next time he‚Äôs out that way. It‚Äôs a fairly simple swap. | Unit fixed in the field via SWO-803. | Customer mis-use and design allows for this failure mode to happen.
1886 | Stiff DDL release lever | closed | ebabaran | Don-Doff Link | fixed | 2014-09-10 17:55 | 2014-09-09 | 14251 | 133 | Uppsala - Sweden | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | I contact you regarding the things I talked about with you last week in Maastricht. 1. The lock for the hip on the right leg (DDL release) is hard to release, it has been like this from the start and needs to be taken care of. 2. Some stitching on the "shark fin" on the right leg is broken, see picture. (for some reason the pictures has turned upside down) Serial number om our Ekso: 021404008 Amount of steps: 14251 | Issue was dealt with by sending instructions and grease to customer (similar to dealing with issue via phone call). No SWO was generated as there were no stock parts required to remedy this issue. | Inadequate lubrication
1883 | No indicator sounds coming from Ekso | closed | ckealoha | Torso | fixed | 2014-09-10 17:26 | 2014-09-08 | 67413 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Customer Use | Good Shepherd Unit #132 Steps: 67,413 Issue: Another issue I'd like to bring to your attention is that the sounds on Ekso 3 are not working. We tried the lateral and forward target and they are not working. The beeps that indicate turning mode and stand still are also not working. Thoughts? | Upon investigation no issues could be found. The speakers in the torso shroud were replaced in case an intermittent cable issue was present. During inspection of the connector on the circuit board all connections appeared to be fully functional. This work was done per SWO-834. | Upon arrival the unit was functional. The unit was producing all sounds as intended.
1879 | Ekso threw LH 21 and then again LH 27 | closed | kmettler | Encoder / Potentiometer | fixed | 2014-09-10 16:56 | 2014-09-05 | 23377 | 131 | ULM - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Workmanship | Sebastian, the physio from RKU Ulm called me this morning with the following info: They received Ekso back from the US and unpacked it today. They walked for about 20 mins with PT s in it in all walk modes to ensure all is working well ‚Äì no issues. They then walked their first patient in ProStep and bilat. Adaptive and after about 25m Ekso threw LH 21 and then again LH 27. The patient was not particularly heavy but did have a lot of spasticity. They will continue to put Ekso through its paces by walking it themselves to see if they can elect the errors again, but it is a bit worrying that they only just received it back and have this issue straight away. | The encoder was reinstalled according to the current manufacturing standards. Work was done per SWO-1215. | Encoder set screw was installed improperly and caused intermittent free movement of the motor rotor relative to the encoder hub.
1878 | Multiple #137 Errors while walking with patients | closed | ajulin | Brake / Release / 9V Battery | fixed | 2014-09-10 16:31 | 2014-09-04 | 72250 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Good Shepherd Unit #132 Step Count: 72,250 Issue: While walking with 2 different patients they experienced 5 #137 errors. 4 occurred with patient #1 and 1 occurred with Patient #2. 1st Patient- walking in ProStep- Adaptive- 120lbs, 5‚Äô5 : Time 8:00AM-9:00AM 2nd Patient- Walking in ProStep- Adaptive- 190lbs, 6‚Äô0 : Time 9:00AM-10:00AM. Response needed: ASAP- PT‚Äôs are using their other device until we get back to them on why they are getting this error repeatedly. | Issue was resolved by sending compressed air as a tool to customer. No SWO was written for this as this resolution is similar to running someone through a joint cal. | Brakes susceptible to contamination
1874 | Right foot was not properly completing step | closed | kmettler | Foot Sensor | fixed | 2014-09-03 18:03 | 2014-09-02 | 52739 | 140 | Distributor - Italy | Matteo Carelli | m.carelli@emac.it | 39 347 5434417 | Italy |  | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Unassigned |  | Design | EMAC have reported an issue with the right leg on device #140. Customer: Villa Beretta (Dr Molteni) Unit 02146001 #140 Attached 2 videos which unfortunately I can‚Äôt open. Adrian is currently travelling and I will liaise with the customer in the morning in order to gather more information. From Chad 9/2: I was able to open the video and to me it looks like the right foot is not properly completing the step. Please walk the customer through a foot cal as that should help with the issue. | Verified that foot sensors were functional during SWO-1179. | The foot sensor design requires calibration of the sensing pads as the foot sensors deteriorates.
1873 | Foot Sensor- separated | closed | kmettler | Foot Sole | fixed | 2014-09-03 17:54 | 2014-09-01 | 140589 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Determined Unknown | Dear friends, Last week we had another problem with EKSO soul. Next one has break. Please see attached photos. We are wondering is it normal that it breaks so often? At this point we have exchanged souls in both our EKSO due to this same problem. Can you kindly sen us new spare one? I am afraid that this kind of problem will occur more in future and we must be ready to fix it in no time. From Reuben 9/1: Hi Adrian We are working on some improvements on the foot sensors that will make them more durable. We will be using a different adhesive that is more ductile. Also we are pushing the vendor of the red sole to deliver them without tiny bubbles. They have not been able to deliver us perfect parts yet. | Replaced foot sensors superceded by work completed via SWO-1408. | The damaged foot sensors were not properly returned to HQ and because of such are now unaccounted for and have not been found.
1870 | Shoulder strap buckle cover stitching pulled away | closed | kmettler | Soft Goods - Shoulder Straps | fixed | 2014-09-03 17:32 | 2014-08-28 | 64645 | 116 | Good Shepherd - IP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Wear | Gs #2 Sn : 021402005 Steps 64645 Back pack straps black fabric cover over strap and sling clips is impaired on both sides. Left side has tear, a right side has stitching starting to pull away from outside seam near where the clip is to release the backpack strap. | Back straps replaced in field via SWO-819 & SWO-1061 CO-11180 updates the B101073/4 shoulder straps to a new design. | Found that stitching through arm sling buckle was not shown through k-text layer. Pictures from diagnosis are attached. No process was defined at time of production of this unit.
1867 | Foot sensor peeling back on feet | closed | kmettler | Foot Sole | fixed | 2014-08-27 17:39 | 2014-08-27 | 10335 | 149 | SDSU | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | Customer SDSU Unit: 149 Steps: 10335 PT sent through these pics of the bottom of their foot sensors. Response: Are we able to send them adhesive so this can be glued back down? Using on carpet can cause this issue. | Adhesive was sent to the customer. | Glue used to attach Vibram inserts can peel with wear over time. Brittle adhesives were previously used which made this failure more likely to occur.
1865 | Plastic housing around power button is elevated | closed | ckealoha | Torso | fixed | 2014-08-27 17:27 | 2014-08-27 | 67707 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | This is device #3 at GS. See attached photo with black plastic housing around power button is elevating. I just noticed this now. Unsure if was like this in AM session. I can wait push it in place and it lifts back up. This is the device that hit the ground when the chair tipped a couple months ago and this same piece broke, and someone came out here to fix. Can I use this device? I need an immediate answer as I have another patient waiting. | Removed old shroud and replaced with new shroud. Work was done per SWO-818. | The method of installation was deemed as insufficient which caused a redesign of the installation method. This change has been implemented by CO-11171.
1864 | Hip width very difficult to adust- both sides | closed | ckealoha | Adjustments - Hip | fixed | 2014-08-27 17:18 | 2014-08-27 | 67707 | 132 | Good Shepherd 3 | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Materials | Good Shep #3, SN attached in picture. Issue #2: Team reports they have already reported that the hip width is getting right to narrow. Have been trialing the tapping on the hip width bolt; Troy said that no longer makes a difference. Both hips intermittently problematic; today, it appears narrowing of right side is the tight. Attached video shows two reps of narrowing/widening, with first rep no issue, but second rep is tight in narrowing. I narrate what I ask them to do and am seeing. | During initial inspection the right hip appeared to have restricted motion. The hip width lock key was replaced and solved the issue. The was no issue with the left side of the device. See photos and videos for further clarification at M:\Manufacturing\Pictures\Good Shepherd\Hip Lock Issues. This work was done per SWO-818. | Hip width lock key missing R.032 fillet around full edge. Known issue. Reference NCR-10995
1862 | Torso Test Required- Error #133 & #121 LH Code 45056 (Motor Short) | closed | rrea | Electrical Interconnect | fixed | 2014-08-27 17:11 | 2014-08-26 | 38617 | 113 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service | Minor |  | Design | The Ekso asked for a torso test, in the morning when they started and also gave the error 133, I went through the system test for torso and legs, this passed both but the error 133 was still there. They have attached a video of the hips and it sounds like this is a mechanical issue. S/No. 021402002 From Katie 8/26: 133 is a motor drive fault. I don‚Äôt have data, but if we can get the error info from the tech screen, we can look into it more‚Ä¶but if it‚Äôs not clearing, then it likely needs serviced. | replace shorted FET, ran through formal opqual process with SWO-00812 | https://eksobionics.centraldesktop.com/p/aQAAAAACHeWW CAPA being created. Not a safety risk. Will cause motor to cog and throw and error requiring a torso test and service.
1861 | Feet clipping during session | closed | jwitherspoon | Lower Leg | fixed | 2014-08-27 16:57 | 2014-08-25 | 350218 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner |  |  | Customer Use | Please see below issue from Mark Pollock. Unit #39 Step Count: 350,218 Response needed ASAP so we can get back to customer with follow up. Mark sent through 2 video‚Äôs and they are both attached to this email. Issue description: I've attached a video which shows my right leg swinging in and clipping my left leg from time to time. Our settings aren't unusual we don't think but since it has come back from the cable being fixed we've had 2 walks with this happening. For today's session Simon widened the stance to take that out of the equation but it was still happening. We wonder is it linked to the problem that Simon flagged previously (it has been happening since May/June). That is, we walk for 3 x 15 min blocks. Then sit down and change the batteries. When we stand up again after walking it is like the device is calibrating or something. Almost every time that we walk, change the batteries and then walk again, one of the legs takes mini-steps or I have to go way out to one side to get it to step. It is like I have to limp until we get to the end of the hall then it gets slightly better. And, then after another length of the hall it is almost back to normal. There does not seem to be any pattern it could be the left one session and the right the next. However, today it was really bad on with the right taking mini-steps until it sorted itself out. Note the video that is attached is just showing the clipping problem that is happening all throughout the sessions but I wonder is it related? It feels like the sensors or software aren't working like they used to. | The right lower shank was re-glued into the correct configuration to remove the toe in issue. The left lower shank was removed and replaced with a shank which was tested for proper alignment | During inspection while the unit was in house on SWO-0003 it was found that the right lower shank's glue bond was broken causing it to toe inwards. The number markings on the left lower shank were inspected and found to be out of alignment relative to the lower pin hole which is used for alignment.
1860 | Software Error- 1241-RH | closed | ebabaran | Electrical Interconnect | fixed | 2014-08-27 16:50 | 2014-08-21 | 3213 | 150 | VA Richmond | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | I had a software error come up with error code, instruction to call Ekso. Error code 1241-RH at 3:54pm EST after the initial stand of the patient's session. I had the team go through the appropriate steps to call support hotline and there was no answer on two attempts. I verified the phone number the PT called and it was accurate. I followed the steps to reboot and then contact Ekso command came up. I reboot a second time and it rebooted appropriately, set appropriate settings and attempted the Second stand. We were able to resume the session as scheduled. | SWO-814 replaced the signal chain for the hip cable to reduce strain on cable. | Previous design allowed for pinching of the hip signal cable that would result in these errors
1859 | Motor Failures when unlocked hip abd | closed | kstrausser | Customer Use | wontfix | 2014-08-25 16:35 | 2014-08-20 | 147523 | 117 | Regi Base - Czech Republic | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | I noticed some motor failures on unit 117, so we asked the clinic for more details. Below is their description followed by the relevant boots from the error log. The 12288 error reported by all 4 Elmos in both these error logs is the motor undervoltage error ---from the clinic--- No fault code presented It happened 4 times in two sessions with 1 hemiplegic patient (already ambulation) once we chose the unlock hip ABD and min. stand support. It was like a disruptive movement with sudden excessive bilateral knee flexion and after 3 second with extension and continue with ambulation. During the incident we had to hold the patient firmly to prevent the fall. Once we had to bring the chair and discontinue the session, next time it resumed the standing position automatically and we could continue. after the accident we lock again the hip ABD and reset the stand support to med. From this time no history of issue but we¬¥ve never continued walking with unlock hip ABD. --/from the clinic-- --Boot log 1 -- 17545, 2787147, 1156, 1, 457, 801, Power off message 17544, 2787121, 1015, 1, 68, 387, I2C comm failed on Battery 17543, 2153403, 165, 2, 88.48132, 0, Stand Failed - Move feet back and retry 17542, 2054111, 90, 1, 1632, 0, Control output exec time too long 17541, 2050241, 121, 3, 12288, LK, Elmo motor failure bit set 17540, 2050241, 121, 3, 12288, RK, Elmo motor failure bit set 17539, 2050239, 121, 3, 12288, LH, Elmo motor failure bit set 17538, 2050239, 121, 3, 12288, RH, Elmo motor failure bit set 17537, 2050189, 133, 3, 0, 0, Motor drive fault 17536, 2050189, 1058, 1, 58, LH, CAN PDO received w/ motors off error 17535, 2050189, 1058, 1, 58, RH, CAN PDO received w/ motors off error 17534, 2050189, 1058, 1, 58, LK, CAN PDO received w/ motors off error 17533, 2050189, 1058, 1, 58, RK, CAN PDO received w/ motors off error 17532, 287023, 164, 2, 0, 0, Crutch Buttons disabled 17531, 16683, 1102, 1, 1, 1, Bluetooth connect changed status 17530, 1201, 12, 1, 1773398550, 74, Battery R info message 17529, 1161, 11, 1, 2476403222, 73, Battery L info message 17528, 21, 1066, 1, 29885422, LK, Joint board info 17527, 21, 1066, 1, 3212270, RK, Joint board info 17526, 13, 1066, 1, 29885422, LH, Joint board info 17525, 13, 1066, 1, 29885422, RH, Joint board info 17524, 1, 1159, 1, 2014/08/08, 12:45:28, Boot time --/boot log 1-- --boot log 2-- 18579, 0, 1150, 1, 489, 2898, Boot message 18564, 2748081, 1156, 1, 488, 1179, Power off message 18563, 1809141, 121, 3, 12288, RH, Elmo motor failure bit set 18562, 1809141, 121, 3, 12288, LH, Elmo motor failure bit set 18561, 1809043, 121, 3, 12288, LK, Elmo motor failure bit set 18560, 1809043, 1058, 1, 58, RH, CAN PDO received w/ motors off error 18559, 1809043, 121, 3, 12288, RK, Elmo motor failure bit set 18558, 1809043, 1058, 1, 58, LK, CAN PDO received w/ motors off error 18557, 1809043, 1058, 1, 58, RK, CAN PDO received w/ motors off error 18556, 1809043, 1058, 1, 58, LH, CAN PDO received w/ motors off error 18555, 611033, 164, 2, 0, 0, Crutch Buttons disabled 18554, 72893, 1102, 1, 1, 1, Bluetooth connect changed status 18553, 1201, 12, 1, 2476403222, 70, Battery R info message 18552, 1161, 11, 1, 1773398550, 71, Battery L info message 18551, 21, 1066, 1, 31917038, LK, Joint board info 18550, 21, 1066, 1, 5243886, RK, Joint board info 18549, 13, 1066, 1, 31917038, LH, Joint board info 18548, 13, 1066, 1, 31917038, RH, Joint board info 18547, 1, 1159, 1, 2014/08/20, 06:05:15, Boot time --/boot log 2-- | Based on response above from Katie unit operated as designed and customer use was cause for issue seen. | no recurrence after clinical training.
1853 | Side Closure- Rivet broke | closed | jwitherspoon | Soft Goods | fixed | 2014-08-20 17:56 | 2014-08-20 | 38617 | 113 | Spandau - Germany | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Solution Not Implemented | Unit #113 popped rivet. While doing training with their second device, the head PT notified me today that ages ago they popped a rivet on their older unit. They have been working with it regardless as it is the left bottom one and hasn't been limiting their practice. This is one of the units that had the rivets fitted with the wrong tool. Please send them a new side piece with instructions on how to fit. | Replaced rivets with stronger rivets per SWO-00809 | Strength of rivet was insufficient for the shear stresses encountered.
1851 | Paint chip on Ekso handle | closed | ckealoha | Torso | wontfix | 2014-08-20 17:43 | 2014-08-18 | 3213 | 150 | VA Richmond | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 3 - Next Service | Minor |  | Design | Upon opening the crate, the safety handle had a paint chip in the bottom right hand corner. Not sure the unit # 150 Step count 3213 Here's a photo | Touch-up done during each service as part of routine inspection. Inspection always done for replacement items. | Known issue. Due to the painting it can chip. This is a low priority, low risk problem. The defect will be corrected in the next version of the device.
1850 | Walker- Difficult to lock in position | closed | kmettler | Cane / Walker / Crutch | fixed | 2014-08-20 17:36 | 2014-08-15 | 166262 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | Customer- Technomex Issue-broken walker, difficulty with locking in place. Just realised that it was stuck in my outbox from last Friday. | Sent new walker via SWO-806. | Walker has difficulty locking into place.
1849 | Residue on batteries when unpacking | closed | kmettler | Shipping / Handling / Packaging | fixed | 2014-08-20 17:23 | 2014-08-15 | 3701 | 149 | SDSU | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 1 - Phone Calls - Single |  |  | Materials | From Melissa 8/15: This is a late report to notify something found on delivery at San Diego State University, for their Ekso GT that was delivered/trained this week. -no response needed -Found while unpacking the new Ekso; steps 3701 -when unpacking the batteries and chargers from the case, I discovered a slightly sticky residue on a couple of the batteries, a charger, and inside the foam casing. I haven‚Äôt seen this before. It was similar to the sticky adhesive that attaches a mailing label to a magazine- same color but thinner substance. It didn‚Äôt rub off easily, and smeared a bit before coming clean. Eventually wore off with repeated attempts to wipe away. Sorry for the late notice. The attached pics show a few locations. The first pic with the gray foam shows one spot where the substance was found inside the foam lining. (look in middle of pic). Please let me know if you have any questions. Thanks, Melissa | Notify supplier and add to quality check. No longer used.NO issue with new packaging. | Excess adhesive on foam in shipping case
1848 | Batteries draining unevenly/ Mismatched error as result | closed | rrea | Battery | fixed | 2014-08-20 17:00 | 2014-08-14 | 166262 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | From Beata 8/14: Customer- Technomex Batteries S/Ns: old 01D09619160000E2 new: 01ECD8EA160000C0 Issue: when two batteries are used in Ekso one battery drains quicker than the other and Ekso turns off. It happened few times and now PT works with one set only to avoid unsafe situation. Just to remind you we replaced one battery recently. From Edmundo 8/14: In talking to Keith and Mark F. sounds like the culprit with this battery is a weak cell which we cannot fixed and has been resolved in the past by taking the battery out of rotation. Please send a replacement battery to the customer. In addition we should look at the datasheet provided by Rose Electronics for this particular battery. Thanks. | replacement battery sent with SWO-00806 | Additional analysis to be done with Ticket #2384 then updated here.
1847 | Step Abort Error- #168 | closed | kmettler | Customer Use | fixed | 2014-08-20 16:52 | 2014-08-14 | 166262 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | From Beata 8/14: Customer- Technomex Unit S/N- 021404004 Step count- 166 262 Issue- Ekso throws error 168 when taking first right step if the patients is are not far forward. It happened with more than 2 patients. PT and myself feel that we need to really go forward to initiate the step (tried to video the position for you to see). PT said that this is happening a lot since Adrian fixed Ekso. We were walking in Prostep Adaptive. Let me know if you need any more information. I am with Technomex today so if you need any more info let me know. From Darrell 8/14: 168 error, it is a ‚ÄúStep Abort‚Äù error. Generally this happens when the foot is not coming off the ground early enough due to tone/spasticity or a patient desire to keep the knee straight. A few ideas to help: 1. ensure a lateral weight shift is completed to unload the foot before initiating the step 2. Increase the step height which will ask for more hip/knee flexion and may improve the first step (and others too). There may be sensor related causes too, I‚Äôll let engineering handle those possibilities. | Clinical explained how to properly take steps with this patient. | Additional training needed.
1836 | Unusual noise coming from the hip area - current dropped on one hip without warning | closed | kmettler | Encoder / Potentiometer | fixed | 2014-08-13 17:00 | 2014-08-10 | 350226 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Process Not Followed | 8/10/14 Customer: Mark Pollock Unit # 39 Steps: 330,117 Response needed: ASAP- Mark is getting a loaner unit to support him while his device is shipped to Adrian‚Äôs workshop for repair. Summary: Mark reported a noise coming from his Ekso. He was able to take a video of the noise and it‚Äôs attached to this email. Below are comments from both Mark and Simon regarding the issue they are currently experiencing. Simon‚Äôs comments (Mark‚Äôs spotter): There's not too much to add to Mark's comments on the left hip. I'd agree that it felt like a stiffness developing in the same joint that we had a problem with the last time Greg was over. I've attached the videos showing the squeaking and the clicking. We had been planning to get in touch on another issue which I'll mention in case they are related: At the moment we change batteries after 45 min and then continue on. Often after the battery change, Mark seemed to have to lean pretty heavily to the left side to get the lateral beep. This would last for about a minute and then it would start to get back to normal. We assumed that it was some kind of calibration glitch. It also happened today at the very beginning of the session. Mark‚Äôs comments: Simon will follow up with 2 videos to show the problem that emerged today. But I wanted to flag this problem ASAP because I am trying to hit a series of targets which require me to walk 7 days a week at the moment. The left hip is squeaking and it feels like it is tight. By the end of the session today I was finding that I was walking poorly - almost like I was limping. Device came back to HQ for eval and found current dropped on L hip. | Original Resolution: replaced hip cable. Parts were sent out on SWO-801. But replacement ended up comingfro EU stock of parts. New Resolution: After the above root cause analysis was completed a new encoder was fitted to the shaft according to the current manufacturing process. | Original Diagnosis:shield wire trapped between hip housing New Diagnosis: During inspection of the device it was found that the left encoder set screw was not properly engaged on to the flat of the motor shaft. The hub of the encoder was held in place by a small shard of metal which was pressed against the hub's inner diameter and the motor shaft. Please see attached photos for further clarification. The insufficient connect between the encoder hub and motor shaft is believed to allow for the encoder hub to slip and eventually cause commutation errors which over time which caused the current to unexpectedly drop. The hub slipping was however not great enough to provide an error.
1835 | #121 & #133 LH Error while walking motor or cable shorted | closed | kmettler | Electrical Interconnect | fixed | 2014-08-12 01:24 | 2014-08-07 | 314000 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 9 - Loaner |  |  | Design | From Josh 8/7 Issue: While walking a patient a #121 LH error was thrown and the patients left leg was suspended in the air due to the knee‚Äôs locking. Troy, who was spotting the patient is unsure of the exact sequence of events after this point but he claims the Ekso was completely frozen. He‚Äôs reporting the knee‚Äôs and hips locked up, he called for assistance and they put the patient on a tether. During this process he pressed the yellow button and this freed the hips and kept the knee‚Äôs locked allowing them to lower the patient down to a chair. Once patient was safely in the chair he noticed error #133 was reported sometime during this process, he‚Äôs unsure when this occurred. Both Patient and Spotter are fine. Followed up with customer after patient was out of device and when they rebooted the error immediately follows. It‚Äôs also reported when the Ekso is off the left hip feels very rigid and doesn‚Äôt feel normal. I need input as quickly as possible for this situation as they are currently down and we need to determine an action plan to get them back up. From Adam 8/7: The left hip Elmo is reporting that the left hip motor is shorted. This also explains why it feels ‚Äúrigid‚Äù when powered off. The joint should still be moveable, but you will feel a (possibly coggy) resistance to motion. Either a motor issue, or the wiring has melted and shorted. From Josh 8/7: We will be shipping them a loaner device and it will be arriving on Monday AM. We should have their device back @ HQ on 8/12-8/13. | replaced transistor via SWO-507 | Phase C transistor on shorty board was failed short. Transistor replaced and tested. Ready to use in 108 again.
1827 | Unit 39 leg calibration needed | closed | adamzoss | Encoder / Potentiometer | fixed | 2014-08-01 15:41 | 2014-07-29 | 315000 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Unassigned |  | Process Not Followed | At 315,000 steps during a sit to stand operation unit requested a joint cal to be done. Customer got out of unit and was walked through a joint cal by FSE. This is unit 39 SN 011209002 and no error was thrown when this occurred. | Issue fixed via email with technician Adrian Poultney. Adrian led the users through a joint calibration. | Joint needed re calibration.
1823 | Ekso not reading battery | closed | lmcneil | Battery | fixed | 2014-07-28 21:16 | 2014-07-28 | 114107 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | One of the batteries is not recognized by the system. It does not charge, regardless of the charger. Any hints what we can check on site ? The battery and the battery socket does not have any visible marks that could suggest some breaks.... | Sent customer new battery via SWO-00469 | One of the batteries was not recognized by the system. It did not hold a charge, regardless of the charger.
1822 | Battery not charging fully | closed | rrea | Battery | fixed | 2014-07-25 19:05 | 2014-07-24 | 188500 | 19 | RIM | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost | Minor |  | Design | We are having a little trouble with 1 set of our batteries. It is only charging up to 87% at best. Is there anything we can do? | Replacement of batteries when cells have become to far out of balance. Trac Ticket #1823 and SWO-00560 replaces battery. | Cell balancing is inefficient. Additional analysis to be done with Ticket #2384 then updated here.
1815 | Knee stuck in fixed position | closed | ebabaran | Brake / Release / 9V Battery | fixed | 2014-07-23 17:30 | 2014-07-19 | 134471 | 63 | Gary Linfoot | Gary Linfoot | glinfoot68@gmail.com | (931) 801-0205 | United States | no injury | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Major |  | Design | I got a call from Gary Linfoot over the weekend while he was finishing up a demo event with his Ekso. When he was going to put his Ekso in the crate he noticed his left leg was stuck in position at the knee. I asked him if this only occurred when the device was off and he stated they turned the Ekso on and freed the legs and the left leg was still stuck in the same position. Luckily it was bent in a way he was still able to pack in the crate and he is currently traveling back from Denver to his home today. If we need him to run any tests he will be able to do so this evening or tomorrow morning. If we feel his unit needs to come back to HQ then I can begin arranging a loaner for him. | Unit being fixed in house per SWO-513. Root cause was found to be known issue which we are fixing with new design of the bearing clamp. | Poor attachment of bearing clamp to housing.
1805 | Toe guard fell off- Error #6 while setting up Ekso | closed | ebabaran | Foot Assy | fixed | 2014-07-16 17:49 | 2014-07-15 | 3790 | 142 | Burke | Helma Zanders | wzanders@Burke.org | 914-548-2857 | United States |  | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Design | 1. The PT wasn‚Äôt sure when or how it came off, it was noticed in the hallway. He will be sending me a picture of the toe guard and I will forward when received. 2. While trying to stand a patient they received error #6 over heated battery. They followed the manual and let the Ekso cool for 10 minutes and it seems to be working appropriately at the moment. They don‚Äôt have a signed data agreement so we won‚Äôt be able to pull their data, is there anything further we should ask customer to determine why they got this error? | Toe guard re-glued when unit serviced in field per SWO-782 | Adhesive used with toe guard can become over-stressed and fail.
1796 | Battery Charge Issue | closed | ckealoha | Battery | fixed | 2014-07-11 16:59 | 2014-07-10 | 12332 | 112 | Heliomare - Netherlands | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | We now use EKSO about 3 hours a week with different clients. There are some questions though: Problems: ‚Ä¢ One of the battery pack just doesn‚Äôt seem to charge a 100%, although we have put them on the charger for days. The charger light is turned on green, than we put it in EKSO and on the display we read battery level 23%. The other pack doesn‚Äôt have that kind of problem. It there anything we can do about this? | Battery has been inspected and accepted back into manufacturing as a refurbished battery. Replacement battery shipped with SWO-00781 | The at the time of production MPI did not explicitly direct to use all batteries to be tested during production. Revision E of the MPI corrects this oversight.
1785 | Unit 25 leg calibration needed | closed | ckealoha | Encoder / Potentiometer | fixed | 2014-07-08 15:45 | 2014-07-08 |  | 25 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Unassigned |  | Undetermined | Unit 25 SN 011307001 populated an error which required a joint recal. | Issue fixed via email with technician Adrian Poultney. Adrian led the users through a joint calibration. | Joint out of calibration
1773 | Hip width slipping during patient session | closed | rsandler | Adjustments - Hip | fixed | 2014-07-01 21:03 | 2014-06-25 | 109500 | 47 | Scripps | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | So we have been using the EKSO regularly with the same gentleman who we noted the overheating motors with and have altered the parameters per Darryl suggestions. We have had improved success with the battery life however its down to about 27%-32% after 50 minutes and just over 2000 steps. Is this typical? Also in the last two sessions with this user on Wednesdays the Left hip width loosens up. It started at 2 this past Wednesday and ended up at 4 even with our standard tightening procedures. And the weirdest thing was that the leg hair on this users outer Left thigh became singed. It curled up and smelled like burnt hair but the pants and skin were intact. Right near the Left mid to upper thigh. These last two events are new since we received the repaired EKSO. Please let me know how to proceed we don‚Äôt have a tether and this user walks all over the hospital. | Josh contacted the customer and walked them through an ideal thigh brace and strap setup to prevent abrasion. | The customer believed that there was a burning smell, however, burning did not happen. The root cause it that there needed to be ideal thigh brace and strap setup to prevent abrasion.
1756 | Step Aborts | closed | ebabaran | Foot Sensor | fixed | 2014-06-23 21:37 | 2014-06-20 | 239101 | 108 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Design | I left you a voice msg but i wanted to let you know that our ekso unit # 1 is throwing several step aborts. This happened to 2 patients on the R leg only. It happened on Friday from 3:00-4:00pm and now it is happening today from 1:00-2:00pm. The patients both had adequate wt shift and does not have any tone issues. I don't think its a safety issue, but need your thoughts on the cause. | It was determined that new foot sensors were need to fix this unit and they were replace din the field via SWO-751. | Sensors may drift over time due to deterioration
1747 | Foot sensor torn at Technomex | closed | kmettler | Foot Sole | fixed | 2014-06-17 22:10 | 2014-06-13 | 51562 | 129 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | This came in from Polish Distributor Technomex. We have already mailed out 2 sets of foot sensors and Adrian will be on site Monday to replace. But this is now the 2nd Ekso that has ripped sensors from this Dist. | Installed new foot sensors per SWO-00661. | Old sensor material has low tear strength and is susceptible to damage due to wear and high shear loading during turns.
1741 | Tear in foot sensor | closed | ebabaran | Foot Sole | fixed | 2014-06-11 17:37 | 2014-06-11 | 3005 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | Technomex customer reported a tear in their foot sensor and having issue's with irregular walking pattern's. Adrian will be going out 6/16 to work on both their devices, new foot sensors are being sent today 6/11 to Distributor in Poland. | New foot sensors sent out per SWO-00727. CAPA #1259 will address foot sensor robustness. | Action - Continual use (walking + turning) on different surfaces Condition - Brittle adhesive causes cracks and tears which can propagate under stress Condition - Capsole material not robust enough
1728 | Vibram coming off the bottom of foot plate | closed | kmettler | Foot Sole | fixed | 2014-06-03 19:03 | 2014-06-02 | 166081 | 25 | Miami Project | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service |  |  | Design | So things have gone from bad to worse. In an attempt to provide a provisory remedy for the loose sole strap (it was tripping up the subjects trying to step) we may have affected the sensors on the sole. We tried to glue the strap on and clamped it for about an hour while the glue was drying to hold it in place. When we just tried to walk another subject the stepping was off by a lot. It seemed to only trigger the left step when the subject was leaning a bit to the inside and much further forward than is usually necessary - we tried to compensate by adjusting the settings but it did not work. Unfortunately this will not allow us to continue walking subjects. Please advise what best to do the fix this asap. | Foot sensor replaced via SWO-723 | Known wear issues with current foot sensor design.
1721 | Step length issue with Ekso- Foot Sensor issue | closed | kmettler | Foot Sensor | fixed | 2014-05-28 17:49 | 2014-05-23 | 3005 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Design | From Barry 5/23: I received the following video from Chris at Technomex. The limited information I have is that they have been practicing with a Physio in the device. The spotter and person in the device notice the issue in the video (see link). I am unsure of what the SN is or what settings were loaded. I have forwarded it to Adrian to follow up. They report no errors and no step aborts. I initially thought it may be a weight shifting problem but looking at the video I think there may be something else. Adrian will follow up with the customer and report back to the team. Take a note at every second step with his left leg he makes with an EKSO. Related to Trac #1741 | New foot sensors sent out per SWO-00727. CAPA #1259 will address foot sensor robustness. | Tear in foot sensor from Trac #1741
1719 | Noise and movement coming from left hip joint | closed | jwitherspoon | Tools | fixed | 2014-05-28 17:43 | 2014-05-27 | 270470 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Clear/Complete | We walked today and the left hip joint was loose. We assumed it was the same problem as before with the hex bolts not tightening but on closer inspection it is the hip. I have attached the video - what do you think? We don't have the right tools to tighten it and don't want to without direction. | SWO-00815 was generated and new tools were sent to Mark. | Tools were out of calibration and customer wasn't able to fully tighten hip width.
1707 | Crutch interface doesn't work at Technomex | closed | kmettler | UI - Crutch Interface | fixed | 2014-05-21 17:16 | 2014-05-21 | 3005 | 125 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Email from Krzysztof: Hi Barry I have a small problem with the crutch interface. Yesterday all went fine but today it does not work... it does npt light up nor it does not make any sound when unplugged from the backpack unit... any hints how can we check that? | SWO- 698 was generated and new crutch interface was sent to the customer. | strain relief not properly attached causing failure
1695 | Broken walker latch @ customer site. | closed | kmettler | Cane / Walker / Crutch | fixed | 2014-05-14 17:33 | 2014-05-12 | 19736 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 6 - Mailed Part - High Cost |  |  | Materials | The other day when I was getting our walker set up, I must have been feeling really strong, because one of the sides sprung open all the way and a piece of the plastic by the locking mechanism broke off. We can still nudge the locking mechanisms back into place with a little extra effort to use it, but not sure how long that will hold. Can we get a new one sent out to us? Anything else you need from me to get it on its way? | New walker sent via SWO-0707 | One of the sides of the walker sprung open and a piece of the plastic by the locking mechanism broke off.
1691 | Bad Aibling DDL clevis marked | closed | ebabaran | Don-Doff Link | fixed | 2014-05-13 20:45 | 2014-05-13 |  | 48 | Bad Aibling - Germany | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 0 - Not Required | Minor |  | Wear | In addition, today noticed that the left Wegklappmechanismus , who recently only wg . the material breach has been replaced, is clearly difficult to move than the right . Here are the paint can still be seen two small strokes and keys ( see photo, at the transition from the block to the two joint parts . They resemble the ones we had at the material fracture . Here I can not judge whether they are normal and for example, come from painting or whether they indicate a material deformation . had the joint not so stiff it would not be noticed . | Affected DDL rebuilt per SWO-528 | Today noticed that the left Wegklappmechanismus , who recently only wg . the material breach has been replaced, is clearly difficult to move than the right . Here are the paint can still be seen two small strokes and keys ( see photo, at the transition from the block to the two joint parts . They resemble the ones we had at the material fracture . Here I can not judge whether they are normal and for example, come from painting or whether they indicate a material deformation . had the joint not so stiff it would not be noticed .
1666 | Left thigh adjustment is harder than right | closed | kmettler | Adjustments - Leg | fixed | 2014-04-21 17:41 | 2014-04-18 | 11914 | 110 | UHS | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 7 - Clinical Visit | Unassigned |  | Design | At UHS and the PT commented that the left thigh adjustment is harder to do than the right. It is manageable and easier after I wiped the tube with a wipe. Rotating in or oscillating it made it easier to adjust as well. It is still more difficult than the right. | Reamed out bore on thigh adjustment structure per CO-11217 on SWO-509. | Dimension on thig adjustment structure to small.
1642 | CI routing Velcro not installed properly | closed | kmettler | UI - Crutch Interface | fixed | 2014-04-09 16:15 | 2014-03-27 | 3104 | 119 | Madonna | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 5 - Mailed Part - Med Cost |  |  | Process Not Followed | From Mike Glover 3/27: I'm confused on the new Velcro pieces for the crutch interface. we have some new pieces that I'm not familiar with and I think the rectangular piece was not threaded through the actual cord. I can't seem to thread it. From Kelly 3/27: The extra rectangular piece was supposed to be on the cable. We are going to have to send a replacement since there is no way to put that on in the field and get that one back. | SWO 617 was generated and new CI was sent to customer. | The CI did not include the extra rectangular piece on the cable.
1637 | tibial pad fabric folded and caused redness and marks on shin | closed | kmettler | Soft Goods | fixed | 2014-04-08 22:27 | 2014-03-28 | 6044 | 115 | NRH - Med Star | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 2 - Phone Calls - Multiple | Unassigned |  | Customer Use | Hi Team, Yesterday, we had patient have complaints of pain on her shin bone and we had to sit twice to readjust. Pressure/pain source identified as the seam of the tibial pad on the inside border of the pad. I noticed the seams on this tibial pad appear thicker and more prominent/ more in line with the curve of the pad inside the tibial shell, leading to being able to feel this on the leg (if the patient has sensation- she did, and was able to notify us of her pain). After walking, patient had visible indent over her bone/tissue from the seam. The resultant pressure indent remained on her skin for just under 2 hours last night. Attached are various angles of the pad in the suit, as well as pressure point site on patient. | Customer determined that it was the pant leg fold instead of the pad. Internally looked at different design options for a later design revision. | Determined was customer pant leg
1636 | Left leg of Ekso making unusual sound | closed | kmettler | Customer Use | wontfix | 2014-04-08 22:21 | 2014-03-28 | 60000 | 30 | Merryweather - So. Africa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 3 - Next Service | Minor |  | Customer Use | Just wanted to let you know ‚Äì the left leg of the suit is making a funny sound when walking (a kind of clicking sound), and also doesn‚Äôt have a smooth gait pattern anymore on that left side ‚Äì we walked Justine yesterday who is walking very well now and on the crutches, and she picked it up straight away ‚Äì as well as after some time walking, the left foot also kept catching the floor. I adjusted the step height to see if it would make a difference but it didn‚Äôt. We haven‚Äôt walked anyone new except the people you have already met Barry ‚Äì and Brandon, the 18 year old who was at the Shepard Centre when you were here ‚Äì I mention this just to illustrate the people who have been walking in the suit ‚Äì no-one new or oversized! If you could let us know soonest what action plan we could take to get this sorted as we were just about to open it up to the public. We could also send a video and sound of the left leg if that would assist you. | Will be inspected via 100-point whenever unit comes in for service. 100 point has walk evaluation to check for any issues including sounds. | Customer use issue, could not reproduce
1628 | Foot binding pad- Incorrectly assembled | closed | jwitherspoon | Soft Goods | fixed | 2014-03-26 18:38 | 2014-03-26 | 3049 | 118 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Followed | Can you please send a new foot binding pad to Technomex today? After borrowing a customers Ekso we could start the training today and we then realised that the velcro was missing on one of them, this is the soft goods kit to the Ekso that was sent to Technomex (Elso serial 021402007). Please see attached pic (part 102630a). | SWO-628 was generated for this part and was hand carried to customer location by FSE. | Process not followed when assembling foot binding.
1624 | Ekso threw error #25 RK | closed | pmcstravick | Electrical Interconnect | fixed | 2014-03-26 17:43 | 2014-03-25 | 3049 | 118 | Technomex - Poland | [5 empty] | Same as contact listed | ekso 1.2 | FALSE | 8 - Field Service |  |  | Materials | We need an urgent response as this will potentially cost us a day in training if not worse if not field fixable. Day 1, walk 1 (therapist). Dennis is leading the training and can provide more details as required. Essentially, from what I understand, the Ekso threw the error in the attached picture on standing. No steps were taken. I ran them through a joint test and it failed (see attached picture). They hit all the hard stops as required. Please advise asap as we are with the distributor on day 1 and this does not look good at all Disassembly of device showed that an inductor on the right thigh board was cracked on the corner. Add picture Nell found one in stock was cracked and Aaron found one in a kit that was broken. Edmundo and I found that there was .03 in of clearance between the shield cover and inductor. Wires can be trapped between the shield and inductor and break. Adding a cable tie can ensure the wires can't be pinched and cause the crack. | Unit fixed in field by replacing knee joint board per SWO-619. System issue still needs to be investigated further. | Disassembly of device showed that an inductor on the right thigh board was cracked on the corner. RSandler: 2/19/2015 This is not likely the cause as Kieth later demonstrated in house that breaking the inductor didn't cause the same failure. A more likely cause is the thigh cable clamping or wire pinching issues described in 2087-CAPA. Materials: Eccentric cable from Minnesota Wire Design: Tight clamping
1617 | Errors thrown after batteries were changed-#1160 | closed | kstrausser | Software | fixed | 2014-03-26 16:18 | 2014-02-27 | 204426 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | 26th February: Error 1160 (we aren't 100% sure that we've taken it down correctly) - We had changed batteries and it happened during the stand sequence. 27 Feb: Error 1110 - happened on reboot. Today we had the following: 1160. The device seemed to pause close to the top of the stand. From Katie 3/20: There is nothing wrong w/ the machine (just our code), so they should carry on. ÔÅä (the code protects itself‚Ä¶thus the error‚Ä¶but we do need to fix it‚Ä¶) Trac #1876 is the Ekso 1.2 version of this ticket. | Per Katie's analysis the unit is functioning as intended. | Error in calculation of the 'c' variable in cases where there is no steady state velocity. checkin for Puma has been done, still need to update 1.1.
1607 | Customer broke Crutch Interface | closed | rsandler | UI - Crutch Interface | fixed | 2014-03-19 16:24 | 2014-03-17 | 215310 | 36 | Gaylord Hospital | Erika Ozdemirer | eozdemirer@gaylord.org | (203) 284-2800 | United States | no injury | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Unassigned |  | Design | PT reported they had a broken Crutch Interface. He wasn't sure how this occurred but he suspects it happened during a stand to sit. | SWO-599 was generated and new CI was sent to customer. | Insufficient strain relief for CI cable.
1588 | Crutch interface broken- cord was pulled during stand to sit | closed | jwitherspoon | UI - Crutch Interface | fixed | 2014-02-26 19:30 | 2014-02-25 | 56380 | 14 | Spaulding | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Customer Use | I just wanted to let you know that we are having an issue with our advanced features. The hand unit that we need to plug in is not being recognized by the computer (Error # 164 ‚Äì saying it is unplugged) when it is in fact plugged in. I think we may have injured the cord - during a sit->stand the cord got caught and pulled a little. Not sure what to do about it.... I apologize. | New CI was sent to customer. SWO-00585 was generated for new part. | If customer not careful, HMI cord can catch and pull out of UI during sit/stand.
1577 | Screws attaching upper leg to hip has loosened | closed | jwitherspoon | Customer Use | fixed | 2014-02-19 19:14 | 2014-02-19 | 569000 | 35 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Customer Training | Hi Manuel, This morning when using Ekso we heard a noise coming from the right side. The Ekso started doing some kind of smooth movement anterior-posterior on the right side (maybe hip?). We find out that one of screws seems to be loosen and the head of this screw is little bit out in relation to the others (pictures attached). However, I'm not sure if this is the problem. My question is if can we keep using Ekso or we must stop? Let me know what can we do. | Customer was sent torque wrench set to 75in lbs, replacement screws and Loctite. They were able to perform procedure and hip screws are now securely fastened. | Customer loosened and then did not re tighten screws to necessary torque.
1553 | Right foot tends to internally rotate | closed | kmettler | Lower Leg | fixed | 2014-02-11 19:24 | 2014-02-10 | 123380 | 31 | Miami Project | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Unassigned |  | Customer Use | Customer reports its seem that with the loaner the right foot tends to internally rotate a bit (we have successfully adjusted the setting for our subjects to compensate) and also the step on the right (left stands) seems to come out short even when properly positioned/aligned. | We never had another report of slipping. SWO-00966 (2/9/2015) PERFORMED A LOT OF WORK INCLUDING CHECKING THE SHANK HOLDING TORQUE. | Floor type and how they adjust the legs.
1543 | Cructh interface no longer working- Step by Step 3rd issue | closed | jwitherspoon | UI - Crutch Interface | fixed | 2014-02-05 17:40 | 2014-01-30 | 615800 | 35 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | The Crutch Interface does not work anymore, so they have to initiate the first step from the User Interface. | SWO-501 was generated for this issue and customer has received new crutch interface. HMI has broken again and will be sent out on SWO-541. | Cable pull out from HMI crutch interface.
1542 | Stand to Sit- Step by Step 2nd issue | closed | bclinch | Battery | fixed | 2014-02-05 17:37 | 2014-01-30 | 615800 | 35 | Step-by-Step - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Materials | In the last two days Ekso they experienced that Ekso unlocked the hips while walking with one patient, without showing any Error code or giving a warning. With another patient they experienced today that whilst they sat the patient down, Ekso bent one leg and kept the other one straight, this caused that the patient lost balance. The Physio put her legs next to it save the patient from falling and got injured on her leg (a bruise on her leg). Manuel has given clarity to this issue. Customer wasn't in safe mode, they are reporting while doing a routine stand to sit one leg failed to bend like the other. | This issue was misreported and was caused by a battery failure. New batteries were sent via SWO 501. | Dead cells
1528 | Adjustment tools- Calibration date expired | closed | kmettler | Tools | fixed | 2014-01-30 20:45 | 2014-01-29 | 14043 | 49 | Mitch Brogan - Canada | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost |  |  | Process Not Clear/Complete | Step Count ‚Äì 14,043 PT at 456 Boler Rd today noted the calibration sticker on both T-wrench says it expired on 9-18-13 and hip wrench expired on 12-20-13. She asked if they are supposed to send back to us and we manage this. I know at one customer site I saw a release note about the customer needing to calibrate the tools themselves. I looked through all Boler Rd‚Äôs Ekso stuff and did not find that document here. Also checked with Mitch (owner), and he does not recall getting any info about calibrating tools. Please let me know how/what I should direct the customer on this. | Customer was sent a new tool kit and it was a non issue. No SWO created. | T-Wrench's were out of calibration
1527 | Error #21 reported- Joint position wasn't seen by PT on UI | closed | kmettler | Actuators | fixed | 2014-01-30 20:37 | 2014-01-29 | 12221 | 49 | Able Bionics - Canada | Mitch Brogan | ablebionics@gmail.com | +1 (226) 700-3654 | Canada | no injury | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Unassigned |  | Design | At 3:05 pm, EST today, Upon powering Ekso off, then right back on, between sessions, device got Error 21, sensor error. (PT did not notice which joint). Cleared it, and carried on as expected. Step count ‚Äì 12,221 | The error was cleared and the PT was able to carry on as expected | The device came up with a sensor error but performed the way it was supposed to.
1526 | DDL band appears to be off | closed | kmettler | Customer Use | fixed | 2014-01-30 20:32 | 2014-01-28 | 51000 | 14 | Spaulding | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Customer Use | I think one of the rubber bands at the hip joint is off a little ‚Äì kind of hard to tell. Pictures show spare DDL band as correctly installed. | Reworked per SWO-562, removed loose stretched out bad and replaced with new. | Lowering the thigh adjustment to low caused the band to brake
1509 | Double level walker- Wheel fell off | closed | kmettler | Cane / Walker / Crutch | fixed | 2014-01-22 17:13 | 2014-01-20 | 30826 | 24 | TIRR | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Materials | Hey team, Five steps into session with PT 1, front left wheel bolt came off. R wheel was loose as well. I forgot to send a ICORD TFF where the walker wheel was too tight and the wheel would not turn. We fixed that with lube to help the wheels roll. | NCR-10797 to notify vendor. Vendor adding loctite to bolts. Inspection done by QA currently. Send 263 Loctite to secure walker wheel. Fixed in field by PT during training. CO-11012 added inspection to MPI. | Need to add "walker" to root cause as a value. Vendor issue but added internal check via walker MPI release via CO-11012.
1492 | MP skin issue- breaking this ticket into 2 issues- #1451 address hex bolts | closed | ajulin | Soft Goods - Torso / Sacral | fixed | 2014-01-08 18:24 | 2014-01-07 | 150000 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Customer Use | I stayed out of the Ekso over Christmas to let my right hip fully heal. It has and there is only a small bit of darkened skin scarring. And, yesterday I got in it for the first time. Unfortunately, my left upper hip at the front got grazed. I've attached a photo. We think that my T-shirt sseam grazed the area during the 60 min walk due to the loosening of the upper part of the unit i.e. the hex bolts are loosening every 10 mins now. The clamps seem to be holding the hip width but the hex bolts are loosening significantly which seems to make the device feel loose to me and to Simon. This looseness is the only thing that has changed out of the entire 150,000 + steps that I've taken in all devices. I have worn the same shoes and trousers and boxer short design every time I've walked in Ekso since Jan 2012. Any thoughts as I'm really keen to keep going but fear we're getting to a 'no walking' position to avoid injury. Thanks Mark | Customer setup was an issue and Ekso clinical member has done a full eval with Mark. They went over proper technique and made changes to his Ekso configuration. All of the details are below in the report Dennis sent over from customer site. See: CO-10949 SWO-00537 | Customer clothing caused skin damage.
1491 | #20/21 LK sensor error | closed | ebabaran | Ball Screw | fixed | 2014-01-08 18:16 | 2014-01-07 | 23049 | 43 | Sunnaas - Norway | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Design | We have come across this error today: #20/21 LK sensor error This initially happened just before toe-off on the left. The patient is 90 kgs and had a lot of spasticity today (MAS 3 in knee flex left=right) but other than that was walking well in Prostep bilat max assist, forward shift 2, lateral shift 1. We did a long-leg sit, rebooted and tried again. The same happened initially during sit to stand (beginning of standing phase) and then a third time after reboot after about three steps. I then went into Ekso and tried it out. Sit to stand was ok, walked about 20 steps and got the #21 LK at toe-off, tried again and got #LK21 in mid swing after about five steps. Please advise on best cause of action. Can one of you email Gyri the head physio and arrange a phone call with her? I know the time difference is a lot and she does not work tomorrow (wed) but will be in Thursday 8-15.00 CET. | Unit found to have broken ballscrew. Unit fixed in house per SWO-533. | Broken ball screw, known issue.
1490 | Missing and Mismatched battery error | closed | kmettler | Battery | fixed | 2014-01-08 18:04 | 2013-12-20 | 181582 | 18 | RIC | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single |  |  | Design | Just powered on RIC‚Äôs old ekso to test foot sensors after talking to Edmundo to reassemble foot. Mismatched battery error-- though individual battery leverls are 93% and 94 % Step count ‚Äì 181,582 | The bad batteries were hand carried back to HQ and the customer received new batteries via SWO-1207 | Could not reliably reproduce problem. See comment:6, comment:8, comment:9
1482 | Crutch interface strap broken | closed | jwitherspoon | UI - Crutch Interface | fixed | 2013-12-18 18:57 | 2013-12-18 | 432000 | 12 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 5 - Mailed Part - Med Cost |  |  | Design | Good Shepherd just reported the band on their CI just broke and they'll need a new CI. | Sent new strap with SWO-477 | Reattaching/detaching CI can cause wear on strap, causing it to break.
1464 | Batteries draining quickly | closed | jwitherspoon | Battery | fixed | 2013-12-12 18:51 | 2013-12-09 | 132000 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 6 - Mailed Part - High Cost |  |  | Customer Use | Our batteries are running out very quickly. We only get 40 mins walking at 500 steps each 10 min block. So, after 40 mins we have to stop, change batteries and reprogram all of our settings. Could we get new batteries? | Customer received 4 new batteries to replace his old set. SWO-580 was generated for these replacement batteries. | Customer walks much more steps than he originally did. This was causing the perception that batteries were draining quicker than they did when unit was first delivered.
1451 | Hex bolts-Loosening during session | closed | kmettler | Adjustments - Hip | fixed | 2013-12-10 22:12 | 2013-12-09 | 132000 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple |  |  | Design | Every 20 mins or so we have to tighten the device which has been ongoing for a good while now. But as you will see from the photo attached I have now grazed my right hip as a result. > > Yesterday I did over 3,000 steps and the thing was loosening throughout the session. And, Simon was tightening throughout the session. And, when the batteries ran out after 40 mins (see other email about this) we tightened again and got going. > > Simone noticed my hip this morning . The skin is broken. > During follow up with Mark it was discovered this was discovered: We've noticed that the left hip width is at 8.5 and the right is at 8. We do not know when this happened. Since you added in the combo of the clamps and the hex bolts we have done the loosening of the hex bolts and clamps followed by tightening the clamps and then the hex bolts. But we still get the loosening. It has also been confirmed by Simon they are tightening the bolts properly-hex bolts followed by secondary clam. | Fixed by implementing nut locking insert per CO-10949 and specifically on this unit by SWO-537 | Hex bolt can loosen during use.
1444 | Missing right battery alarm | closed | kmettler | Battery | wontfix | 2013-12-04 18:49 | 2013-12-03 | 11338 | 52 | Walter Reed | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Customer Use | Walter reed step count 11338 While in Stand Still with PT in device in pro plus, battery level had read 59%. Got missing right battery alarm with right battery fully engaged. About screen battery level right 0%, left 59%. Right Battery fully engaged in ekso torso. I was holding remote, teaching a point, and no movement had been occurring when alarm signaled. Had to safe mode sit. Reboot and device operating properly, overall level was 62% on power up. Batteries even at 62. | Issue was monitored and did not reoccur | Battery may not have been fully engaged as indicated by customer.
1378 | Load last settings issue | closed | adamzoss | Software | fixed | 2013-11-06 18:30 | 2013-11-06 | 113491 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 1 - Phone Calls - Single | Minor |  | Design | Sandy was letting us know that Mark‚Äôs device does not load last settings when using VA. I already let her know it‚Äôs a known issue that will be fixed, and to see if setting him at .9 will change it. Same issue as Amanda. Bianca has advised Sandy this is a known issue and avoided if the swing time is more than .8 | The range checks on loaded parameter values now accept values slightly smaller and larger than the valid range to account for possible float rounding issues. | It occurs on at least one of our stored setting values. The values are scaled to ints, stored in NV, and then when they are read and scaled back to floats upon "load last settings" the value fails the range check and thus fails to load.
1373 | Ekso threw an #22 LH sensor error | closed | ebabaran | Don-Doff Link | fixed | 2013-11-06 16:34 | 2013-11-06 | 113491 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Minor |  | Design | In today's training session Mark was doing well when mid flow we heard a crack coming from the left hip and Ekso threw an #22 LH sensor error. We sat down in safe mode but Ekso requires Joint Test diagnostics. No obvious damage or noise when passively moving the leg but we noticed on Monday that the left leg generally sounded less smooth than the right. As it is still in the middle of the night in the US and my flight back to the UK is in a couple of hours I'm unable to go through the diagnostics. (I am travelling again tomorrow morning so couldn't delay my flight either) Can one of the engineers call Mark and walk AJ or Simon through the diagnostics or discuss the next steps with them? | Superceded by DDL fixes applied by SWO-492 | DDL wall not thick enough.
1371 | Sensor Error #21 RH | closed | kmettler | Electrical Interconnect | fixed | 2013-11-05 17:45 | 2013-09-09 | 13931 | 38 | VA Tampa | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Minor |  | Workmanship | In short, because I can't step away too long, I got a sensor error #21-RH. After rebooting I can't get the error to clear. I even tried under the technician screen to recalibrate the hip/knee sensors that all read PASSED. I still got the hip error. Luckily Tampa has 3 backups that we will use for the last patient but I leave in 2 hours and would like a answer to help them out. Customer has 4 devices and the issue occured on the device designated by customer as #4. | MFG found Loctite in one of the connectors. This part was replaced and error was cleared. SW0-457 shows the work completed to repair unit. | Dog house connectors
1356 | R knee abruptly into extension | closed | kmettler | Customer Use | wontfix | 2013-10-29 02:14 | 2013-09-09 | 377556 | 12 | Good Shepherd - OP | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 2 - Phone Calls - Multiple | Minor |  | Determined Unknown | Hey Josh, When walking today the R knee appeared to abruptly go into R knee extension, making a quick jerk. Today is the first day i'm noticing it. I'm going to send you a video to see if you have any suggestions. I know Melissa will be here tomorrow, maybe she can check it out too? Thanks, Nicole | could not identify issue, issue has not recurred. | Error not present in data and not repeated.
1337 | UI button not working | closed | kmettler | UI - Controller | fixed | 2013-10-17 21:04 | 2013-10-17 | 30929 | 33 | WakeMed | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 4 - Mailed Part - Low Cost | Unassigned |  | Materials | Just spoke to our new contact at WakeMed Health and they are having an issue with their UI. One of the buttons has become intermittent and when triggered it will stay depressed and not return to proper position. It‚Äôs affecting their walking sessions | SWO-440 was opened and new UI was sent to customer. They have reported no further issues since failure. | Customer abuse/misuse caused failure.
1335 | Don/Doff lower link fracture | closed | bclinch | Don-Doff Link | fixed | 2013-10-17 20:49 | 2013-10-17 | 113491 | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 8 - Field Service | Major |  | Design | I just hit my record steps yesterday and today at about 3 mins into the session we heard a large crack. A piece on the left hip/upper leg cracked. I have attached the photos and no error message came up but it is very definitely out of order! It is physically cracked. Also, the software doesn't allow us to select previous settings since Greg was over. Simon has to re-enter the settings every time. | Repaired multiple times with SWO-442 (steel insert) and SWO-492 (thick wall DDL) | DDL did not meet strenght requirements.
1284 | Guttman- noise from knee | closed | duanee | Actuators | worksforme | 2013-09-03 16:16 | 2013-09-03 |  | 46 | Guttmann - Spain | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 9 - Loaner | Unassigned |  | Solution Not Implemented | 1 - "The right knee engine makes a strange noise" 2 - "In the sitting maneuver, the system will get stuck for a while, making little fluid manoeuvre" This is all information I have. Manuel will follow up with the clinic and then we can update ticket with appropriate information. | Unit was updated to VA per SWO-296. Issue was not reproducable in house. Upgrades are felt to have addressed issues seen by customer. Have not heard back from customer regarding this issue. | Issue was not reproducable in house. Upgrades are felt to have addressed issues seen by customer.
1240 | Thigh strap extensions failed with patient | closed | kmettler | Soft Goods | fixed | 2013-07-03 15:30 | 2013-07-03 |  | 45 | Justin Smith - So. Africa | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 4 - Mailed Part - Low Cost | Minor |  | Wear | Please see the attached video. I have never seen this happen before. I completed level 2 today in South Africa and they asked if they could bring in a patient for me to trouble shoot with as they are having difficulty with him. In short he has an increased tonal pattern in the lower limbs with MAS level 3 in the major muscles. He is pushing max height and max weight also. The thigh girth is such that we needed to use the extensions on the strap and thigh brace. We assessed and all appeared fine with no contra-indications. I was aware of the tonal issues and made adjustments in the setting (stand time, swing time and built a small degree of knee flexion in to Ekso) to try and accommodate this. We got 3/4 of the way up into a stand and the Velcro on the extension straps bilaterally failed. We paused and sat down with no injury to the patient or therapist. We set up the ekso again as per usual and the same thing happened again. We had more than enough capture on the extension to the thigh brace bilaterally. However the Velcro didn‚Äôt seem to hold in place (see video). We were unable to walk this patient. I would suggest we send out some new extension straps and get these back for analysis. For your information the SN of the device was 011211002 with 11443 steps. The problem occurred in sit to stand. Unfortunately I cant recall the exact patient leg length measurements but they were reaching max. | Final Extender redesigned with CO-10992 & CO-11014 - REV F CO-10950 - Temporary fix increasing length of extenders to allow for more overlap. SWO-00531 installed new thigh straps | Known wear issue
1234 | PT finger pinched between Ekso and chair | closed | kmettler | System | fixed | 2013-06-28 17:38 | 2013-06-28 |  | 51 | QLI | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 2 - Phone Calls - Multiple | Major |  | Mishandling | We had a sit to stand issue here at QLI during level 2 testing. The PT was working with a tall patient (same patient that created stand error on 6/27) and is using a chair that has been in play since the initial training. During the lean forward, the chair tipped forward and the patient almost fell to his knees. The PT did a nice job of pulling the patient back into the chair and in doing so, smashed her right first finger between the Ekso and the chair. The PT was crying and needed a few minutes to gather herself after this issue. She completed the session with discomfort and has at least a sprain and possibly a fracture based on stress testing the finger. Reported by Darrell Musick. | The pilot was trying to push the opposite direction of the machine and the user stumbled but there was no harm to the patient, however the PT crushed her finger in the process. The arm sling was also not being used and should be used in the future to prevent injury. | Instructions for use not followed.
1177 | FLM has a 2nd faulty battery | closed | kmettler | Battery | fixed | 2013-06-03 16:11 | 2013-06-03 |  | 44 | FLM - Spain | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 3 - Next Service | Unassigned |  | Design | Manuel called me today to report that FLM in Madrid has had a second battery failure. I had him report the Serial Number (017C4C201500001F) and Part Number (102505). Based on the part number it is my understanding that this is a new battery board. He proposed I go to Madrid to fix it but I am not sure how to fix it, I have a few modified battery boards remaining | All batteries go through a refurb process because the Unit was sent back and they upgraded to a GT Unit | Battery Failures, the issue is unknown now however the device and batteries were retrofitted.
1161 | Fractured Actuator Cart | closed | kmettler | Actuators | fixed | 2013-05-21 15:28 | 2013-05-21 |  | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | 10 - Field Service & Loaner | Minor |  | Design | So I met Mark this morning for testing for the creakiness. I applied vibratite first to the plastic and carbon covers and flexed the legs laterally slightly and found the creaking seemed to be reduced. Then when testing with Mark during the initial stand I couldn‚Äôt really tell what went wrong at the time, he just leaned forward for the three beeps then when he started to stand and he just rolled forward with more force than I could resist; I ended up lowering him down into a crouching position. His assistant and I then picked him up and put him back in the chair. We inspected his knees and hands and could find no bruises or red marks; luckily he rolled forward very slowly and I was able to lower him down gently. As neither of us understood what went wrong, we decided it was not a good idea to attempt again; we took Mark out of the Ekso and back into his chair. He seemed in very good spirits and said it was a good experience and good to know that Simon was doing more than we thought (as we assumed it must be something I had done wrong). We decided that I should try and tackle the creaking without walking him and wait for Simon to return on Saturday. Upon investigating the Ekso I found the right hip actuator cart hard fractured; this is likely the cause of Mark rolling forward as the hips lock up in the flexed position and the knees continue to roll forward. Katie, is this possible or will the knee stop extending if the hips lock up? I had inspected the cart during the retrofit and couldn‚Äôt see any signs of cracking; I think either: - I missed signs of fracture during retrofit - I induced the fracture while applying lateral force before testing to evaluate crackling - The cart had fatigued and I was just unlucky (or lucky because I‚Äôm still in Ireland) that it fractured at the first stand I attempted Luckily, I have access to the Ekso for the rest of the week. I can get in as early as 7AM each day and work until as late as 10PM. As the fracture may be caused by fatigue I think I should replace both right and left hip actuator carts. My proposed schedule is: Today: Work with Turbo to create a procedure for hip actuator cart field replacement Tomorrow through Friday: Replace carts, implement additional gusset cover fix. One Concern: This unit has the hip hard stops. My guess is the lateral loading of the actuator cart due to the difference in flexibility between the joint and the rail through the links or the actuator housing may be causing this. I think this failure in a unit with the hard stops implemented makes a very strong case for implementing the steel actuator cart across the fleet. Reuben, please work with Patrick to determine if this is possible however I understand it probably impossible to fabricate while I am still in Ireland. Beth, please check with Turbo if there are any other tools I need to replace actuator carts; I think I have everything and already have two actuator carts in my spare kit. | CO 10720 replaces the aluminum actuator cart design with a steel version which is ~8 times stronger. Unit fixed with SWO-00309. | Strength of aluminum actuator carts was estimated without accounting for lateral flex during anomalous sit to stands. Strength requirements were correspondingly set below real world usage requirements.
1077 | UI continues to read right foot stepping | closed | kmettler | Foot Sensor | fixed | 2013-03-22 16:20 | 2013-03-22 |  | 44 | FLM - Spain | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Wear | They stand someone up for walking and everything is alright. But when they start walking and do the FIRST right step Ekso only puts the right foot forward and stops! Ekso doesn't allow the user to come forward over the right foot and get the weight forward. The display stays in "stepping right" though the step is taken already. It is not possible to trigger the next (left) step. They tried it with different patients and with the physios as well. It happened every time. They say their is no obvious mechanical problem visible. Reported by Dennis Veit Ekso EU PT. | CO 10628 updates the foot sensor assembly so that the wire routing of the hook up cable is in a position not prone to fatigue. | It is most likely that this is caused by a high resistance solder joint due to fatigue. This joint will be reworked along with the other changes made in CO 10628.
1041 | Patient fell at RIC | closed | kmettler | Clinical Training | fixed | 2013-02-13 20:40 | 2013-02-13 |  | 31 | RIC | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 2 - Phone Calls - Multiple | Major |  | Mishandling | I spoke with a PT at RIC who reported to me that they had a patient fall down forwards, all the way down to the ground last Thursday. Occurrence: On Thurs, Feb 7 Patient (pt) is a strong para; has walked in ekso before Was practicing turning in a small area to prepare for safely turning inside an elevator pt was off tether, on crutches, in ProStep pt had been instructed before to not turn/unweight until they were in Stand Mode PTs know to have pt complete a tight turn in Stand Mode, and identified that they had told the pt this before, yet he went too quickly While still in Walk Mode, pt lifted himself off ground to turn, depressing through his crutches; per PT estimate, about 6-8 inches of ground clearance likely occurred Legs of Ekso split apart into split stance pt lost balance forward, and went down to ground A tall PT (6‚Äô2‚Äù) was guarding pt, and was able to slowly lower him. Pt also caught himself first on a nearby chair with one arm, then braced other arm on ground; ended up on ground, belly down, on both his forearms. They rolled him on to his side, then on to back (in ekso), then went to safe mode, to Free, then transferred pt up into chair. Neither patient nor PT hurt. Ekso not damaged on their inspection. They have used Ekso 3-4 time since this incident, without trouble. This Ekso that hit the ground was the LOANER unit they have, and not their own device. Clinical Input on how/why this happened: EB clinical input/conversation: -sounds like multiple factors lead to this loss of balance and fall: 1) pt unweighting before entering Stand Mode; 2) pt unweighting too high; 3) crutch placement may have helped OR hindered the loss of balance- and this cannot be determined at this point as RIC PTs are unaware of how placement was at the time. -good that the RIC PTs do know the correct sequencing and safety concerns for correct teaching -this highlights need for clear communication and instruction, and how patient interaction/behavior is as important a factor -Clinical and Controls team started discussions into whether a software change could help identify unweighting when not appropriate, or other considerations surrounding turns. Customer input/feedback: Their input/comments: -They did not have to report this to their IRB board; they have concern though about keeping this patient in their research protocol if he continues this behavior. 1) Coordinating PT feels that because a tall PT was guarding, they had the best chance for success with catching/controlling the pt fall. Concerned how it would have turned out had the PT not had height advantage. 2) Coordinating PT asks: is there a code that could be written into software to prevent ‚Äúthis‚Äù from happening | Multiple factors lead to this loss of balance and fall: 1) pt unweighting before entering Stand Mode; 2) pt unweighting too high; 3) crutch placement may have helped OR hindered the loss of balance- and this cannot be determined at this point as RIC PTs are unaware of how placement was at the time. | Customer training
1032 | RIC 164 error CI | closed | kmettler | UI - Crutch Interface | fixed | 2013-02-08 20:14 | 2013-02-08 |  | 31 | RIC | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | RIC called and said they are getting error 164, crutch interface isn't connected. We are sending a new CI with their unit and this is going out today 2/8/13. This is happening on a loaner unit. #31 Corrective Action: The miss functioning crutch interface is being brought back for analysis. If an underlying failure mode is found the design will be updated to fix the issue. | CO 10592 updates the software to improve the electrical stability of the HMI detection testing in software 1.1.3 | This issue was subsequently investigated again and found to be a heat related electrical instability in the HMI board.
958 | Shark fin causing cut on Mark P. Ekso | closed | ajulin | Soft Goods | fixed | 2012-11-29 20:24 | 2012-11-29 |  | 39 | Mark Pollock - Ireland | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | Unassigned | Major |  | Design | Email from Mark: The only other thing that we have noticed is to do with the little ‚Äòshark‚Äôs fin‚Äô which is just below the hip joint when sitting. With the legs open wide in preparation for the transfer the top (the point) of the shark‚Äôs fin is exposed. I can stop myself hitting the fin with the leading hip (left hip to left shark‚Äôs fin for example. But the trailing hip sometimes lands on top of the sharks fin. Right hip on right shark‚Äôs fin. I‚Äôm not quite sure how to fix this problem. Maybe padding it a bit more. But I think it might need to bend out of the way completely. I cut myself slightly on it and couldn‚Äôt work out for ages what had happened. Corrective Action: Shark fin design is being reviewed to determine if changes can be made to eliminate this issue as well as improving the robustness of the rubber band retention. | This edge was trimmed during SWO-996. No design changes needed | The shark fin design includes a pointy edge which caused the minor scrape reported.
846 | Crutch interface connectivity issues at RIM | closed | kmettler | UI - Crutch Interface | fixed | 2012-10-05 16:35 | 2012-10-05 |  | 19 | RIM | [5 empty] | Same as contact listed | Pro 1.0 | FALSE | 5 - Mailed Part - Med Cost | Minor |  | Design | Email from Darrell at RIM: The crutch interface appears to have a loose pin or wire. We are getting the #164 error (CI unplugged) 7-15 times before we are able to get it to work. Of note, the light has a faint glow when plugged in, even the alarm is sounding. Moving the wire at top appears to eventually solve it. We need a new CI ASAP. Corrective Action: Edmundo to investigate failed crutch interface and implement changes to the MPI which elliminate the failure mode. | Customer hasn't seen this error since Edmundo performed the fix. 7-29-2013: CO 10652 releases software 1.1.4 which resolves this issue by addressing the thermal instability of the crutch interface detection. | Ultimately, it was found that this error was due to software and electronics becoming unstable at higher temperatures.
710 | Plastic clip popped off chest strap piping in Miami | closed | plong | HW - System | fixed | 2012-06-27 16:51 | 2012-06-27 |  | 25 | Ekso Bionics | [5 empty] | Same as contact listed | ekso 1.1 | FALSE | Unassigned | Unassigned |  | Undetermined | Original email from Melissa: Hey team, This morning we had a failure of the plastic clip that holds the chest strap to the backpack, along the red piping. The chest strap was fastened, and the backpack straps were about to get tightened. The plastic clip piece popped off the red piping when someone was tightening the backpack strap to secure to patient. The plastic piece did not break. To manage it for the time being, we have tied to the whole chest strap around the backpack strap. We need to know how to replace that clip, if it can be done in the field. I'm going to call in to the red phone to report this now as well. See attached picture of the intact clip....not sure if it has widened out a bit from the pull off. Melissa | The latest version of the piping includes a section that allows for easy re-installation of the clip.